Guidance for the Selection of Analytical Methods and the Evaluation of MDLs and PQLs

List Referenced in Chapter 62-4.246(4), F.A.C.

Prepared by

Division of Environmental Assessment and Restoration

November 10, 2020

# Intended Use of This Document

This guidance is published to assist DEP staff and other data reviewers with the evaluation of appropriate analytical methodologies used to generate water quality data.  Subsection 62-4.246(4), Florida Administrative Code (F.A.C.), requires DEP to establish and publish a list of target Method Detection Limits (MDLs) and Practical Quantitation Limits (PQLs) for published analytical methods that may be used for the analysis of wastewater, surface water and groundwater.  Data reviewers must compare MDLs and PQLs reported for permit compliance data with the listed MDLs and PQLs to ensure that adequately sensitive analytical techniques were utilized to meet permit requirements or other data quality objectives.

If the target MDL or PQL was not achieved for a given report, there may be valid reasons why this occurred, and the reviewer should refer to justification provided by the laboratory. Situations that may prevent achievement of target limits include, but are not limited to, environmental samples that contain constituents that cause analytical interferences or other quality control problems and sample collection anomalies (e.g., the collection of less than ideal sample volumes). In some cases, DEP may approve alternative MDLs and PQLs for the duration of a specific permit if sufficient justification is provided. Alternative MDLs and PQLs may be approved if alternative target limits remain below compliance limits or if matrix-specific MDLs and PQLs are developed.

# Sources of Listed Target MDLs

The listed target MDLs were obtained from the following sources:

- MDLs published in approved methods. The EPA and the APHA (“Standard Methods”) were the primary sources used.

- MDLs published in the Environmental Monitoring Methods Index (historical); and, the National Environmental Monitoring Index (NEMI).

- Consultations from laboratories in Florida, including the DEP laboratory in Tallahassee.

MDLs from the sources above may differ for the same analyte-technique combination. In these cases, the MDL included in the 62-4.246(4) list was judged to be the most appropriate target value for the listed method.

# Calculation of Target PQLs

- Except as noted below, many of the listed PQLs were calculated as 4 X MDL, per the default calculation of the PQL indicated at subsection 62-160.120(23), F.A.C. For some listings, the PQL is taken from the information in the method about minimum levels applicable to quantitation.

- PQLs are listed as equivalent to the MDL or as some other factor of the MDL for certain analytes, where the detection or quantitation value is affected by practical considerations associated with the analytical technique or where the concept of a statistically derived MDL may not be appropriate. The analytical techniques for these analytes are specifically designated in the MDL/PQL list as described in the next section. Examples of these analytes include:

Biochemical Oxygen Demand (BOD)

Carbonaceous Biochemical Oxygen Demand (CBOD)

Chlorine, Total

Chlorophyll

Coliforms, Fecal and Total

Color

Dissolved Oxygen (DO)

E. Coli

Enterococci

Hardness (by calculation)

Hydrogen ion (pH)

Hydrogen sulfide

Nitrate (by calculation)

Nitrogen, Total (by calculation)

Odor

Radiochemistry analytes

Residues (TDS, TSS)

Specific conductance

Temperature

Titrimetric methods

Turbidity

Un-ionized ammonia

# MDL/PQL Listings Where MDL May Not Be Applicable

An “X” in the column titled “MDL Not Required” indicates an analytical technique for which the EPA procedure for MDL determination may not be appropriate (see 40 CFR, Part 136, Appendix B). The laboratory may instead indicate the PQL or laboratory Reporting Limit as an equivalent detection level for a specified analyte, as determined by alternative procedures. Field organizations may report field test results as a minimum increment based on the field sensor specifications. Per paragraph 62-160.340(3)(c), F.A.C., analytical results below the laboratory reporting limit shall be indicated as the reporting limit value followed by the data qualifier code "U" (see Table 1, Rule 62-160.700, F.A.C.).

# Reported Versus Target MDLs and PQLs

The listed MDLs and PQLs are targets that should be achievable under optimal conditions of quality control by most environmental laboratories. However, the laboratory cannot always be expected to meet the MDLs or PQLs listed in the table. Often, environmental samples may contain constituents that cause analytical interferences or other quality control problems that can prevent achieving the target limits (e.g., cyanide, semi-volatiles, and phenols). Even sample collection anomalies (e.g., the collection of less than ideal sample volumes) can affect the MDL or PQL reported by the laboratory. Laboratory representatives should be able to explain the rationale behind the MDLs and PQLs reported with the analytical results, especially if the reported MDLs and PQLs are above the required compliance limits. If necessary, data reviewers should request documentation that demonstrates how the reported MDLs and PQLs were derived or adjusted. In general, laboratories are required to calculate or determine MDLs and PQLs using established procedures that are required by rules (e.g. 40 CFR, Part 136, Appendix B), certification standards (e.g. TNI, 2016), and the laboratories’ internal Standard Operating Procedures. Note that some laboratories may use allowable modifications or optimizations of approved analytical methods or use improved instrumentation that provide MDLs and PQLs lower than those found in the 62-4.246(4) list.

# Reporting Data Below the MDL or PQL

Chapter 62-160, F.A.C., addresses reporting requirements for data submitted to the FDEP programs. Laboratory analytical reports should present the sample data in accordance with these requirements. Additionally, data reporting instructions specific to DEP-required forms must be followed for regulatory compliance.

“Non-detect” results

Whenever an analyte is not detected at or above the MDL, the MDL for the measurement must be reported along with the qualifier code “U” indicating that the analyte was not detected at the reported detection limit.

Results reported below the MDL

Alternately, laboratories have the option of reporting the analytical value from the sample analysis followed by the qualifier code “T” indicating the reported value was below the MDL.

Results between the MDL and PQL

If an analyte was detected at or above the MDL but was below the PQL, the result may be reported in one of two ways:

  - the value from the sample analysis can be reported with the qualifier code “I,” indicating that the analyte was detected but could not be quantified with certainty.

  - the PQL associated with the sample analysis can be reported followed by the qualifier code “M,” indicating that the analyte was detected but was below the reported PQL.

Refer to Table 1 in Rule 62-160.700, F.A.C., for a complete list of appropriate data qualifier codes.

If the analyte is not detected at or above the MDL or is detected at a level below the PQL, the reviewer should compare the reported MDL and PQL with the 62-4.246(4) list to determine if an appropriately sensitive method was utilized for the permit or other project objective. If the target MDL or PQL was not achieved, there may be valid reasons why this occurred, and the reviewer should refer to justification provided by the laboratory.

# MDLs and PQLs for “Calculated” Quantities

Some analytes are not tested by direct analysis but are calculated from the results of other analyses; for example:

Hardness (by calculation)

Hydrogen sulfide

Nitrate (by calculation)

Total nitrogen (by calculation)

Un-ionized ammonia

MDLs and PQLs for calculated results are therefore not directly determinable. Limits reported for these analytical quantities are dependent on the limits for the component analyses and the number of significant figures carried in the calculations. The “MDLs” and “PQLs” in the 62-4.246(4) list for these calculated analytes were selected based on the above considerations.

# Laboratory Certification Requirements

Laboratories must be certified by the DOH Environmental Laboratory Certification Program (ELCP) for all reported matrix-analyte-method combinations. Per Rule 62-160.300, F.A.C., DEP may temporarily waive this requirement while the laboratory pursues certification with the DOH ELCP.

# Revisions to Analytical Methods

All citations of analytical methods in the 62-4.246(4) list refer to the edition of the method listed in 40 CFR Part 136.3. If the method is not listed in this federal rule, the most current, published edition of the method should be used.

Letter suffixes associated with EPA SW-846 method revisions have been omitted from the applicable entries in the 62-4.246(4) method list. The most current published edition of the SW- 846 method should be used.

# Use of Additional Methods in 40 CFR Part 136.3 Not Found in the 62-4.246(4) Method List

Laboratories may use additional analytical methods that are approved in 40 CFR part 136.3 but which are not found in the 62-4.246(4) list.  The MDL and PQL for the alternative method must be at least as sensitive as the MDL and PQL for the equivalent analytical technique listed in the 62-4.246(4) list, if applicable.

# Use of Methods Not Found in 40 CFR Part 136.3

If no analytical method is listed in 40 CFR Part 136.3 for the analyte, the laboratory may use any method specified for the analyte in the 62-4.246 list, in any applicable DEP rule, or as approved by the applicable DEP regulatory program.  In some instances, approval for use of these alternative methods may be required by EPA, and approval must be obtained prior to analyzing samples for NPDES permits.  The MDL and PQL for the alternative method must be at least as sensitive as the MDL and PQL for the equivalent analytical technique listed in the 62-4.246(4) list, if applicable.

# Analysis of Samples for Gross Alpha Activity with High Total Dissolved Solids (TDS)

The co-precipitation methods for gross alpha analysis found in the 62-4.246 list may be used instead of 40 CFR Part 136.3 methods for high-TDS samples where the sample matrix is otherwise expected to raise the MDL. For compliance with NPDES permits, the laboratory should contact DEP with a request for approval of co-precipitation methods before analyzing compliance samples. The Department will forward the request to the EPA Regional coordinator for review as an alternative test procedure or method modification, per 40 CFR, Sections 136.4 and 136.6.

# Reporting Total Nitrogen

Total Nitrogen is defined as the sum of the results of the sample analysis for total Kjeldahl nitrogen (TKN), nitrate (NO3) and nitrite (NO2), using approved testing methods to perform the calculation.

Direct-measurement techniques for the analysis of Total Nitrogen are described in some published methods and may be approved for specific uses of data by DEP programs, such as for the assessment of surface waters. The laboratory should contact DEP for approval prior to using direct-measurement methods for Total Nitrogen analysis. Approval by EPA of a direct-measurement test method for Total Nitrogen may be required for compliance monitoring for NPDES permits.

All concentrations for Total Nitrogen should be reported as nitrogen (N) in mg/L.

# Reporting Nutrients for Ambient Samples and Compliance Monitoring

In permits with effluent limits for nutrients, the effluent limits will often be significantly higher than the nutrient criteria in Chapter 62-302, F.A.C.  In these cases, the MDL selected should provide detectable results below the effluent limit instead of below the nutrient criteria, and all target MDLs on the 62-4.246 list likely meet that description. However, if the permit requires ambient monitoring of the receiving waters, the MDLs for that monitoring should be selected to provide detectable results below the applicable criteria, which for some waters (many estuaries, for example) would require selecting the method with the lowest MDL. If there are no permit limits for nutrients but nutrients are monitored for “report only,” methods from the list with higher MDLs may be used.

# Achieving Lower MDLs for Polycyclic Aromatic Hydrocarbon (PAH) Methods

Some methods for PAHs have an alternative prep method known as the Selected (Single) Ion Monitoring (SIM) method with which lower MDLs can be reached. Laboratories may use this prep method if DEP requires MDLs and PQLs lower than those found in this list for permit compliance or other objectives. Methods with the SIM option include: EPA 1625B, SM 6410 B, EPA 8270E, EPA 8260D, EPA 624.1, and EPA 625.1. Lower MDLs and PQLs achievable with the SIM option included in this list are EPA 1625B, SM 6410 B, and EPA 625.1.

# How to Report Total Polycyclic Aromatic Hydrocarbons (PAHs) When One or More Were Non-detect

The water quality criterion for total polycyclic aromatic hydrocarbons (PAHs) in Chapter 62-302, F.A.C., is the total of the results for the following ten PAHs:

Acenaphthylene

Benzo(a)anthracene

Benzo(a)pyrene

Benzo(b)fluoranthene

Benzo(ghi)perylene

Benzo(k)fluoranthene

Chrysene

Dibenzo(a,h)anthracene

Indeno(1,2,3-cd)pyrene

Phenanthrene

When an individual PAH result is below the MDL, use a factor of 1/20th the MDL or 1/20th the water quality criterion (WQC), whichever is lower, when calculating the total PAHs.  This modification from Rule 62-4.246(8), F.A.C., is necessary because it is possible that the sum of ten individual PAH values, each set to the lower ½ the WQC or ½ the respective MDL, could exceed the WQC even if all the component PAHs were not detected. The modified approach applies 1/20th of the WQC or 1/20th of the MDL, whichever is lower, to replace all “U” qualified data.  Data qualified with an “I” are handled consistent with the rule; the reported value is replaced with the MDL.  An MDL for total PAHs is not determined as each individual PAH has its own MDL.

# How to Report Total Chlorinated Phenols and Chlorinated Cresols When One or More Were Non-detect

The water quality criterion for total chlorinated phenols and chlorinated cresols in Chapter 62-302, F.A.C., is the total of the results for the following seventeen phenols and cresols.

| 4- | Chloro-2-methylphenol |
| --- | --- |
| 4- | Chloro-3-methylphenol (4-chloro-m-cresol) |
| 2- | Chloro-5-methylphenol |
| 3- | Chlorophenol |
| 4- | Chlorophenol |
| 2,3- | Dichlorophenol |
| 2,5- | Dichlorophenol |
| 2,6- | Dichlorophenol |
| 3,4- | Dichlorophenol |
| 3,5- | Dichlorophenol |
| 2,3,4,5- | Tetrachlorophenol |
| 2,3,4,6- | Tetrachlorophenol |
| 2,3,5,6- | Tetrachlorophenol |
| 2,3,4- | Trichlorophenol |
| 2,3,5- | Trichlorophenol |
| 2,3,6- | Trichlorophenol |
| 2,4,5- | Trichlorophenol |

When an individual result is below the MDL, use the formula WQC/(2*n), where WQC = water quality criterion and n = total number of congeners reported, or MDL/(2*n) if the MDL is lower than the WQC. For example, if the total consists of all the analytes above, a factor of 1/34th the MDL or 1/34th the water quality criterion (WQC) is used.  This modification from Rule 62-4.246(8), F.A.C., is necessary because it is possible that the sum of individual phenol and cresol values, each set to the lower ½ the WQC or ½ the respective MDL, could exceed the WQC even if all the component phenol and cresols were not detected. The modified approach applies WQC/(2*n) or MDL/(2*n), whichever is lower, to replace all “U” qualified data.  Data qualified with an “I” are handled consistent with the rule; the reported value is replaced with the MDL.  An MDL for total chlorinated phenols and chlorinated cresols is not determined as each individual analyte has its own MDL.

# How to Report Total Phthalates When One or More Were Non-detect

The water quality criterion for phthalate esters in Chapter 62-302, F.A.C., is the total of the results for the following six phthalates:

Butyl benzyl phthalate

bis (2-ethylhexyl) phthalate (DEHP)

Diethyl phthalate

Dimethyl phthalate

Di-n-butyl phthalate

Di-n-octyl phthalate

When an individual phthalate result is below the MDL, use a factor of 1/12th the MDL or 1/12th the water quality criterion (WQC), whichever is lower, when calculating the total phthalates.  This modification from Rule 62-4.246(8), F.A.C., is necessary because it is possible that the sum of six individual phthalate values, each set to the lower ½ the WQC or ½ the respective MDL, could exceed the WQC even if all the component phthalates were not detected. The modified approach applies 1/12th of the WQC or 1/12th of the MDL, whichever is lower, to replace all “U” qualified data.  Data qualified with an “I” are handled consistent with the rule; the reported value is replaced with the MDL.  An MDL for total phthalate esters is not determined as each individual phthalate has its own MDL.

# How to Report Total Polychlorinated Biphenyls (PCBs) When One or More Were Non-detect

The water quality criteria for Total Polychlorinated Biphenyls (PCBs) in Chapter 62-302, F.A.C., are the total of the results of individual congeners tested.  The actual number of congeners analyzed for PCBs will vary, but when an individual result is below the MDL, calculate the total by setting the value of any congener below detection limit as = WQC/(2*n), where WQC = water quality criterion and n = total number of congeners reported, or MDL/(2*n) if the MDL is lower than the WQC.  This modification from Rule 62-4.246(8), F.A.C., is necessary because it is possible that the sum of individual values could exceed the WQC even if all the components were not detected. The modified approach applies the formula above, to replace all “U” qualified data.  Data qualified with an “I” are handled consistent with the rule; the reported value is replaced with the MDL.  An MDL for total PCBs is not determined as each individual PCB has its own MDL.

# List of Acronyms and Abbreviations for Analytical Techniques

The following definitions apply to the analytical technique descriptions found in the 62-4.246(4) list.

| A-GPC | Gas Flow Proportional Counter |
| --- | --- |
| AMP | Amperometric |
| A-SCINT | Scintillation Counter |
| Auto | Automated Analyzer |
| BAC-MF | Membrane Filter Technique |
| BAC-MPN | Most Probable Number (Multiple Tube/Multiple Well Plate) |
| CALC | Calculated by Indirect Analysis |
| Chelation-FLAAS | Flame Atomic Absorption Spectroscopy with Sample Chelation Preparation |
| CIE-UV | Capillary Electrophoresis-Ultraviolet Detector |
| COLOR | Colorimetric |
| COND | Conductivity |
| CO-PRECIP | Co-precipitation Sample Preparation |
| COUL | Coulometric Titration |
| CVAF CVAAS | Cold Vapor Atomic Fluorescence Cold Vapor Atomic Absorption Spectrophotometry |
| DCP-ES DIFFUSION-Auto | Direct Current Plasma-Emission Spectrophotometry Gas Diffusion – Automated Analyzer |
| FID | Flame Ionization Detector |
| FLAAS | Flame Atomic Absorption Spectroscopy |
| FLAAS-HYDRIDE | Flame Atomic Absorption Spectroscopy with Sample Hydride Preparation |
| FLUOR | Flourescence Spectrophotometry |
| GAS-FLOW-P | Gas Flow Proportional Counter |
| GC | Gas Chromatography |
| GC-ECD | Gas Chromatography - Electron Capture Detector |
| GC-ELCD | Gas Chromatography - Electrolytic Conductivity Detector |
| GC-ELCD-MCTD | Gas Chromatography- Electrolytic Conductivity or Microcoulometric Detector |
| GC-ELCD-PID | Electrolytic Conductivity and Photoionization Detector |
| GC-FID | Gas Chromatography-Flame Ionization Detector |
| GC-FPD GC-MS GC-MS (SIM) | Gas Chromatography-Flame Photometric Detector Gas Chromatography-Mass Spectrometry Gas Chromatography-Mass Spectrometry (Selected Ion Monitoring) |
| GC-NPD | Gas Chromatography-Nitrogen-Phosphorus Detector |
| GC-PID | Gas Chromatography-Photoionization Detector |
| GC-TBD GFAAS | Gas Chromatography-Thermionic Bead Detector Graphite Furnace - Atomic Absorption Spectrophotometry |
| GM | General Chemistry Method |
| GRAV | Gravimetry |
| HPLC | High Performance Liquid Chromatography |
| HPLC-UV | High Performance Liquid Chromatography-Ultraviolet Detector |
| HPLC-UV-FLUOR | High Performance Liquid Chromatography-Ultraviolet or Fluorescence Detector |
| Hydride | Flame Atomic Absorption Spectroscopy with Sample Hydride Preparation |
| IC | Ion Chromatography |
| ICP-ES | Inductively Coupled Argon Plasma Emission Spectrophotometry |
| ICP-MS | Inductively Coupled Argon Plasma Mass Spectrometry |
| IMS-FA | Immunomagnetic Separation-Immunofluorescent Assay |
| IR | Infrared Detector |
| IR-COUL | Infrared Detector or Coulometric Titration |
| IR-COUL-COND | Infrared Detector or Coulometric Titration or Conductivity Detector |
| IR-FID | Infrared Detector or Flame Ionization Detector |
| ISE | Ion-specific Electrode |
| ISE-Auto | Ion-specific Electrode-Automated |
| LC-ION | Ion Chromatography |
| LDO | Luminescence-Based Dissolved Oxygen Sensor |
| NESSLER | Nesslerization |
| POT | Potentiometry |
| SPEC | Spectrophotometry |
| SPEC-Auto | Spectrophotometry-Automated Analyzer |
| ST-GFAAS | Stabilized Temperature-Graphite Furnace Atomic Absorption Spectrophotometry |
| TEMP | Temperature |
| TITR | Titration |
| TITR-Auto | Titration-Automated Analyzer |
| TSGF | Thermally Stabilized Platform Graphite Furnace |
| TURBID | Turbidity/Turbidimetric |
| UV-VIS | Ultraviolet-Visible Detector |
| UV-VIS-Auto | Ultraviolet-Visible Detector-Automated Analyzer |
| Voltammetry | Pulse Anodic Stripping Voltammetry |

List of Acronyms for Organizations and Sources of Methods

   The following definitions apply to the acronyms for organizations cited in the 62-4.246(4) list.

| ANSI | American National Standard Institute |
| --- | --- |
| AOAC | AOAC International |
| ASTM | American Society for Testing and Materials |
| Bran Luebbe | Bran & Luebbe Analyzing Technologies, Inc. (Industrial Methods) |
| Brooks-Blanchard | “A Procedure for the Determination of 228Ra,” I.B. Brooks & R.L. Blanchard |
| CFR | Code of Federal Regulations (U.S.) |
| DEP | Florida Department of Environmental Protection |
| DOE | U.S. Dept. of Energy, Environmental Measurements Laboratory Procedures Manual |
| DOH | Florida Department of Health |
| EPA | United States Environmental Protection Agency |
| EPA, 1978 | EPA Reference Method* |
| HACH | “Hach Handbook of Water Analysis,” Hach Chemical Company |
| NCASI | National Council for Air and Stream Improvement, Inc. |
| OIA ALPKEM | OI Analytical/ALPKEM Corporation |
| OIC | Oceanography International Corporation |
| ORION | Orion Research, Inc. (Instruction Manual for 97-70 electrode) |
| SM | “Standard Methods for the Examination of Water and Wastewater” (APHA) |
| USGS | “Techniques of Water-Resources Investigation of the U.S. Geological Survey” |

*Methods for Benzidine: Chlorinated Organic Compounds, Pentachlorophenol and Pesticides in Water and Wastewater. September 1978.

# List of Acronyms and Abbreviations for Analytes, Units and Miscellaneous Terms

The following definitions apply to the acronyms and abbreviations for miscellaneous terms found in the 62-4.246(4) list.

| ADMI | American Dye Manufacturers’ Institute Color Value |
| --- | --- |
| BHC | Hexachlorocyclohexanes |
| COD | Chemical Oxygen Demand |
| Col/100 mL | Colonies per 100 mL of sample |
| D (2,4-) | 2,4-Dichlorophenoxyacetic acid |
| DDE | 1,1-Dichloro-2,2-di(p-chlorophenyl)ethylene |
| DDT | 4,4'-Dichlorodiphenyltrichloroethane |
| degree C | Degrees Celsius |
| DEP-SOP | DEP Standard Operating Procedures for Field Activities (DEP-SOP-001/01, January 2017) |
| DEP-SOP (NH3) | “Calculation of Un-Ionized Ammonia in Fresh Water” (DEP Chemistry Laboratory, 2/12/01) |
| FAC | Florida Administrative Code |
| mg/L | milligrams per liter |
| mL | milliliter(s) |
| MPN/100 mL | Most Probable Number per 100 mL |
| NTU | Nephelometric Turbidity Units |
| PCB | Polychlorinated biphenyl |
| pci/L | picocuries per liter |

| PCU | Platinum Cobalt Unit |
| --- | --- |
| pg/L | picograms per liter |
| SU | Standard Units (pH) |
| TCDD | Tetrachloro-dibenzo-p-dioxin |
| TDS | Total Dissolved Solids |
| TON | Threshold Odor Number |
| TP(2,4,5-) | 2-(2,4,5-Trichlorophenoxy)propionic acid (Silvex) |
| TSS | Total Suspended Solids |
| μg/L | micrograms per liter |
| μmhos/cm | micromhos per centimeter; equivalent to microsiemens per centimeter (μS/cm) |