## Final 5-31-19
Analyte Group | Analyte | Method | TechniqueID* | Target MDL | Target PQL | Units
Non-Pesticide Organics | 1,1,1-Trichloroethane | EPA 601 | GC-ELCD | 0.2 | 0.8 | ug/L
Non-Pesticide Organics | 1,1,1-Trichloroethane | SM 6200 C-2011 | GC-ELCD | 0.2 | 0.8 | ug/L
Non-Pesticide Organics | 1,1,1-Trichloroethane | EPA 1624B | GC-MS | 0.2 | 0.8 | ug/L
Non-Pesticide Organics | 1,1,1-Trichloroethane | EPA 624.1 | GC-MS | 0.2 | 0.8 | ug/L
Non-Pesticide Organics | 1,1,1-Trichloroethane | SM 6200 B-2011 | GC-MS | 0.2 | 0.8 | ug/L
Non-Pesticide Organics | 1,1,2,2-Tetrachloroethane | SM 6200 C-2011 | GC | 0.2 | 0.8 | ug/L
Non-Pesticide Organics | 1,1,2,2-Tetrachloroethane | EPA Method for Benzidine, pg. 130 | GC | 0.2 | 0.8 | ug/L
Non-Pesticide Organics | 1,1,2,2-Tetrachloroethane | EPA 601 | GC-ELCD | 0.2 | 0.8 | ug/L
Non-Pesticide Organics | 1,1,2,2-Tetrachloroethane | EPA 1624B | GC-MS | 0.5 | 0.8 | ug/L
Non-Pesticide Organics | 1,1,2,2-Tetrachloroethane | EPA 624.1 | GC-MS | 0.5 | 0.8 | ug/L
Non-Pesticide Organics | 1,1,2,2-Tetrachloroethane | SM 6200 B-2011 | GC-MS | 0.2 | 0.8 | ug/L
Non-Pesticide Organics | 1,1,2-Trichloroethane | EPA Method for Benzidine, pg. 130 | GC | 0.2 | 0.8 | ug/L
Non-Pesticide Organics | 1,1,2-Trichloroethane | EPA 601 | GC-ELCD | 0.2 | 0.8 | ug/L
Non-Pesticide Organics | 1,1,2-Trichloroethane | SM 6200 C-2011 | GC-ELCD | 0.2 | 0.8 | ug/L
Non-Pesticide Organics | 1,1,2-Trichloroethane | EPA 1624B | GC-MS | 0.2 | 0.8 | ug/L
Non-Pesticide Organics | 1,1,2-Trichloroethane | EPA 624.1 | GC-MS | 0.2 | 0.8 | ug/L
Non-Pesticide Organics | 1,1,2-Trichloroethane | SM 6200 B-2011 | GC-MS | 0.2 | 0.8 | ug/L
Non-Pesticide Organics | 1,1-Dichloroethane | SM 6200 C-2011 | GC | 0.2 | 0.8 | ug/L
Non-Pesticide Organics | 1,1-Dichloroethane | EPA 601 | GC-ELCD | 0.2 | 0.8 | ug/L
Non-Pesticide Organics | 1,1-Dichloroethane | EPA 1624B | GC-MS | 0.2 | 0.8 | ug/L
Non-Pesticide Organics | 1,1-Dichloroethane | EPA 624.1 | GC-MS | 0.2 | 0.8 | ug/L
Non-Pesticide Organics | 1,1-Dichloroethane | SM 6200 B-2011 | GC-MS | 0.2 | 0.8 | ug/L
Non-Pesticide Organics | 1,1-Dichloroethene | SM 6200 C-2011 | GC | 0.2 | 0.8 | ug/L
Non-Pesticide Organics | 1,1-Dichloroethene | EPA 601 | GC-ELCD | 0.2 | 0.8 | ug/L
Non-Pesticide Organics | 1,1-Dichloroethene | EPA 1624B | GC-MS | 0.5 | 2 | ug/L
Non-Pesticide Organics | 1,1-Dichloroethene | EPA 624.1 | GC-MS | 0.5 | 2 | ug/L
Non-Pesticide Organics | 1,1-Dichloroethene | SM 6200 B-2011 | GC-MS | 0.2 | 0.8 | ug/L
Non-Pesticide Organics | 1,2,4-Trichlorobenzene | EPA Method for Benzidine, p.130 | GC | 1 | 4 | ug/L
Non-Pesticide Organics | 1,2,4-Trichlorobenzene | EPA 612 | GC-ECD | 1 | 4 | ug/L
Non-Pesticide Organics | 1,2,4-Trichlorobenzene | EPA 1625B | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | 1,2,4-Trichlorobenzene | EPA 625.1 | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | 1,2,4-Trichlorobenzene | SM 6410 B-2000 | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | 1,2,4-Trichlorobenzene | USGS 0-3116-87 p.27 | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | 1,2-Dichlorobenzene | SM 6200 C-2011 | GC | 1 | 4 | ug/L
Non-Pesticide Organics | 1,2-Dichlorobenzene | EPA 601 | GC-ELCD | 0.2 | 0.8 | ug/L
Non-Pesticide Organics | 1,2-Dichlorobenzene | EPA 1625B | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | 1,2-Dichlorobenzene | EPA 624.1 | GC-MS | 0.5 | 2 | ug/L
Non-Pesticide Organics | 1,2-Dichlorobenzene | SM8 6200 B-2011 | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | 1,2-Dichlorobenzene | USGS 0-3116-87 p.27 | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | 1,2-Dichlorobenzene | EPA 602 | GC-PID | 0.2 | 0.8 | ug/L
Non-Pesticide Organics | 1,2-Dichloroethane | SM 6200 C-2011 | GC | 0.2 | 0.8 | ug/L
Non-Pesticide Organics | 1,2-Dichloroethane | EPA 601 | GC-ELCD | 0.2 | 0.8 | ug/L
Non-Pesticide Organics | 1,2-Dichloroethane | EPA 1624B | GC-MS | 0.2 | 0.8 | ug/L
Non-Pesticide Organics | 1,2-Dichloroethane | EPA 624.1 | GC-MS | 0.2 | 0.8 | ug/L
Non-Pesticide Organics | 1,2-Dichloroethane | SM 6200 B-2011 | GC-MS | 0.2 | 0.8 | ug/L
Non-Pesticide Organics | 1,2-Dichloropropane | SM 6200 C-2011 | GC | 0.2 | 0.8 | ug/L
Non-Pesticide Organics | 1,2-Dichloropropane | EPA 601 | GC-ELCD | 0.2 | 0.8 | ug/L
Non-Pesticide Organics | 1,2-Dichloropropane | EPA 1624B | GC-MS | 0.5 | 2 | ug/L
Non-Pesticide Organics | 1,2-Dichloropropane | EPA 624.1 | GC-MS | 0.5 | 2 | ug/L
Non-Pesticide Organics | 1,2-Dichloropropane | SM 6200 B-2011 | GC-MS | 0.2 | 0.8 | ug/L
Non-Pesticide Organics | 1,2-Diphenylhydrazine | EPA 1625 | GC-MS | 20 | 80 | ug/L
Non-Pesticide Organics | 1,2-Diphenylhydrazine | EPA 625.1 | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | 1,2-Diphenylhydrazine | EPA 8270 E | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | 1,3-Dichlorobenzene | SM 6200 C-2011 | GC | 0.2 | 0.8 | ug/L
Non-Pesticide Organics | 1,3-Dichlorobenzene | EPA 601 | GC-ELCD | 0.2 | 0.8 | ug/L
Non-Pesticide Organics | 1,3-Dichlorobenzene | EPA 1625B | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | 1,3-Dichlorobenzene | EPA 624.1 | GC-MS | 0.5 | 2 | ug/L
Non-Pesticide Organics | 1,3-Dichlorobenzene | SM 6200 B-2011 | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | 1,3-Dichlorobenzene | USGS 0-3116-87 p.27 | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | 1,3-Dichlorobenzene | EPA 602 | GC-PID | 0.2 | 0.8 | ug/L
Non-Pesticide Organics | 1,3-Dichloropropene- Total | SM 6200 C-2011 | GC | 0.2 | 0.8 | ug/L
Non-Pesticide Organics | 1,3-Dichloropropene- Total | EPA 601 | GC-ELCD | 0.2 | 0.8 | ug/L
Non-Pesticide Organics | 1,3-Dichloropropene- Total | SM 6200 B-2011 | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | 1,3-Dichloropropene- Total | EPA 1624B | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | 1,3-Dichloropropene- Total | EPA 624.1 | GC-MS | 0.2 | 0.8 | ug/L
Non-Pesticide Organics | 1,4-Dichlorobenzene | SM 6200 C-2011 | GC | 0.2 | 0.8 | ug/L
Non-Pesticide Organics | 1,4-Dichlorobenzene | EPA 601 | GC-ELCD | 0.2 | 0.8 | ug/L
Non-Pesticide Organics | 1,4-Dichlorobenzene | EPA 1625B | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | 1,4-Dichlorobenzene | EPA 624.1 | GC-MS | 0.5 | 2 | ug/L
Non-Pesticide Organics | 1,4-Dichlorobenzene | SM 6200 B-2011 | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | 1,4-Dichlorobenzene | USGS 0-3116-87 p.27 | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | 1,4-Dichlorobenzene | EPA 602 | GC-PID | 0.2 | 0.8 | ug/L
Non-Pesticide Organics | 1-Methylnaphthalene | EPA 610-GC | GC-FID | 1 | 4 | ug/L
Non-Pesticide Organics | 1-Methylnaphthalene | EPA 8100 | GC-FID | 1 | 4 | ug/L
Non-Pesticide Organics | 1-Methylnaphthalene | SM 6440 B-GC | GC-FID | 1 | 4 | ug/L
Non-Pesticide Organics | 1-Methylnaphthalene | EPA 1625 | GC-MS | 0.05 | 0.2 | ug/L
Non-Pesticide Organics | 1-Methylnaphthalene | EPA 625.1 | GC-MS | 0.05 | 0.2 | ug/L
Non-Pesticide Organics | 1-Methylnaphthalene | EPA 8270 E | GC-MS | 0.05 | 0.2 | ug/L
Non-Pesticide Organics | 1-Methylnaphthalene | SM 6410B | GC-MS | 0.05 | 0.2 | ug/L
Non-Pesticide Organics | 1-Methylnaphthalene | EPA 8310 | HPLC | 0.1 | 0.4 | ug/L
Non-Pesticide Organics | 1-Methylnaphthalene | EPA 610-HPLC | HPLC-UV-FLUOR | 0.1 | 0.4 | ug/L
Non-Pesticide Organics | 1-Methylnaphthalene | SM 6440 B-HPLC | HPLC-UV-FLUOR | 0.1 | 0.4 | ug/L
Non-Pesticide Organics | 2,3,4,5-Tetrachlorophenol | EPA 604 | GC-FID | 1 | 4 | ug/L
Non-Pesticide Organics | 2,3,4,5-Tetrachlorophenol | SM 6420 B | GC-FID | 1 | 4 | ug/L
Non-Pesticide Organics | 2,3,4,5-Tetrachlorophenol | EPA 8041 A | GC-FID or ECD | 1 | 4 | ug/L
Non-Pesticide Organics | 2,3,4,5-Tetrachlorophenol | EPA 1625 | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | 2,3,4,5-Tetrachlorophenol | EPA 8270 E | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | 2,3,4,5-Tetrachlorophenol | SM 6410 B | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | 2,3,4,6-Tetrachlorophenol | EPA 604 | GC-FID | 1 | 4 | ug/L
Non-Pesticide Organics | 2,3,4,6-Tetrachlorophenol | SM 6420 B | GC-FID | 1 | 4 | ug/L
Non-Pesticide Organics | 2,3,4,6-Tetrachlorophenol | EPA 8041 A | GC-FID or ECD | 1 | 4 | ug/L
Non-Pesticide Organics | 2,3,4,6-Tetrachlorophenol | EPA 1625 | GC-MS | 2 | 8 | ug/L
Non-Pesticide Organics | 2,3,4,6-Tetrachlorophenol | EPA 625.1 | GC-MS | 2 | 8 | ug/L
Non-Pesticide Organics | 2,3,4,6-Tetrachlorophenol | EPA 8270 E | GC-MS | 2 | 8 | ug/L
Non-Pesticide Organics | 2,3,4,6-Tetrachlorophenol | SM 6410 B | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | 2,3,4-Trichlorophenol | EPA 604 | GC-FID | 1 | 4 | ug/L
Non-Pesticide Organics | 2,3,4-Trichlorophenol | EPA 6420 B | GC-FID | 1 | 4 | ug/L
Non-Pesticide Organics | 2,3,4-Trichlorophenol | EPA 8041 A | GC-FID or ECD | 2 | 8 | ug/L
Non-Pesticide Organics | 2,3,4-Trichlorophenol | EPA 1625 | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | 2,3,4-Trichlorophenol | 8270 E | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | 2,3,4-Trichlorophenol | SM 6410 B | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | 2,3,5,6-Tetrachlorophenol | EPA 604 | GC-FID | 1 | 4 | ug/L
Non-Pesticide Organics | 2,3,5,6-Tetrachlorophenol | EPA 6420 B | GC-FID | 1 | 4 | ug/L
Non-Pesticide Organics | 2,3,5,6-Tetrachlorophenol | EPA 8041 A | GC-FID or ECD | 1 | 4 | ug/L
Non-Pesticide Organics | 2,3,5,6-Tetrachlorophenol | EPA 1625 | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | 2,3,5,6-Tetrachlorophenol | EPA 8270 E | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | 2,3,5,6-Tetrachlorophenol | SM 6410 B | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | 2,3,5-Trichlorophenol | EPA 604 | GC-FID | 1 | 4 | ug/L
Non-Pesticide Organics | 2,3,5-Trichlorophenol | SM 6420 B | GC-FID | 1 | 4 | ug/L
Non-Pesticide Organics | 2,3,5-Trichlorophenol | EPA 8041 A | GC-FID or ECD | 2 | 8 | ug/L
Non-Pesticide Organics | 2,3,5-Trichlorophenol | EPA 1625 | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | 2,3,5-Trichlorophenol | 8270 E | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | 2,3,5-Trichlorophenol | SM 6410 B | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | 2,3,6-Trichlorophenol | EPA 604 | GC-FID | 1 | 4 | ug/L
Non-Pesticide Organics | 2,3,6-Trichlorophenol | SM6420 B | GC-FID | 1 | 4 | ug/L
Non-Pesticide Organics | 2,3,6-Trichlorophenol | EPA 8041 A | GC-FID or ECD | 2 | 8 | ug/L
Non-Pesticide Organics | 2,3,6-Trichlorophenol | EPA 1625 | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | 2,3,6-Trichlorophenol | 8270 E | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | 2,3,6-Trichlorophenol | EPA 1625 | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | 2,3,6-Trichlorophenol | SM 6410 B | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | 2,3,6-Trichlorophenol | EPA 625.1 | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | 2,3,7,8-Tetrachloro-dibenzo-p-dioxin | EPA 613 | GC-MS | 0.002 | 0.008 | ug/L
Non-Pesticide Organics | 2,3,7,8-Tetrachloro-dibenzo-p-dioxin | EPA 625.1 | GC-MS | 0.002 | 0.008 | ug/L
Non-Pesticide Organics | 2,3,7,8-Tetrachloro-dibenzo-p-dioxin | EPA 1613B | GC-MS | 4.4 | 18 | pg/L
Non-Pesticide Organics | 2,3-Dichlorophenol | EPA 604 | GC-FID | 1 | 4 | ug/L
Non-Pesticide Organics | 2,3-Dichlorophenol | EPA 6420 B | GC-FID | 1 | 4 | ug/L
Non-Pesticide Organics | 2,3-Dichlorophenol | EPA 8041 A | GC-FID or ECD | 1 | 4 | ug/L
Non-Pesticide Organics | 2,3-Dichlorophenol | EPA 1625 | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | 2,3-Dichlorophenol | EPA 8270 E | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | 2,3-Dichlorophenol | SM 6410 B | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | 2,4 Dimethylphenol | SM 6420 B | GC-ECD | 0.63 | 2.5 | ug/L
Non-Pesticide Organics | 2,4 Dimethylphenol | SM 6420 B-2000 | GC-FID | 0.32 | 1.3 | ug/L
Non-Pesticide Organics | 2,4 Dimethylphenol | SM 6410 B-2000 | GC-MS | 2.7 | 11 | ug/L
Non-Pesticide Organics | 2,4 Dimethylphenol | EPA 604 | GC-MS | 2.7 | 11 | ug/L
Non-Pesticide Organics | 2,4 Dimethylphenol | EPA 625.1 | GC-MS | 2.7 | 11 | ug/L
Non-Pesticide Organics | 2,4 Dimethylphenol | EPA 1625B | GC-MS | 2.7 | 11 | ug/L
Non-Pesticide Organics | 2,4 Dimethylphenol | USGS O-3116-87 p.28 | GC-MS | 2.7 | 11 | ug/L
Pesticide | 2,4,5-TP | EPA, 1978, p. 115 | GC | 0.2 | 0.6 | ug/L
Pesticide | 2,4,5-TP | USGS O-3105-83 | GC | 0.2 | 0.6 | ug/L
Pesticide | 2,4,5-TP | SM 6640 B-2006 | GC-ECD | 0.2 | 0.6 | ug/L
Pesticide | 2,4,5-TP | EPA 615 | GC-ECD | 0.2 | 0.6 | ug/L
Non-Pesticide Organics | 2,4,5-Trichlorophenol | EPA 604 | GC-FID or ECD | 1 | 4 | ug/L
Non-Pesticide Organics | 2,4,5-Trichlorophenol | EPA 8041 A | GC-FID or ECD | 2 | 8 | ug/L
Non-Pesticide Organics | 2,4,5-Trichlorophenol | SM 6420 B | GC-FID or ECD | 1 | 4 | ug/L
Non-Pesticide Organics | 2,4,5-Trichlorophenol | EPA 1625 | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | 2,4,5-Trichlorophenol | EPA 8270 E | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | 2,4,5-Trichlorophenol | SM 6410 B | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | 2,4,5-Trichlorophenol | EPA 625.1 | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | 2,4,6-Trichlorophenol | EPA 604 | GC-FID or ECD | 2 | 8 | ug/L
Non-Pesticide Organics | 2,4,6-Trichlorophenol | EPA 8041 A | GC-FID or ECD | 2 | 8 | ug/L
Non-Pesticide Organics | 2,4,6-Trichlorophenol | SM 6420 B | GC-FID or ECD | 2 | 8 | ug/L
Non-Pesticide Organics | 2,4,6-Trichlorophenol | EPA 1625B | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | 2,4,6-Trichlorophenol | EPA 625.1 | GC-MS | 2.7 | 11 | ug/L
Non-Pesticide Organics | 2,4,6-Trichlorophenol | EPA 8270 E | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | 2,4,6-Trichlorophenol | SM 6410 B | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | 2,4,6-Trichlorophenol | USGS O-3116-87 p.27 | GC-MS | 1 | 4 | ug/L
Pesticide | 2,4-D | EPA 615 | GC | 0.2 | 0.6 | ug/L
Pesticide | 2,4-D | EPA Method for Benzidine p.115 | GC | 0.2 | 0.6 | ug/L
Pesticide | 2,4-D | USGS O-3105-83 | GC | 0.2 | 0.6 | ug/L
Pesticide | 2,4-D | SM 6640 B-2006 | GC-ECD | 0.2 | 0.6 | ug/L
Pesticide | 2,4-D | USGS O-2060-01 | HPLC | 0.2 | 0.6 | ug/L
Non-Pesticide Organics | 2,4-Dichlorophenol | SM 6420 B | GC-ECD | 0.68 | 2.7 | ug/L
Non-Pesticide Organics | 2,4-Dichlorophenol | EPA 604 | GC-FID | 1 | 4 | ug/L
Non-Pesticide Organics | 2,4-Dichlorophenol | SM 6420 B-2000 | GC-FID | 1 | 4 | ug/L
Non-Pesticide Organics | 2,4-Dichlorophenol | EPA 8041 A | GC-FID or ECD | 1 | 4 | ug/L
Non-Pesticide Organics | 2,4-Dichlorophenol | EPA 1625B | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | 2,4-Dichlorophenol | EPA 625.1 | GC-MS | 2.7 | 11 | ug/L
Non-Pesticide Organics | 2,4-Dichlorophenol | EPA 8270 E | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | 2,4-Dichlorophenol | SM 6410 B-2000 | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | 2,4-Dichlorophenol | USGS O-3116-87 p.27 | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | 2,4-Dinitrophenol | EPA 604 | GC-FID | 15 | 60 | ug/L
Non-Pesticide Organics | 2,4-Dinitrophenol | SM 6420B-2000 | GC-FID | 15 | 60 | ug/L
Non-Pesticide Organics | 2,4-Dinitrophenol | EPA 8041 A | GC-FID or ECD | 15 | 60 | ug/L
Non-Pesticide Organics | 2,4-Dinitrophenol | EPA 1625B | GC-MS | 15 | 60 | ug/L
Non-Pesticide Organics | 2,4-Dinitrophenol | EPA 625.1 | GC-MS | 42 | 170 | ug/L
Non-Pesticide Organics | 2,4-Dinitrophenol | EPA 8270 E | GC-MS | 15 | 60 | ug/L
Non-Pesticide Organics | 2,4-Dinitrophenol | SM 6410B-2000 | GC-MS | 15 | 60 | ug/L
Non-Pesticide Organics | 2,4-Dinitrophenol | USGS O-3116-87 p.28 | GC-MS | 15 | 60 | ug/L
Non-Pesticide Organics | 2,4-Dinitrotoluene | EPA 609 | GC | 1 | 4 | ug/L
Non-Pesticide Organics | 2,4-Dinitrotoluene | EPA 1625B | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | 2,4-Dinitrotoluene | EPA 625.1 | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | 2,4-Dinitrotoluene | SM 6410B-2000 | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | 2,4-Dinitrotoluene | USGS O-3116-87 p.28 | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | 2,5-Dichlorophenol | EPA 604 | GC-FID | 1 | 4 | ug/L
Non-Pesticide Organics | 2,5-Dichlorophenol | EPA 8041 A | GC-FID or ECD | 1 | 4 | ug/L
Non-Pesticide Organics | 2,5-Dichlorophenol | EPA 8270 E | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | 2,5-Dichlorophenol | EPA 1625 | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | 2,5-Dichlorophenol | SM 6410 B | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | 2,5-Dichlorophenol | SM 6420 B | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | 2,6-Dichlorophenol | EPA 604 | GC-FID | 1 | 4 | ug/L
Non-Pesticide Organics | 2,6-Dichlorophenol | SM 6420 B | GC-FID | 1 | 4 | ug/L
Non-Pesticide Organics | 2,6-Dichlorophenol | EPA 8041 A | GC-FID or ECD | 1 | 4 | ug/L
Non-Pesticide Organics | 2,6-Dichlorophenol | EPA 1625 | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | 2,6-Dichlorophenol | EPA 8270 E | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | 2,6-Dichlorophenol | SM 6410 B | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | 2,6-Dichlorophenol | EPA 625.1 | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | 2,6-Dinitrotoluene | EPA 609 | GC | 2 | 8 | ug/L
Non-Pesticide Organics | 2,6-Dinitrotoluene | EPA 1625B | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | 2,6-Dinitrotoluene | EPA 625.1 | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | 2,6-Dinitrotoluene | SM 6410 B-2000 | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | 2,6-Dinitrotoluene | USGS O-3116-87 p.28 | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | 2-Chloro-5-methylphenol | EPA 604 | GC-FID | 1 | 4 | ug/L
Non-Pesticide Organics | 2-Chloro-5-methylphenol | SM 6420 B | GC-FID | 1 | 4 | ug/L
Non-Pesticide Organics | 2-Chloro-5-methylphenol | EPA 8041 A | GC-FID or ECD | 1 | 4 | ug/L
Non-Pesticide Organics | 2-Chloro-5-methylphenol | EPA 1625 | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | 2-Chloro-5-methylphenol | EPA 8270 E | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | 2-Chloro-5-methylphenol | SM 6410 B | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | 2-Chloroethylvinyl ether | EPA 601 | GC-ELCD | 0.5 | 2 | ug/L
Non-Pesticide Organics | 2-Chloroethylvinyl ether | EPA 1624B | GC-MS | 0.5 | 2 | ug/L
Non-Pesticide Organics | 2-Chloroethylvinyl ether | EPA 624.1 | GC-MS | 0.5 | 2 | ug/L
Non-Pesticide Organics | 2-Chloronaphthalene | EPA 612 | GC-ECD | 1 | 4 | ug/L
Non-Pesticide Organics | 2-Chloronaphthalene | EPA 1625B | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | 2-Chloronaphthalene | EPA 625.1 | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | 2-Chloronaphthalene | SM 6410 B-2000 | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | 2-Chloronaphthalene | USGS O-3116-87 p.27 | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | 2-Chlorophenol | EPA 604 | GC-FID | 1 | 4 | ug/L
Non-Pesticide Organics | 2-Chlorophenol | SM 6420 B-2000 | GC-FID | 1 | 4 | ug/L
Non-Pesticide Organics | 2-Chlorophenol | EPA 8041 A | GC-FID or ECD | 1 | 4 | ug/L
Non-Pesticide Organics | 2-Chlorophenol | EPA 1625B | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | 2-Chlorophenol | EPA 625.1 | GC-MS | 3.3 | 13 | ug/L
Non-Pesticide Organics | 2-Chlorophenol | EPA 8270 E | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | 2-Chlorophenol | SM 6410 B-2000 | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | 2-Chlorophenol | USGS O-3116-87 p.27 | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | 2-Methyl-4,6-dinitrophenol | EPA 604 | GC-FID | 3 | 12 | ug/L
Non-Pesticide Organics | 2-Methyl-4,6-dinitrophenol | SM 6420B-2000 | GC-FID | 3 | 12 | ug/L
Non-Pesticide Organics | 2-Methyl-4,6-dinitrophenol | EPA 625.1 | GC-MS | 24 | 96 | ug/L
Non-Pesticide Organics | 2-Methyl-4,6-dinitrophenol | SM 6410B-2000 | GC-MS | 3 | 12 | ug/L
Non-Pesticide Organics | 2-Methyl-4,6-dinitrophenol | EPA 1625B | GC-MS | 3 | 12 | ug/L
Non-Pesticide Organics | 2-Methyl-4,6-dinitrophenol | USGS O-3116-87 p.27 | GC-MS | 3 | 12 | ug/L
Non-Pesticide Organics | 2-Methylnaphthalene | EPA 610-GC | GC-FID | 1 | 4 | ug/L
Non-Pesticide Organics | 2-Methylnaphthalene | EPA 8100 | GC-FID | 1 | 4 | ug/L
Non-Pesticide Organics | 2-Methylnaphthalene | EPA 1625 | GC-MS | 0.05 | 0.2 | ug/L
Non-Pesticide Organics | 2-Methylnaphthalene | EPA 625.1 | GC-MS | 0.05 | 0.2 | ug/L
Non-Pesticide Organics | 2-Methylnaphthalene | EPA 8270 E | GC-MS | 0.05 | 0.2 | ug/L
Non-Pesticide Organics | 2-Methylnaphthalene | SM 6410 B | GC-MS | 0.05 | 0.2 | ug/L
Non-Pesticide Organics | 2-Methylnaphthalene | EPA 8310 | HPLC-UV-FLUOR | 0.1 | 0.4 | ug/L
Non-Pesticide Organics | 2-Methylnaphthalene | EPA 610-HPLC | HPLC-UV-FLUOR | 0.1 | 0.4 | ug/L
Non-Pesticide Organics | 2-Methylnaphthalene | SM 6440 B-HPLC | HPLC-UV-FLUOR | 0.1 | 0.4 | ug/L
Non-Pesticide Organics | 2-Nitrophenol | EPA 604 | GC-FID | 1 | 4 | ug/L
Non-Pesticide Organics | 2-Nitrophenol | SM 6420B-2000 | GC-FID | 1 | 4 | ug/L
Non-Pesticide Organics | 2-Nitrophenol | EPA 8041 A | GC-FID or ECD | 15 | 60 | ug/L
Non-Pesticide Organics | 2-Nitrophenol | EPA 1625B | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | 2-Nitrophenol | EPA 625.1 | GC-MS | 3.6 | 14 | ug/L
Non-Pesticide Organics | 2-Nitrophenol | SM 6410B-2000 | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | 2-Nitrophenol | USGS O-3116-87 p.27 | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | 3,3'-Dichlorobenzidine | EPA 1625B | GC-MS | 40 | 160 | ug/L
Non-Pesticide Organics | 3,3'-Dichlorobenzidine | EPA 625.1 | GC-MS | 40 | 160 | ug/L
Non-Pesticide Organics | 3,3'-Dichlorobenzidine | SM 6410 B-2000 | GC-MS | 40 | 160 | ug/L
Non-Pesticide Organics | 3,3'-Dichlorobenzidine | EPA 605 | HPLC | 40 | 160 | ug/L
Non-Pesticide Organics | 3,4-Dichlorophenol | EPA 604 | GC-FID | 1 | 4 | ug/L
Non-Pesticide Organics | 3,4-Dichlorophenol | SM 6420 B | GC-FID | 1 | 4 | ug/L
Non-Pesticide Organics | 3,4-Dichlorophenol | EPA 8041 A | GC-FID or ECD | 1 | 4 | ug/L
Non-Pesticide Organics | 3,4-Dichlorophenol | EPA 1625 | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | 3,4-Dichlorophenol | EPA 8270 E | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | 3,4-Dichlorophenol | SM 6410 B | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | 3,5-Dichlorophenol | EPA 604 | GC-FID | 1 | 4 | ug/L
Non-Pesticide Organics | 3,5-Dichlorophenol | SM 6420 B | GC-FID | 1 | 4 | ug/L
Non-Pesticide Organics | 3,5-Dichlorophenol | EPA 8041 A | GC-FID or ECD | 1 | 4 | ug/L
Non-Pesticide Organics | 3,5-Dichlorophenol | EPA 1625 | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | 3,5-Dichlorophenol | EPA 8270 E | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | 3,5-Dichlorophenol | SM 6410 B | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | 3-Chlorophenol | EPA 604 | GC-FID | 1 | 4 | ug/L
Non-Pesticide Organics | 3-Chlorophenol | SM 6420 B | GC-FID | 1 | 4 | ug/L
Non-Pesticide Organics | 3-Chlorophenol | EPA 8041 A | GC-FID or ECD | 1 | 4 | ug/L
Non-Pesticide Organics | 3-Chlorophenol | EPA 1625 | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | 3-Chlorophenol | EPA 8270 E | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | 3-Chlorophenol | SM 6410 B | GC-MS | 1 | 4 | ug/L
Pesticide | 4,4'-DDD | ASTM D5812-96 (02) | GC | 0.02 | 0.08 | ug/L
Pesticide | 4,4'-DDD | EPA 617 | GC | 0.02 | 0.08 | ug/L
Pesticide | 4,4'-DDD | EPA, 1978, p.7 | GC | 0.02 | 0.08 | ug/L
Pesticide | 4,4'-DDD | USGS O-3105-83 | GC | 0.02 | 0.08 | ug/L
Pesticide | 4,4'-DDD | 3M Corp 3M0222 (10/28/94) | GC | 0.02 | 0.08 | ug/L
Pesticide | 4,4'-DDD | ASTM D3086-90 | GC-ECD | 0.02 | 0.08 | ug/L
Pesticide | 4,4'-DDD | EPA 608.3 | GC-ECD | 0.02 | 0.08 | ug/L
Pesticide | 4,4'-DDD | SM 6630 B-2007 | GC-ECD | 0.02 | 0.08 | ug/L
Pesticide | 4,4'-DDD | SM 6630 C-2007 | GC-ECD | 0.02 | 0.08 | ug/L
Pesticide | 4,4'-DDD | EPA 625.1 | GC-MS | 2.8 | 11.2 | ug/L
Pesticide | 4,4'-DDD | SM 6410 B-2000 | GC-MS | 2.8 | 11.2 | ug/L
Pesticide | 4,4'-DDE | ASTM D5812-96 (02) | GC | 0.02 | 0.08 | ug/L
Pesticide | 4,4'-DDE | EPA 617 | GC | 0.02 | 0.08 | ug/L
Pesticide | 4,4'-DDE | EPA, 1978, p.7 | GC | 0.02 | 0.08 | ug/L
Pesticide | 4,4'-DDE | USGS O-3104-83 | GC | 0.02 | 0.08 | ug/L
Pesticide | 4,4'-DDE | 3M Corp 3M0222 (10/28/94) | GC | 0.02 | 0.08 | ug/L
Pesticide | 4,4'-DDE | USGS O-1126-95 | GC | 0.02 | 0.08 | ug/L
Pesticide | 4,4'-DDE | ASTM D3086-90 | GC-ECD | 0.02 | 0.08 | ug/L
Pesticide | 4,4'-DDE | EPA 608.3 | GC-ECD | 0.02 | 0.08 | ug/L
Pesticide | 4,4'-DDE | SM 6630 B-2007 | GC-ECD | 0.02 | 0.08 | ug/L
Pesticide | 4,4'-DDE | SM 6630 C-2007 | GC-ECD | 0.02 | 0.08 | ug/L
Pesticide | 4,4'-DDE | EPA 625.1 | GC-MS | 2 | 8 | ug/L
Pesticide | 4,4'-DDE | SM 6410 B-2000 | GC-MS | 2 | 8 | ug/L
Pesticide | 4,4'-DDT | EPA, 1978, p.7 | GC | 0.02 | 0.08 | ug/L
Pesticide | 4,4'-DDT | USGS O-3104-83 | GC | 0.02 | 0.08 | ug/L
Pesticide | 4,4'-DDT | 3M Corp 3M0222 (10/28/94) | GC | 0.02 | 0.08 | ug/L
Pesticide | 4,4'-DDT | ASTM D3086-90 | GC-ECD | 0.02 | 0.08 | ug/L
Pesticide | 4,4'-DDT | ASTM D5812-96 (02) | GC-ECD | 0.02 | 0.08 | ug/L
Pesticide | 4,4'-DDT | EPA 608.3 | GC-ECD | 0.02 | 0.08 | ug/L
Pesticide | 4,4'-DDT | SM 6630 B-2007 | GC-ECD | 0.02 | 0.08 | ug/L
Pesticide | 4,4'-DDT | SM 6630 C-2007 | GC-ECD | 0.02 | 0.08 | ug/L
Pesticide | 4,4'-DDT | EPA 625.1 | GC-MS | 2 | 8 | ug/L
Pesticide | 4,4'-DDT | SM 6410 B-2000 | GC-MS | 2 | 8 | ug/L
Non-Pesticide Organics | 4-Bromophenyl phenyl ether | EPA 611 | GC-ELCD-MCTD | 1 | 4 | ug/L
Non-Pesticide Organics | 4-Bromophenyl phenyl ether | EPA 1625B | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | 4-Bromophenyl phenyl ether | EPA 625.1 | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | 4-Bromophenyl phenyl ether | SM 6410 B-2000 | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | 4-Bromophenyl phenyl ether | USGS O-3116-87 p.27 | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | 4-Chloro-2-methylphenol | EPA 604 | GC-FID | 1 | 4 | ug/L
Non-Pesticide Organics | 4-Chloro-2-methylphenol | SM 6420 B | GC-FID | 1 | 4 | ug/L
Non-Pesticide Organics | 4-Chloro-2-methylphenol | EPA 8041 A | GC-FID or ECD | 1 | 4 | ug/L
Non-Pesticide Organics | 4-Chloro-2-methylphenol | EPA 1625 | GC-MNS | 1 | 4 | ug/L
Non-Pesticide Organics | 4-Chloro-2-methylphenol | EPA 8270 E | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | 4-Chloro-2-methylphenol | SM 6410 B | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | 4-Chloro-3-methylphenol | EPA 604 | GC-FID | 1 | 4 | ug/L
Non-Pesticide Organics | 4-Chloro-3-methylphenol | SM 6420 B | GC-FID | 1 | 4 | ug/L
Non-Pesticide Organics | 4-Chloro-3-methylphenol | EPA 8041 A | GC-FID or ECD | 1 | 4 | ug/L
Non-Pesticide Organics | 4-Chloro-3-methylphenol | EPA 1625 | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | 4-Chloro-3-methylphenol | EPA 625.1 | GC-MS | 3.0 | 12 | ug/L
Non-Pesticide Organics | 4-Chloro-3-methylphenol | EPA 8270 E | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | 4-Chloro-3-methylphenol | SM 6410 B | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | 4-Chlorophenol | EPA 604 | GC-FID | 1 | 4 | ug/L
Non-Pesticide Organics | 4-Chlorophenol | SM 6420 B | GC-FID | 1 | 4 | ug/L
Non-Pesticide Organics | 4-Chlorophenol | EPA 8041 A | GC-FID or ECD | 1 | 4 | ug/L
Non-Pesticide Organics | 4-Chlorophenol | EPA 1625 | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | 4-Chlorophenol | EPA 8270 E | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | 4-Chlorophenol | SM 6410 B | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | 4-Chlorophenyl phenyl ether | EPA 611 | GC-ELCD-MCTD | 2 | 8 | ug/L
Non-Pesticide Organics | 4-Chlorophenyl phenyl ether | EPA 1625B | GC-MS | 2 | 8 | ug/L
Non-Pesticide Organics | 4-Chlorophenyl phenyl ether | EPA 625.1 | GC-MS | 2 | 8 | ug/L
Non-Pesticide Organics | 4-Chlorophenyl phenyl ether | SM 6410 B-2000 | GC-MS | 2 | 8 | ug/L
Non-Pesticide Organics | 4-Chlorophenyl phenyl ether | USGS O-3116-87 p.27 | GC-MS | 2 | 8 | ug/L
Non-Pesticide Organics | 4-Nitrophenol | EPA 604 | GC-FID | 15 | 60 | ug/L
Non-Pesticide Organics | 4-Nitrophenol | SM 6420B-2000 | GC-FID | 15 | 60 | ug/L
Non-Pesticide Organics | 4-Nitrophenol | EPA 8041 A | GC-FID or ECD | 15 | 60 | ug/L
Non-Pesticide Organics | 4-Nitrophenol | EPA 1625B | GC-MS | 15 | 60 | ug/L
Non-Pesticide Organics | 4-Nitrophenol | EPA 625.1 | GC-MS | 2.4 | 10 | ug/L
Non-Pesticide Organics | 4-Nitrophenol | EPA 8270 E | GC-MS | 15 | 60 | ug/L
Non-Pesticide Organics | 4-Nitrophenol | SM 6410B-20000 | GC-MS | 15 | 60 | ug/L
Non-Pesticide Organics | 4-Nitrophenol | USGS O-3116-87 p.27 | GC-MS | 15 | 60 | ug/L
Non-Pesticide Organics | 4-Nonylphenol (Nonylphenol) | ASTM D7065-11 | GC-MS | 0.9 | 3.6 | ug/L
Pesticide | a-BHC | EPA 617 | GC | 0.01 | 0.04 | ug/L
Pesticide | a-BHC | ASTM D5812-96 (02) | GC | 0.01 | 0.04 | ug/L
Pesticide | a-BHC | EPA, 1978, p.7 | GC | 0.01 | 0.04 | ug/L
Pesticide | a-BHC | 3M Corp 3M0222 (10/28/94) | GC | 0.01 | 0.04 | ug/L
Pesticide | a-BHC | ASTM D3086-90 | GC-ECD | 0.01 | 0.04 | ug/L
Pesticide | a-BHC | EPA 608.3 | GC-ECD | 0.01 | 0.04 | ug/L
Pesticide | a-BHC | SM 6630 B-2007 | GC-ECD | 0.01 | 0.04 | ug/L
Pesticide | a-BHC | SM 6630 C-2007 | GC-ECD | 0.01 | 0.04 | ug/L
Pesticide | a-BHC | EPA 625.1 | GC-MS | 2 | 8 | ug/L
Pesticide | a-BHC | SM 6410 B | GC-MS | 2 | 8 | ug/L
Pesticide | a-BHC | USGS O-1126-95 | GC-MS | 2 | 8 | ug/L
Non-Pesticide Organics | Acenaphthene | EPA 610-GC | GC-FID | 1 | 4 | ug/L
Non-Pesticide Organics | Acenaphthene | EPA 1625B | GC-MS | 0.05 | 0.2 | ug/L
Non-Pesticide Organics | Acenaphthene | EPA 625.1 | GC-MS | 0.05 | 0.2 | ug/L
Non-Pesticide Organics | Acenaphthene | USGS O-3116-87 p.27 | GC-MS | 0.05 | 0.2 | ug/L
Non-Pesticide Organics | Acenaphthene | SM 6410 B-2000 | GC-MS | 0.05 | 0.2 | ug/L
Non-Pesticide Organics | Acenaphthene | ASTM D4657-92 | HPLC-UV-FLUOR | 0.3 | 1.2 | ug/L
Non-Pesticide Organics | Acenaphthene | EPA 610-HPLC | HPLC-UV-FLUOR | 0.3 | 1.2 | ug/L
Non-Pesticide Organics | Acenaphthene | SM 6440 B-2005-HPLC | HPLC-UV-FLUOR | 0.3 | 1.2 | ug/L
Non-Pesticide Organics | Acenaphthylene | EPA 610-GC | GC-FID | 1 | 4 | ug/L
Non-Pesticide Organics | Acenaphthylene | EPA 1625B | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | Acenaphthylene | EPA 625.1 | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | Acenaphthylene | SM 6410 B | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | Acenaphthylene | USGS O-3116-87 p.27 | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | Acenaphthylene | ASTM D4657-92 | HPLC-UV-FLUOR | 0.3 | 1.2 | ug/L
Non-Pesticide Organics | Acenaphthylene | EPA 610-HPLC | HPLC-UV-FLUOR | 0.3 | 1.2 | ug/L
Non-Pesticide Organics | Acenaphthylene | SM 6440 B-2005-HPLC | HPLC-UV-FLUOR | 0.3 | 1.2 | ug/L
Inorganic | Acid Mine Drainage | EPA 1627 | LEACH | N/A* | N/A* | mg/L
Non-Pesticide Organics | Acrolein | EPA 603 | GC-FID | 5 | 20 | ug/L
Non-Pesticide Organics | Acrolein | EPA 1624B | GC-MS | 10 | 40 | ug/L
Non-Pesticide Organics | Acrolein | EPA 624.1 | GC-MS | 10 | 40 | ug/L
Non-Pesticide Organics | Acrylonitrile | EPA 603 | GC-FID | 5 | 20 | ug/L
Non-Pesticide Organics | Acrylonitrile | EPA 1624B | GC-MS | 4 | 16 | ug/L
Non-Pesticide Organics | Acrylonitrile | EPA 624.1 | GC-MS | 4 | 16 | ug/L
Pesticide | Aldrin | EPA, 1978, p.7 | GC | 0.01 | 0.04 | ug/L
Pesticide | Aldrin | USGS 0-3104-83 | GC | 0.01 | 0.04 | ug/L
Pesticide | Aldrin | 3M Corp 3M0222 (10/28/94) | GC | 0.01 | 0.04 | ug/L
Pesticide | Aldrin | ASTM D3086-90 | GC-ECD | 0.01 | 0.04 | ug/L
Pesticide | Aldrin | D5812-96 (02) | GC-ECD | 0.01 | 0.04 | ug/L
Pesticide | Aldrin | EPA 608.3 | GC-ECD | 0.01 | 0.04 | ug/L
Pesticide | Aldrin | EPA 617 | GC-ECD | 0.01 | 0.04 | ug/L
Pesticide | Aldrin | SM 6630 B-2007 | GC-ECD | 0.01 | 0.04 | ug/L
Pesticide | Aldrin | SM 6630 C-2007 | GC-ECD | 0.01 | 0.04 | ug/L
Pesticide | Aldrin | EPA 625.1 | GC-MS | 2 | 8 | ug/L
Pesticide | Aldrin | SM 6410 B-2000 | GC-MS | 2 | 8 | ug/L
Inorganic | Alkalinity | EPA 310.2 (Rev. 1974) | COLOR-Auto | 10 | 40 | mg/L
Inorganic | Alkalinity | USGS I-2030-85 | COLOR-Auto | 10 | 40 | mg/L
Inorganic | Alkalinity | SM 2320B-11 | TITR* | 20 | 20 | mg/L
Inorganic | Alkalinity | ASTM D1067-11 | TITR* | 20 | 20 | mg/L
Inorganic | Alkalinity | AOAC 973.43 | TITR* | 20 | 20 | mg/L
Inorganic | Alkalinity | USGS I-1030-85 | TITR* | 20 | 20 | mg/L
Radiologic | Alpha - Gross (Total) | EPA 900.0 (Alpha) | A-SCINT* or GPC* | 3 | 3 | pCi/L
Radiologic | Alpha - Gross (Total) | SM 7110 B (Alpha) | A-SCINT* or GPC* | 3 | 3 | pCi/L
Radiologic | Alpha - Gross (Total) | SM 7110 B-00 (Alpha) | A-SCINT* or GPC* | 3 | 3 | pCi/L
Radiologic | Alpha - Gross (Total) | USGS 76-177 (1976), pp. 75 & 78 | A-SCINT* or GPC* | 3 | 3 | pCi/L
Radiologic | Alpha - Gross (Total) | ASTM D1943-90 | A-SCINT* or GPC* | 3 | 3 | pCi/L
Radiologic | Alpha - Gross (Total) | ASTM D1943-96 | A-SCINT* or GPC* | 3 | 3 | pCi/L
Radiologic | Alpha - Gross (Total) | EPA 00-02 | CO-PRECIP* | 3 | 3 | pCi/L
Radiologic | Alpha - Gross (Total) | SM 7110 C | CO-PRECIP* | 3 | 3 | pCi/L
Inorganic | Aluminum | SM 3500-Al B-11 | COLOR | 6 | 20 | ug/L
Inorganic | Aluminum | ASTM D4190­08 | DCP | 50 | 200 | mg/L
Inorganic | Aluminum | AES0029, Rev. 1991 Thermo Jarrell Ash Corp. | DCP | 50 | 200 | mg/L
Inorganic | Aluminum | SM 3111 D-11 | FLAAS | 100 | 400 | ug/L
Inorganic | Aluminum | SM 3111 E-11 | FLAAS | 100 | 400 | ug/L
Inorganic | Aluminum | USGS I­3051­852 | FLAAS | 100 | 400 | ug/L
Inorganic | Aluminum | SM 3113 B-10 | GFAAS | 3 | 12 | ug/L
Inorganic | Aluminum | ASTM D1976-12 | ICP-AES | 2.2 | 8.8 | ug/L
Inorganic | Aluminum | EPA 200.7, Rev. 4.4 (1994) | ICP-AES | 20 | 80 | ug/L
Inorganic | Aluminum | EPA 200.5, Rev 4.2 (2003) | ICP-AES | 2.2 | 8.8 | ug/L
Inorganic | Aluminum | USGS I­4471­9750 | ICP-AES | 1 | 4 | ug/L
Inorganic | Aluminum | SM 3120 B-11 | ICP-ES | 2.2 | 8.8 | ug/L
Inorganic | Aluminum | AOAC 993.14 3 | ICP-MS | 1 | 4 | ug/L
Inorganic | Aluminum | ASTM D5673­10 | ICP-MS | 1 | 4 | ug/L
Inorganic | Aluminum | EPA 200.8, Rev. 5.4 (1994) | ICP-MS | 1 | 4 | ug/L
Inorganic | Aluminum | SM 3125 B-11 | ICP-MS | 1 | 4 | ug/L
Inorganic | Aluminum | EPA 200.9, Rev. 2.2 (1994) | STGFAA | 7.8 | 31.2 | ug/L
Inorganic | Ammonia | Timberline Ammonia-001 (2011) | DIFFUSION-COND-Auto | 0.002 | 0.008 | mg/L
Inorganic | Ammonia (N) | EPA 350.1, Rev. 2.0 (1993) | COLOR-Auto | 10 | 40 | ug/L
Inorganic | Ammonia (N) | USGS I-4523-85 | COLOR-Auto | 10 | 40 | ug/L
Inorganic | Ammonia (N) | SM 4500-NH3 G-11 and 4500-NH3 H -11 | COLOR-Auto | 0.002 | 0.008 | mg/L
Inorganic | Ammonia (N) | ASTM D6919-09 | IC | 0.4 | 1.6 | mg/L
Inorganic | Ammonia (N) | ASTM D1426-08 (B) | ISE | 50 | 200 | ug/L
Inorganic | Ammonia (N) | SM 4500-NH3 D or E -11 | ISE | 0.02 | 0.08 | mg/L
Inorganic | Ammonia (N) | 4500­NH3 F-11 | ISE | 0.02 | 0.08 | mg/L
Inorganic | Ammonia (N) | Bran Luebbe 379-75 WE | ISE-Auto | 50 | 200 | ug/L
Inorganic | Ammonia (N) | USGS I-3520-85 | NESSLER | 50 | 200 | ug/L
Inorganic | Ammonia (N) | AOAC 973.49 (1998) | NESSLER | 50 | 200 | ug/L
Inorganic | Ammonia (N) | ASTM D1426-08 (A) | NESSLER | 50 | 200 | ug/L
Inorganic | Ammonia (N) | SM 4500-NH3 C-11 | TITR* | 0.02 | 0.08 | mg/L
Inorganic | Ammonia- Un-ionized (NH3) | DEP-SOP (NH3) | CALC* | 0.4 | 0.16 | ug/L
Non-Pesticide Organics | Anthracene | EPA 610-GC | GC-FID | 1 | 4 | ug/L
Non-Pesticide Organics | Anthracene | EPA 1625B | GC-MS | 0.05 | 0.2 | ug/L
Non-Pesticide Organics | Anthracene | EPA 625.1 | GC-MS | 0.05 | 0.2 | ug/L
Non-Pesticide Organics | Anthracene | SM 6410 B-2000 | GC-MS | 0.05 | 0.2 | ug/L
Non-Pesticide Organics | Anthracene | USGS O-3116-87, P.27 | GC-MS | 0.05 | 0.2 | ug/L
Non-Pesticide Organics | Anthracene | SM 6440B-2005 | HPLC | 0.8 | 3.2 | ug/L
Non-Pesticide Organics | Anthracene | ASTM D4657-92 | HPLC-UV | 0.8 | 3.2 | ug/L
Non-Pesticide Organics | Anthracene | EPA 610-HPLC | HPLC-UV-FLUOR | 0.8 | 3.2 | ug/L
Inorganic | Antimony | SM 3111 B -99(Sb) | FLAAS | 200 | 800 | mg/L
Inorganic | Antimony | SM 3113 B-10 | FLAAS | 0.8 | 3.2 | ug/L
Inorganic | Antimony | 200.7, Rev. 4.4 (1994) | ICP | 0.008 | 0.032 | mg/L
Inorganic | Antimony | 200.5, Rev. 4.2 (2003) | ICP/AES | 0.9 | 3.6 | ug/L
Inorganic | Antimony | ASTM 1976-07 | ICP/AES | 0.9 | 3.6 | ug/L
Inorganic | Antimony | USGS I-4471-97 | ICP/AES or ICP-MS | 0.4 | 1.6 | ug/L
Inorganic | Antimony | SM 3125 B-09 | ICP/MS | 0.4 | 1.6 | ug/L
Inorganic | Antimony | 206.5 (1978) | ICP/MS | 0.4 | 1.6 | ug/L
Inorganic | Antimony | SM 3120 B-11 | ICP-ES | 30 | 120 | ug/L
Inorganic | Antimony | AOAC 993.14 | ICP-MS | 0.4 | 1.6 | ug/L
Inorganic | Antimony | ASTM D5673­05 | ICP-MS | 0.4 | 1.6 | ug/L
Inorganic | Antimony | 200.8, Rev. 5.4 (1994) | ICP-MS | 0.4 | 1.6 | ug/L
Inorganic | Antimony | 200.9, Rev. 2.2 (1994) | TSGF | 0.8 | 3.2 | ug/L
Inorganic | Arsenic - Total | SM 3114 B-09 or 3114C-09 | AA | 1 | 4 | ug/L
Inorganic | Arsenic - Total | USGS I-4063-98 | AA | 1 | 4 | ug/L
Inorganic | Arsenic - Total | ASTM D2972-08 (B) | COLOR | 30 | 120 | ug/L
Inorganic | Arsenic - Total | SM 3500-As B-97 | COLOR | 1 | 4 | ug/L
Inorganic | Arsenic - Total | USGS I-3060-85 | COLOR | 10 | 40 | ug/L
Inorganic | Arsenic - Total | SM 3113 B-10 | FLAAS | 1 | 4 | ug/L
Inorganic | Arsenic - Total | ASTM D2972-93 (C) | GFAAS | 1 | 4 | ug/L
Inorganic | Arsenic - Total | ASTM D2972-08 (C) | Hydride | 1 | 4 | ug/L
Inorganic | Arsenic - Total | SM 3114 B 4.d | Hydride | 1 | 4 | ug/L
Inorganic | Arsenic - Total | ASTM D2972-93 (B) | Hydride | 1 | 4 | ug/L
Inorganic | Arsenic - Total | USGS I-3062-85 | Hydride | 1 | 4 | ug/L
Inorganic | Arsenic - Total | EPA 200.7, Rev. 4.4 (1994) | ICP | 0.008 | 0.032 | mg/L
Inorganic | Arsenic - Total | ASTM 1976-07 | ICP/AES | 1.4 | 5.6 | ug/L
Inorganic | Arsenic - Total | 200.5, Rev. 4.2 (2003) (68) | ICP/AES | 1.4 | 5.6 | ug/L
Inorganic | Arsenic - Total | SM 3125B-09 | ICP/MS | 1.4 | 5.6 | ug/L
Inorganic | Arsenic - Total | USGS I-4020-05 | ICP/MS | 1 | 4 | ug/L
Inorganic | Arsenic - Total | SM 3120 B-11 | ICP-ES | 4 | 16 | ug/L
Inorganic | Arsenic - Total | AOAC 993.14 | ICP-MS | 1.4 | 5.6 | ug/L
Inorganic | Arsenic - Total | ASTM D5673­05 | ICP-MS | 1.4 | 5.6 | ug/L
Inorganic | Arsenic - Total | 200.8, Rev. 5.4 (1994) | ICP-MS | 1.4 | 5.6 | ug/L
Inorganic | Arsenic - Total | 200.9, Rev. 2.2 (1994) | TSGF | 0.5 | 2 | ug/L
Inorganic | Arsenic - Trivalent | EPA 1632 | GFAAS | 3 | 10 | ng/L
Inorganic | Arsenic - Trivalent | USGS I-1190 | GFAAS | 0.3 | 1 | ug/L
Inorganic | Arsenic - Trivalent | USGS I-2191 | ICP-MS | 0.1 | 0.4 | ug/L
Inorganic | Arsenic - Trivalent | USGS I-2192 | ICP-MS | 0.6 | 2 | ug/L
Inorganic | Arsenic - Trivalent | USGS I-2193 | ICP-MS | 0.3 | 1 | ug/L
Inorganic | Barium | AES0029, Rev. 1991 Thermo Jarrell Ash Corp. | DCP | 50 | 200 | ug/L
Inorganic | Barium | SM 3111 D-99 | FLAAS | 0.4 | 0.44 | mg/L
Inorganic | Barium | USGS I-3084-85 | FLAAS | 100 | 400 | ug/L
Inorganic | Barium | SM 3113 B-10 | GFAAS | 2 | 8 | ug/L
Inorganic | Barium | ASTM D4382-02(07) | GFAAS | 2 | 8 | ug/L
Inorganic | Barium | 200.7, Rev. 4.4 (1994) | ICP | 0.001 | 0.004 | mg/L
Inorganic | Barium | 200.5, Rev. 4.2 (2003) (68) | ICP | 0.05 | 0.2 | ug/L
Inorganic | Barium | SM 3125B-09 | ICP/MS | 0.8 | 3.2 | ug/L
Inorganic | Barium | USGS I-4471-97 | ICP/MS ICP/AES | 0.8 | 3.2 | ug/L
Inorganic | Barium | SM 3120 B-11 | ICP-ES | 0.05 | 0.2 | ug/L
Inorganic | Barium | 200.8, Rev. 5.4 (1994) | ICP-MS | 0.8 | 3.2 | ug/L
Inorganic | Barium | ASTM D5673­05 | ICP-MS | 0.8 | 3.2 | ug/L
Inorganic | Barium | AOAC 993.14 3 | ICP-MS | 0.8 | 3.2 | ug/L
Pesticide | b-BHC | EPA 617 | GC | 0.01 | 0.04 | ug/L
Pesticide | b-BHC | ASTM 5812-96 (02) | GC | 0.01 | 0.04 | ug/L
Pesticide | b-BHC | 3M Corp 3M0222 (10/28/94) | GC | 0.01 | 0.04 | ug/L
Pesticide | b-BHC | ASTM D3086-90 | GC-ECD | 0.02 | 0.08 | ug/L
Pesticide | b-BHC | EPA 608.3 | GC-ECD | 0.01 | 0.04 | ug/L
Pesticide | b-BHC | SM 6630 B-2007 | GC-ECD | 0.01 | 0.04 | ug/L
Pesticide | b-BHC | SM 6630 C-2007 | GC-ECD | 0.01 | 0.04 | ug/L
Pesticide | b-BHC | EPA 625.1 | GC-MS | 2 | 8 | ug/L
Pesticide | b-BHC | SM 6410 B-2000 | GC-MS | 2 | 8 | ug/L
Non-Pesticide Organics | Benzene | EPA 1624B | GC-MS | 0.2 | 0.8 | ug/L
Non-Pesticide Organics | Benzene | EPA 624.1 | GC-MS | 0.2 | 0.8 | ug/L
Non-Pesticide Organics | Benzene | SM 6200 B-2011 | GC-MS | 0.5 | 2 | ug/L
Non-Pesticide Organics | Benzene | EPA 602 | GC-PID | 0.2 | 0.8 | ug/L
Non-Pesticide Organics | Benzidine | EPA 1625B | GC-MS | 100 | 400 | ug/L
Non-Pesticide Organics | Benzidine | EPA 625.1 | GC-MS | 100 | 400 | ug/L
Non-Pesticide Organics | Benzidine | SM 6410 B-2000 | GC-MS | 100 | 400 | ug/L
Non-Pesticide Organics | Benzidine | EPA 605 | HPLC | 100 | 400 | ug/L
Non-Pesticide Organics | Benzidine | EPA, 1978, p.1 | SPEC | 0.2 | 0.8 | ug/L
Non-Pesticide Organics | Benzo(a)anthracene | EPA 610-GC | GC-FID | 1 | 4 | ug/L
Non-Pesticide Organics | Benzo(a)anthracene | EPA 1625B | GC-MS | 0.05 | 0.2 | ug/L
Non-Pesticide Organics | Benzo(a)anthracene | EPA 625.1 | GC-MS | 0.05 | 0.2 | ug/L
Non-Pesticide Organics | Benzo(a)anthracene | SM 6410 B-2000 | GC-MS | 0.05 | 0.2 | ug/L
Non-Pesticide Organics | Benzo(a)anthracene | USGS 0-3116-87, p. 27 | GC-MS | 0.05 | 0.2 | ug/L
Non-Pesticide Organics | Benzo(a)anthracene | ASTM D4657-92 (98) | HPLC-UV | 0.05 | 0.2 | ug/L
Non-Pesticide Organics | Benzo(a)anthracene | EPA 610-HPLC | HPLC-UV-FLUOR | 0.05 | 0.2 | ug/L
Non-Pesticide Organics | Benzo(a)anthracene | SM 6440 B-2005 HPLC | HPLC-UV-FLUOR | 0.05 | 0.2 | ug/L
Non-Pesticide Organics | Benzo(a)pyrene | EPA 610-GC | GC-FID | 1 | 4 | ug/L
Non-Pesticide Organics | Benzo(a)pyrene | EPA 1625B | GC-MS | 0.05 | 0.2 | ug/L
Non-Pesticide Organics | Benzo(a)pyrene | EPA 625.1 | GC-MS | 0.05 | 0.2 | ug/L
Non-Pesticide Organics | Benzo(a)pyrene | SM 6410 B-2000 | GC-MS | 0.05 | 0.2 | ug/L
Non-Pesticide Organics | Benzo(a)pyrene | USGS 0-3116-87, p. 27 | GC-MS | 0.05 | 0.2 | ug/L
Non-Pesticide Organics | Benzo(a)pyrene | ASTM D4657-92 (98) | HPLC-UV | 0.03 | 0.12 | ug/L
Non-Pesticide Organics | Benzo(a)pyrene | EPA 610-HPLC | HPLC-UV-FLUOR | 0.03 | 0.12 | ug/L
Non-Pesticide Organics | Benzo(a)pyrene | SM 6440 B-2005 HPLC | HPLC-UV-FLUOR | 0.03 | 0.12 | ug/L
Non-Pesticide Organics | Benzo(b)fluoranthene | EPA 610-GC | GC-FID | 1 | 4 | ug/L
Non-Pesticide Organics | Benzo(b)fluoranthene | EPA 1625B | GC-MS | 0.05 | 0.2 | ug/L
Non-Pesticide Organics | Benzo(b)fluoranthene | EPA 625.1 | GC-MS | 0.05 | 0.2 | ug/L
Non-Pesticide Organics | Benzo(b)fluoranthene | SM 6410 B-2000 | GC-MS | 0.05 | 0.2 | ug/L
Non-Pesticide Organics | Benzo(b)fluoranthene | USGS 0-3116-87, p. 27 | GC-MS | 0.05 | 0.2 | ug/L
Non-Pesticide Organics | Benzo(b)fluoranthene | ASTM D4657-92 | HPLC-UV | 0.1 | 0.4 | ug/L
Non-Pesticide Organics | Benzo(b)fluoranthene | EPA 610-HPLC | HPLC-UV-FLUOR | 0.1 | 0.4 | ug/L
Non-Pesticide Organics | Benzo(b)fluoranthene | SM 6440 B-2005-HPLC | HPLC-UV-FLUOR | 0.1 | 0.4 | ug/L
Non-Pesticide Organics | Benzo(g,h,i)perylene | EPA 610-GC | GC-FID | 1 | 4 | ug/L
Non-Pesticide Organics | Benzo(g,h,i)perylene | EPA 1625B | GC-MS | 0.05 | 0.2 | ug/L
Non-Pesticide Organics | Benzo(g,h,i)perylene | EPA 625.1 | GC-MS | 0.05 | 0.2 | ug/L
Non-Pesticide Organics | Benzo(g,h,i)perylene | SM 6410 B-2000 | GC-MS | 0.05 | 0.2 | ug/L
Non-Pesticide Organics | Benzo(g,h,i)perylene | USGS 0-3116-87, p. 27 | GC-MS | 0.05 | 0.2 | ug/L
Non-Pesticide Organics | Benzo(g,h,i)perylene | ASTM D4657-92 | HPLC-UV | 0.05 | 0.2 | ug/L
Non-Pesticide Organics | Benzo(g,h,i)perylene | EPA 610-HPLC | HPLC-UV-FLUOR | 0.05 | 0.2 | ug/L
Non-Pesticide Organics | Benzo(g,h,i)perylene | SM 6440 B-2500-HPLC | HPLC-UV-FLUOR | 0.05 | 0.2 | ug/L
Non-Pesticide Organics | Benzo(k)fluoranthene | EPA 610-GC | GC-FID | 1 | 4 | ug/L
Non-Pesticide Organics | Benzo(k)fluoranthene | EPA 1625B | GC-MS | 0.05 | 0.2 | ug/L
Non-Pesticide Organics | Benzo(k)fluoranthene | EPA 625.1 | GC-MS | 0.05 | 0.2 | ug/L
Non-Pesticide Organics | Benzo(k)fluoranthene | SM 6410 B-2000 | GC-MS | 0.05 | 0.2 | ug/L
Non-Pesticide Organics | Benzo(k)fluoranthene | USGS 0-3116-87, p. 27 | GC-MS | 0.05 | 0.2 | ug/L
Non-Pesticide Organics | Benzo(k)fluoranthene | ASTM D4657-92 | HPLC-UV | 0.03 | 0.12 | ug/L
Non-Pesticide Organics | Benzo(k)fluoranthene | EPA 610-HPLC | HPLC-UV-FLUOR | 0.03 | 0.12 | ug/L
Non-Pesticide Organics | Benzo(k)fluoranthene | SM 6440 B-2500-HPLC | HPLC-UV-FLUOR | 0.03 | 0.12 | ug/L
Inorganic | Beryllium | SM 3500-Be D (19th ed.) | COLOR | 5 | 20 | ug/L
Inorganic | Beryllium | ASTM 4190-08 | DCP | 50 | 200 | ug/L
Inorganic | Beryllium | AES0029, Rev. 1991 Thermo Jarrell Ash Corp. | DCP | 50 | 200 | ug/L
Inorganic | Beryllium | ASTM D3645-08 (A) | FLAAS | 10 | 40 | ug/L
Inorganic | Beryllium | SM 3111 D -99 | FLAAS | 0.03 | 0.12 | mg/L
Inorganic | Beryllium | USGS I-3095-85 | FLAAS | 10 | 40 | ug/L
Inorganic | Beryllium | SM 3111 E-2011 | FLAAS | 10 | 40 | ug/L
Inorganic | Beryllium | SM 3113 B-10 | GFAAS | 0.02 | 0.08 | ug/L
Inorganic | Beryllium | ASTM D3645-08 (B) | GFAAS | 0.5 | 2 | ug/L
Inorganic | Beryllium | EPA 200.7, Rev. 4.4 (1994) | ICP/AES | 0.3 | 1.2 | ug/L
Inorganic | Beryllium | ASTM 1976-07 | ICP/AES | 0.3 | 1.2 | ug/L
Inorganic | Beryllium | USGS I-4471-97 | ICP/AES | 0.3 | 1.2 | ug/L
Inorganic | Beryllium | 200.5, Rev. 4.2 (2003) | ICP/AES-Axial | 0.02 | 0.08 | ug/L
Inorganic | Beryllium | USGS I-4471-97 (I-4472-97, USGS, 2003) | ICP/MS | 0.3 | 1.2 | ug/L
Inorganic | Beryllium | SM 3125B-09 | ICP/MS | 0.3 | 1.2 | ug/L
Inorganic | Beryllium | ASTM 5673-05 | ICP/MS | 0.3 | 1.2 | ug/L
Inorganic | Beryllium | EPA 200.8, Rev. 5.4 (1994) | ICP/MS | 0.3 | 1.2 | ug/L
Inorganic | Beryllium | AOAC 993.14 | ICP/MS | 0.3 | 1.2 | ug/L
Inorganic | Beryllium | SM 3120 B-11 | ICP-ES | 0.3 | 1.2 | ug/L
Inorganic | Beryllium | EPA 200.9, Rev. 2.2 (1994) | TSGF | 0.02 | 0.08 | ug/L
Radiologic | Beta - Gross (Total) | EPA 900.0 (Beta) | GAS-FLOW-P* | 4 | NA | pCi/L
Radiologic | Beta - Gross (Total) | SM 7110 B (Beta) | GAS-FLOW-P* | 4 | NA | pCi/L
Radiologic | Beta - Gross (Total) | SM 7110 B-00 (Beta) | GAS-FLOW-P* | 4 | NA | pCi/L
Radiologic | Beta - Gross (Total) | USGS 76-177 (1976), pp. 75 & 78 | GAS-FLOW-P* | 4 | NA | pCi/L
Radiologic | Beta - Gross (Total) | ASTM D1943-90 | GAS-FLOW-P* | 4 | NA | pCi/L
Radiologic | Beta - Gross (Total) | ASTM D1943-96 | GAS-FLOW-P* | 4 | NA | pCi/L
Non-Pesticide Organics | Bis(2-chloroethoxy)methane | EPA 611 | GC-ELCD-MCTD | 1 | 4 | ug/L
Non-Pesticide Organics | Bis(2-chloroethoxy)methane | EPA 1625B | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | Bis(2-chloroethoxy)methane | EPA 625.1 | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | Bis(2-chloroethoxy)methane | SM 6410 B-2000 | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | Bis(2-chloroethoxy)methane | USGS 0-3116-87, p. 27 | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | Bis(2-chloroethyl)ether | EPA 611 | GC-ELCD-MCTD | 0.3 | 1.2 | ug/L
Non-Pesticide Organics | Bis(2-chloroethyl)ether | EPA 1625B | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | Bis(2-chloroethyl)ether | EPA 625.1 | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | Bis(2-chloroethyl)ether | SM 6410 B-2000 | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | Bis(2-chloroethyl)ether | USGS 0-3116-87, p. 27 | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | Bis(2-chloroisopropyl)ether | EPA 625.1 | GC-MS | 1.9 | 7.6 | ug/L
Non-Pesticide Organics | Bis(2-chloroisopropyl)ether | EPA 611 | GC-ELCD-MCTD | 0.2 | 0.8 | ug/L
Non-Pesticide Organics | Bis(2-chloroisopropyl)ether | EPA 1625 | GC-MS | 3 | 12 | ug/L
Non-Pesticide Organics | Bis(2-chloroisopropyl)ether | SM 6410 B | GC-MS | 3 | 12 | ug/L
Non-Pesticide Organics | Bis(2-ethylhexyl)phthalate | EPA 606 | GC-ECD | 3 | 12 | ug/L
Non-Pesticide Organics | Bis(2-ethylhexyl)phthalate | EPA 1625B | GC-MS | 15 | 60 | ug/L
Non-Pesticide Organics | Bis(2-ethylhexyl)phthalate | EPA 625.1 | GC-MS | 15 | 60 | ug/L
Non-Pesticide Organics | Bis(2-ethylhexyl)phthalate | SM 6410 B-2000 | GC-MS | 15 | 60 | ug/L
Non-Pesticide Organics | Bis(2-ethylhexyl)phthalate | USGS 0-3116-87, p. 27 | GC-MS | 15 | 60 | ug/L
Non-Pesticide Organics | Bisphenol A (BPA) | ASTM D7065-11 | GC-MS | 0.2 | 0.8 | ug/L
Inorganic | BOD | ANSI, 1975, p. 17 | GM* | N/A | 2 | mg/L
Inorganic | BOD | AOAC 973.44 | GM* | N/A | 2 | mg/L
Inorganic | BOD | SM 5210B-01 | GM* | N/A | 2 | mg/L
Inorganic | BOD | USGS I-1578-78 | GM* | N/A | 2 | mg/L
Inorganic | BOD | Method 1003-8-2009, In-Situ, Inc. | GM* | N/A | 2 | mg/L
Inorganic | BOD | Method 10360, Hach Co. | GM* | N/A | 2 | mg/L
Inorganic | BOD-Carbonaceous (CBOD) | SM 5210B-01 | GM* | N/A | 2 | mg/L
Inorganic | BOD-Carbonaceous (CBOD) | Method 1003-8-2009, In-Situ, Inc. | GM* | N/A | 2 | mg/L
Inorganic | BOD-Carbonaceous (CBOD) | Method 10360, Hach Co. | GM* | N/A | 2 | mg/L
Inorganic | Boron | SM 4500-B B-00 | COLOR | 100 | 400 | ug/L
Inorganic | Boron | USGS I-3112-85(2) | COLOR | 100 | 400 | ug/L
Inorganic | Boron | ASTM D4190-08 | DCP | 50 | 200 | ug/L
Inorganic | Boron | AES0029, Rev. 1991 Thermo Jarrell Ash Corp. | DCP | 50 | 200 | ug/L
Inorganic | Boron | EPA 200.7, Rev. 4.4 (1994) | ICP | 0.003 | 0.012 | mg/L
Inorganic | Boron | 200.5, Rev. 4.2 (2003) | ICP | 0.3 | 1.2 | ug/L
Inorganic | Boron | USGS I-4471-97 | ICP | 10 | 40 | ug/L
Inorganic | Boron | ASTM D1976-07 | ICP/AES | 0.3 | 1.2 | ug/L
Inorganic | Boron | SM 3125B-09 | ICP/MS | 1 | 4 | ug/L
Inorganic | Boron | SM 3120 B-11 | ICP-ES | 0.3 | 1.2 | ug/L
Inorganic | Boron | ASTM D5673-05 | ICP-MS | 1 | 4 | ug/L
Inorganic | Boron | EPA 200.8 (Boron) | ICP-MS | 1 | 4 | ug/L
Inorganic | Bromate | EPA 300.1 (Bromate) | IC | 1.44 | 5.76 | ug/L
Inorganic | Bromate | EPA 300.0 B (Bromate), rev 2.1 | LC-ION | 0.02 | 0.08 | mg/L
Inorganic | Bromide | SM 4140 B-97 | CIE/UV | 0.09 | 0.36 | mg/L
Inorganic | Bromide | ASTM D6508-00 (05) | CIE/UV | 0.1 | 0.4 | mg/L
Inorganic | Bromide | EPA 300.0, Rev 2.1 (1993) | IC | 0.01 | 0.04 | mg/L
Inorganic | Bromide | EPA 300.1-1, Rev 1.0 (1997) | IC | 0.014 | 0.056 | mg/L
Inorganic | Bromide | SM 4110 B | IC | 0.11 | 0.44 | mg/L
Inorganic | Bromide | SM 4110 B-11, C-11, D-11 | IC | 0.11 | 0.44 | mg/L
Inorganic | Bromide | ASTM D4327­ 03 | IC | 0.11 | 0.44 | mg/L
Inorganic | Bromide | ASTM D1246-05 | IC | 0.11 | 0.44 | mg/L
Inorganic | Bromide | AOAC 993.30 | IC | 0.11 | 0.44 | mg/L
Inorganic | Bromide | USGS I-1125-85 | ISE | 0.11 | 0.44 | mg/L
Non-Pesticide Organics | Bromodichloromethane | EPA 601 | GC-ELCD | 0.2 | 0.8 | ug/L
Non-Pesticide Organics | Bromodichloromethane | SM 6200 C-2011 | GC-ELCD | 0.2 | 0.8 | ug/L
Non-Pesticide Organics | Bromodichloromethane | EPA 1624B | GC-MS | 0.2 | 0.8 | ug/L
Non-Pesticide Organics | Bromodichloromethane | EPA 624.1 | GC-MS | 0.2 | 0.8 | ug/L
Non-Pesticide Organics | Bromodichloromethane | SM 6200 B-2011 | GC-MS | 0.2 | 0.8 | ug/L
Non-Pesticide Organics | Bromoform | EPA 601 | GC-ELCD | 0.2 | 0.8 | ug/L
Non-Pesticide Organics | Bromoform | SM 6200 C-2011 | GC-ELCD | 0.2 | 0.8 | ug/L
Non-Pesticide Organics | Bromoform | EPA 1624B | GC-MS | 0.5 | 2 | ug/L
Non-Pesticide Organics | Bromoform | EPA 624.1 | GC-MS | 0.5 | 2 | ug/L
Non-Pesticide Organics | Bromoform | SM 6200 B-2011 | GC-MS | 0.2 | 0.8 | ug/L
Non-Pesticide Organics | Bromomethane | EPA 601 | GC-ELCD | 0.5 | 2 | ug/L
Non-Pesticide Organics | Bromomethane | SM 6200 C-2011 | GC-ELCD | 0.5 | 2 | ug/L
Non-Pesticide Organics | Bromomethane | EPA 1624B | GC-MS | 0.5 | 2 | ug/L
Non-Pesticide Organics | Bromomethane | EPA 624.1 | GC-MS | 0.5 | 2 | ug/L
Non-Pesticide Organics | Bromomethane | SM 6200 B-2011 | GC-MS | 0.5 | 2 | ug/L
Non-Pesticide Organics | Butyl benzyl phthalate | EPA 606 | GC-ECD | 3 | 12 | ug/L
Non-Pesticide Organics | Butyl benzyl phthalate | EPA 1625B | GC-MS | 5 | 20 | ug/L
Non-Pesticide Organics | Butyl benzyl phthalate | EPA 625.1 | GC-MS | 5 | 20 | ug/L
Non-Pesticide Organics | Butyl benzyl phthalate | SM 6410 B-2000 | GC-MS | 5 | 20 | ug/L
Non-Pesticide Organics | Butyl benzyl phthalate | USGS 0-3116-87, p. 27 | GC-MS | 5 | 20 | ug/L
Inorganic | Cadmium | USGS I-4138-89 | AA | 0.025 | 0.1 | mg/L
Inorganic | Cadmium | SM 3500-Cd D | COLOR | 5 | 20 | ug/L
Inorganic | Cadmium | D4190-08 | DCP | 1 | 4 | mg/L
Inorganic | Cadmium | AES0029, Rev. 1991 Thermo Jarrell Ash Corp. | DCP | 50 | 200 | ug/L
Inorganic | Cadmium | ANS P. 37 (Cadmium) | FLAAS | 5 | 20 | ug/L
Inorganic | Cadmium | AOAC 974.27 (Cadmium) | FLAAS | 5 | 20 | ug/L
Inorganic | Cadmium | ASTM D3557-12 (A or B) | FLAAS | 5 | 20 | ug/L
Inorganic | Cadmium | SM 3111 B -99 | FLAAS | 0.025 | 0.1 | mg/L
Inorganic | Cadmium | SM 3111 C -99 | FLAAS | 0.025 | 0.1 | mg/L
Inorganic | Cadmium | USGS I-3135-85 | FLAAS | 5 | 20 | ug/L
Inorganic | Cadmium | USGS I-3136-85 | FLAAS | 5 | 20 | ug/L
Inorganic | Cadmium | ASTM D3557­12 (D) | GFAAS | 0.1 | 0.4 | ug/L
Inorganic | Cadmium | EPA 200.9, Rev. 2.2 (1994) | ICP | 0.05 | 0.2 | ug/L
Inorganic | Cadmium | USGS I-1472-85 | ICP | 1 | 4 | ug/L
Inorganic | Cadmium | ASTM D1976-07 | ICP/AES | 0.1 | 0.4 | ug/L
Inorganic | Cadmium | 200.5, Rev. 4.2 (2003) (68) | ICP/AES | 0.1 | 0.4 | ug/L
Inorganic | Cadmium | EPA 200.7, Rev. 4.4 (1994) | ICP/AES | 0.001 | 0.004 | mg/L
Inorganic | Cadmium | AOAC 993.14 | ICP/MS | 0.5 | 2 | ug/L
Inorganic | Cadmium | ASTM D5673-05 | ICP/MS | 0.5 | 2 | ug/L
Inorganic | Cadmium | SM 3125 B -09 | ICP/MS | 0.5 | 2 | ug/L
Inorganic | Cadmium | USGS I-4471-97 | ICP=MS/AES | 0.5 | 2 | ug/L
Inorganic | Cadmium | SM 3120 B-11 | ICP-ES | 0.1 | 0.4 | ug/L
Inorganic | Cadmium | EPA 200.8 | ICP-MS | 0.5 | 2 | ug/L
Inorganic | Cadmium | EPA 200.8, Rev. 5.4 (1994) | ICP-MS | 0.5 | 2 | ug/L
Inorganic | Cadmium | EPA 200.9 | TSGF | 0.1 | 0.4 | ug/L
Inorganic | Cadmium | ASTM D3557­12 (C) | Voltammetry | 1 | 4 | ug/L
Inorganic | Calcium | AES0029, Rev. 1991 Thermo Jarrell Ash Corp. | DCP | 50 | 200 | ug/L
Inorganic | Calcium | ASTM D511-08(B) (Calcium) | FLAAS | 10 | 40 | ug/L
Inorganic | Calcium | SM 3111 B -99 | FLAAS | 0.08 | 0.32 | mg/L
Inorganic | Calcium | USGS I-3152-85 | FLAAS | 10 | 40 | ug/L
Inorganic | Calcium | ASTM D6919­09 | IC | 4 | 16 | mg/L
Inorganic | Calcium | USGS I-4471-97 (50) | ICP/AES | 10 | 40 | ug/L
Inorganic | Calcium | EPA 200.7, Rev. 4.4 (1994) ,200.5, Rev. 4.2 (2003) (68) | ICP/AES | 0.01 | 0.04 | mg/L
Inorganic | Calcium | SM 3125 B-09 | ICP/MS | 2 | 8 | ug/L
Inorganic | Calcium | SM 3120 B-11 | ICP-ES | 10 | 40 | ug/L
Inorganic | Calcium | ASTM D5673-05 | ICP-MS | 2 | 8 | ug/L
Inorganic | Calcium | EPA 200.8, Rev. 5.4 (1994) | ICP-MS | 2 | 8 | ug/L
Inorganic | Calcium | AOAC 993.14 | ICP-MS | 2 | 8 | ug/L
Inorganic | Calcium | ASTM D511-08 (A) | TITR* | 1 | 4 | mg/L
Inorganic | Calcium | SM 3500-Ca B-97 | TITR* | 0.5 | 2 | mg/L
Non-Pesticide Organics | Carbon Tetrachloride | EPA, 1978, p.130 | GC | 0.2 | 0.8 | ug/L
Non-Pesticide Organics | Carbon Tetrachloride | EPA 601 | GC-ELCD | 0.2 | 0.8 | ug/L
Non-Pesticide Organics | Carbon Tetrachloride | SM 6200 C-2011 | GC-ELCD | 0.2 | 0.8 | ug/L
Non-Pesticide Organics | Carbon Tetrachloride | EPA 1624B | GC-MS | 0.2 | 0.8 | ug/L
Non-Pesticide Organics | Carbon Tetrachloride | EPA 624.1 | GC-MS | 0.2 | 0.8 | ug/L
Non-Pesticide Organics | Carbon Tetrachloride | SM 6200 B-2011 | GC-MS | 0.2 | 0.8 | ug/L
Non-Pesticide Organics | Carbon- Total Organic | SM 5310 B-2011 | Combustion-IR | 1 | 4 | mg/L
Non-Pesticide Organics | Carbon- Total Organic | ASTM D7573-09 | Combustion-IR | 0.5 | 2 | mg/L
Non-Pesticide Organics | Carbon- Total Organic | AOAC 973.47 | IR | 1 | 4 | mg/L
Non-Pesticide Organics | Carbon- Total Organic | USGS Bk. 5, Ch. A3, Rev. 1987, p. 14 | IR | 1 | 4 | mg/L
Non-Pesticide Organics | Carbon- Total Organic | SM 5310 D-2011 | Oxidation-IR | 1 | 4 | mg/L
Non-Pesticide Organics | Carbon- Total Organic | ASTM 4839-03 | Oxidation-IR | 0.1 | 0.4 | mg/L
Non-Pesticide Organics | Carbon- Total Organic | SM 5310 C-2011 | Oxidation-IR-COUL-COND | 1 | 4 | mg/L
Inorganic | Chlorate | EPA 300.0 | IC | 0.003 | 0.012 | mg/L
Inorganic | Chlorate | EPA 300.1 | IC | 1.31 | 5.24 | ug/L
Pesticides | Chlordane | ASTM D5812-96 (02) | GC | 0.01* | 0.04 | ug/L
Pesticides | Chlordane | EPA 617 | GC | 0.01* | 0.04 | ug/L
Pesticides | Chlordane | EPA, 1978, p.7 | GC | 0.01* | 0.04 | ug/L
Pesticides | Chlordane | USGS O-3104-83 | GC | 0.01* | 0.04 | ug/L
Pesticides | Chlordane | 3M Corp 3M0222 (10/28/94) | GC | 0.01* | 0.04 | ug/L
Pesticides | Chlordane | ASTM D3086-90 | GC-ECD | 0.01* | 0.04 | ug/L
Pesticides | Chlordane | EPA 608.3 | GC-ECD | 0.01* | 0.04 | ug/L
Pesticides | Chlordane | SM 6630 B-2007 | GC-ECD | 0.01* | 0.04 | ug/L
Pesticides | Chlordane | SM 6630 C-2007 | GC-ECD | 0.01* | 0.04 | ug/L
Pesticides | Chlordane | EPA 625.1 | GC-MS | 40 | 160 | ug/L
Pesticides | Chlordane | SM 6410 B-2000 | GC-MS | 40 | 160 | ug/L
Inorganic | Chloride | ASTM 6508,-00 (05) | CIE/UV | 0.1 | 0.4 | mg/L
Inorganic | Chloride | SM 4140 B-11 | CIE/UV | 0.1 | 0.4 | mg/L
Inorganic | Chloride | USGS I-1187-85 | COLOR | 0.5 | 2 | mg/L
Inorganic | Chloride | AOAC 993.303 | IC | 0.02 | 0.08 | mg/L
Inorganic | Chloride | ASTM D4327 03 | IC | 0.02 | 0.08 | mg/L
Inorganic | Chloride | USGS I-2057-90(51) | IC | 0.02 | 0.08 | mg/L
Inorganic | Chloride | EPA 300.0, Rev 2.1 (1993) | IC | 0.02 | 0.08 | mg/L
Inorganic | Chloride | EPA 300.1-1, Rev 1.0 (1997) | IC | 0.004 | 0.016 | mg/L
Inorganic | Chloride | SM 4110 B -11 or C-11 | IC | 0.035 | 0.14 | mg/L
Inorganic | Chloride | ASTM D512­04(C ) | ISE | 0.5 | 2 | mg/L
Inorganic | Chloride | USGS I-2187-85 | SPEC-Auto | 1 | 4 | mg/L
Inorganic | Chloride | SM 4500-Cl- E-11 | SPEC-Auto | 1 | 4 | mg/L
Inorganic | Chloride | AOAC 973.51 | TITR* | 0.5 | 2 | mg/L
Inorganic | Chloride | ASTM D512-04(B) | TITR* | 0.5 | 2 | mg/L
Inorganic | Chloride | SM 4500 Cl D-11 | TITR* | 0.5 | 2 | mg/L
Inorganic | Chloride | SM 4500-Cl - B-11 | TITR* | 0.5 | 2 | mg/L
Inorganic | Chloride | SM 4500-Cl- C-11 | TITR* | 0.5 | 2 | mg/L
Inorganic | Chloride | USGS I-1183-85 | TITR* | 0.5 | 2 | mg/L
Inorganic | Chloride | USGS I-1184-85 | TITR* | 0.5 | 2 | mg/L
Inorganic | Chlorine- Total Residual | ASTM D1253­08 | AMP* | 0.01 | 0.04 | mg/L
Inorganic | Chlorine- Total Residual | DEP-SOP FT 2000-Amperometric | AMP* | 0.01 | 0.04 | mg/L
Inorganic | Chlorine- Total Residual | SM 4500-Cl D-00 | AMP* | 0.01 | 0.04 | mg/L
Inorganic | Chlorine- Total Residual | DEP-SOP FT 2000-Ion Selective Electrode | ISE | 0.2 | 0.8 | mg/L
Inorganic | Chlorine- Total Residual | ORION Orion | ISE | 0.1 | 0.4 | mg/L
Inorganic | Chlorine- Total Residual | SM 4500-Cl G-00 | SPEC | 0.2 | 0.8 | mg/L
Inorganic | Chlorine- Total Residual | DEP-SOP FT 2000-Spectrophotometric | SPEC* | 0.2 | 0.8 | mg/L
Inorganic | Chlorine- Total Residual | DEP-SOP FT 2000-Titration | TITR* | 0.04 | 0.16 | mg/L
Inorganic | Chlorine- Total Residual | SM 4500-Cl E -00 | TITR* | 0.04 | 0.16 | mg/L
Inorganic | Chlorine- Total Residual | SM 4500-Cl B-00 | TITR* | 0.04 | 0.16 | mg/L
Inorganic | Chlorine- Total Residual | SM 4500-Cl C-00 | TITR* | 0.04 | 0.16 | mg/L
Inorganic | Chlorine- Total Residual | SM 4500-Cl F-00 | TITR* | 0.04 | 0.16 | mg/L
Non-Pesticide Organics | Chlorobenzene | EPA, 1978, p.130 | GC | 0.2 | 0.8 | ug/L
Non-Pesticide Organics | Chlorobenzene | EPA 601 | GC-ELCD | 0.2 | 0.8 | ug/L
Non-Pesticide Organics | Chlorobenzene | SM 6200 C-2011 | GC-ELCD | 0.2 | 0.8 | ug/L
Non-Pesticide Organics | Chlorobenzene | EPA 1624B | GC-MS | 0.2 | 0.8 | ug/L
Non-Pesticide Organics | Chlorobenzene | EPA 624.1 | GC-MS | 0.2 | 0.8 | ug/L
Non-Pesticide Organics | Chlorobenzene | SM 6200 B-2011 | GC-MS | 0.2 | 0.8 | ug/L
Non-Pesticide Organics | Chlorobenzene | EPA 602 | GC-PID | 0.2 | 0.8 | ug/L
Non-Pesticide Organics | Chloroethane | EPA 601 | GC-ELCD | 0.5 | 2 | ug/L
Non-Pesticide Organics | Chloroethane | SM 6200 C-2011 | GC-ELCD | 0.5 | 2 | ug/L
Non-Pesticide Organics | Chloroethane | EPA 1624B | GC-MS | 0.5 | 2 | ug/L
Non-Pesticide Organics | Chloroethane | EPA 624.1 | GC-MS | 0.5 | 2 | ug/L
Non-Pesticide Organics | Chloroethane | SM 6200 B-2011 | GC-MS | 0.5 | 2 | ug/L
Non-Pesticide Organics | Chloroform | EPA, 1978, p.130 | GC | 0.2 | 0.8 | ug/L
Non-Pesticide Organics | Chloroform | EPA 601 | GC-ELCD | 0.2 | 0.8 | ug/L
Non-Pesticide Organics | Chloroform | SM 6200 C-2011 | GC-ELCD | 0.2 | 0.8 | ug/L
Non-Pesticide Organics | Chloroform | EPA 1624B | GC-MS | 0.2 | 0.8 | ug/L
Non-Pesticide Organics | Chloroform | EPA 624.1 | GC-MS | 0.2 | 0.8 | ug/L
Non-Pesticide Organics | Chloroform | SM 6200 B-2011 | GC-MS | 0.2 | 0.8 | ug/L
Non-Pesticide Organics | Chloromethane | EPA 601 | GC-ELCD | 0.5 | 2 | ug/L
Non-Pesticide Organics | Chloromethane | SM 6200 C-2011 | GC-ELCD | 0.5 | 2 | ug/L
Non-Pesticide Organics | Chloromethane | EPA 1624B | GC-MS | 0.5 | 2 | ug/L
Non-Pesticide Organics | Chloromethane | EPA 624.1 | GC-MS | 0.5 | 2 | ug/L
Non-Pesticide Organics | Chloromethane | SM 6200 B-2011 | GC-MS | 0.5 | 2 | ug/L
Biological | Chlorophyll a (corrected for pheophytin) | SM 10200 H.3. (acidification) | FLUOR* | N/A | 0.11 | ug/L
Biological | Chlorophyll a (corrected for pheophytin) | EPA 445.0 (acidification) | FLUOR* | N/A | 0.11 | ug/L
Biological | Chlorophyll a (corrected for pheophytin) | SM 10200 H.2.b. (acidification) | UV-VIS* | N/A | 3 | ug/L
Biological | Chlorophyll a (corrected for pheophytin) | EPA 446.0 (acidification) | UV-VIS* | N/A | 3 | ug/L
Biological | Chlorophyll a (free of pheophytin) | EPA 445.0 (narrow-bandpass filters) | FLUOR* | N/A | 0.0042 | ug/L
Biological | Chlorophyll a (free of pheophytin) | EPA 447.0 | HPLC* | N/A | 0.77 | ug/L
Biological | Chlorophyll a (free of pheophytin) | SM 10200 H.4. | HPLC* | N/A | 0.77 | ug/L
Inorganic | Chromium - Total | AOAC 974.27 | AA | 0.1 | 0.4 | mg/L
Inorganic | Chromium - Total | USGS I-3233-93 | AA | 0.1 | 0.4 | mg/L
Inorganic | Chromium - Total | SM 3500-Cr B-09 | COLOR | 0.1 | 0.4 | mg/L
Inorganic | Chromium - Total | ASTM 4190 -08 | DCP | 500 | 2000 | ug/L
Inorganic | Chromium - Total | AES0029, Rev. 1991 Thermo Jarrell Ash Corp. | DCP | 500 | 2000 | ug/L
Inorganic | Chromium - Total | USGS I-3236-85 | FLAAS | 50 | 200 | ug/L
Inorganic | Chromium - Total | ASTM D1687-02 (12) (B) | FLAAS | 50 | 200 | ug/L
Inorganic | Chromium - Total | SM 3111 C -11 | FLAAS | 0.1 | 0.4 | mg/L
Inorganic | Chromium - Total | SM 3111 B -99 | FLAAS | 0.1 | 0.4 | mg/L
Inorganic | Chromium - Total | ASTM D1687-02 (12) (C) | GFAAS | 1 | 4 | ug/L
Inorganic | Chromium - Total | SM 3113 B-10 | GFAAS | 0.1 | 0.4 | ug/L
Inorganic | Chromium - Total | ASTM D1976-07 | ICP/AES | 0.2 | 0.8 | ug/L
Inorganic | Chromium - Total | SM 3125 B-09 | ICP/MS | 0.9 | 3.6 | ug/L
Inorganic | Chromium - Total | 200.5 Rev 4.2 (2003) (68) | ICP/MS | 0.2 | 0.8 | ug/L
Inorganic | Chromium - Total | EPA 200.7 Rev 4.4 (1994) | ICP/MS | 0.004 | 0.016 | mg/L
Inorganic | Chromium - Total | AOAC 993.14 | ICP/MS | 0.8 | 3.2 | ug/L
Inorganic | Chromium - Total | ASTM D5673--05 | ICP/MS | 0.9 | 3.6 | ug/L
Inorganic | Chromium - Total | USGS I-4020-05 (70) | ICP/MS | 0.9 | 3.6 | ug/L
Inorganic | Chromium - Total | USGS I-4471-97 | ICP-AES | 0.2 | 0.8 | ug/L
Inorganic | Chromium - Total | SM 3120 B-11 | ICP-ES | 0.2 | 0.8 | ug/L
Inorganic | Chromium - Total | EPA 200.8, Rev. 5.4 (1994) | ICP-MS | 0.9 | 3.6 | ug/L
Inorganic | Chromium - Total | EPA 200.9, Rev. 2.2 (1994) | TSGF | 0.1 | 0.4 | ug/L
Inorganic | Chromium III | EPA 1669/1639 | GFAAS | 1 | 4 | ug/L
Inorganic | Chromium VI | SM 3111 C (Cr+6) | Chelation-FLAAS | 10 | 40 | ug/L
Inorganic | Chromium VI | SM 3111 C -99 | Chelation-FLAAS | 10 | 40 | ug/L
Inorganic | Chromium VI | USGS I-1232-85 | Chelation-FLAAS | 10 | 40 | ug/L
Inorganic | Chromium VI | ASTM D1687-92 (A) | COLOR | 10 | 40 | ug/L
Inorganic | Chromium VI | ASTM D1687­02(A) | COLOR | 50 | 200 | ug/L
Inorganic | Chromium VI | SM 3500 Cr B-09 | COLOR | 50 | 200 | ug/L
Inorganic | Chromium VI | USGS I-1230-85 | COLOR | 50 | 200 | ug/L
Inorganic | Chromium VI | EPA 218.6, Rev. 3.3 (1994) | IC | 0.3 | 1.2 | ug/L
Inorganic | Chromium VI | AOAC 993.23 | IC | 1.2 | 4.8 | ug/L
Inorganic | Chromium VI | ASTM D5257­03 | IC | 1.2 | 4.8 | ug/L
Inorganic | Chromium VI | SM 3500-Cr C -09 | IC | 1.2 | 4.8 | ug/L
Non-Pesticide Organics | Chrysene | EPA 610-GC | GC-FID | 1 | 4 | ug/L
Non-Pesticide Organics | Chrysene | EPA 1625B | GC-MS | 0.05 | 0.2 | ug/L
Non-Pesticide Organics | Chrysene | EPA 625.1 | GC-MS | 0.05 | 0.2 | ug/L
Non-Pesticide Organics | Chrysene | USGS 0-3116-87 P. 27 | GC-MS | 0.05 | 0.2 | ug/L
Non-Pesticide Organics | Chrysene | SM 6410 B-2000 | GC-MS | 0.05 | 0.2 | ug/L
Non-Pesticide Organics | Chrysene | ASTM D4657-92 (98) | HPLC-UV | 0.05 | 0.2 | ug/L
Non-Pesticide Organics | Chrysene | EPA 610-HPLC | HPLC-UV-FLUOR | 0.05 | 0.2 | ug/L
Non-Pesticide Organics | Chrysene | SM 6440 B-HPLC | HPLC-UV-FLUOR | 0.05 | 0.2 | ug/L
Inorganic | COD | ASTM D1252-06(B) | SPEC | 10 | 40 | mg/L
Inorganic | COD | EPA 410.4, Rev. 2.0 (1993) | SPEC | 10 | 40 | mg/L
Inorganic | COD | HACH 8000 (1979) | SPEC | 5 | 20 | mg/L
Inorganic | COD | OIC COD (1978) | SPEC | 5 | 20 | mg/L
Inorganic | COD | SM 5220 D-11 | SPEC | 10 | 40 | mg/L
Inorganic | COD | USGS I-3561-85 | SPEC | 10 | 40 | mg/L
Inorganic | COD | ANS P. 17 (COD) | TITR* | 10 | 40 | mg/L
Inorganic | COD | AOAC 973.46 | TITR* | 10 | 40 | mg/L
Inorganic | COD | ASTM D1252­ 06 (A) | TITR* | 10 | 40 | mg/L
Inorganic | COD | EPA 410.3 (Rev. 1978)1 | TITR* | 250 | 1000 | mg/L
Inorganic | COD | SM 5220B-2011 | TITR* | 5 | 20 | mg/L
Inorganic | COD | SM 5220C-2011 | TITR* | 5 | 20 | mg/L
Inorganic | COD | USGS I-3560-85 | TITR* | 5 | 20 | mg/L
Microbiological | Coliform Fecal | EPA P. 124 (MF) | BAC-MF* | 1 | 4 | Col/100 mL
Microbiological | Coliform Fecal | SM 9222 D-2006 (MF) | BAC-MF* | 1 | 4 | Col/100 mL
Microbiological | Coliform Fecal | USGS B-0050-85 (Fecal) (MF) | BAC-MF* | 1 | 4 | Col/100 mL
Microbiological | Coliform Fecal | Colilert-18® | BAC-MPN* | 2 | 8 | MPN/100 mL
Microbiological | Coliform Fecal | EPA P. 132 (MPN) | BAC-MPN* | 2 | 8 | MPN/100 mL
Microbiological | Coliform Fecal | EPA 1680 (MPN) | BAC-MPN* | 2 | 8 | MPN/100 mL
Microbiological | Coliform Fecal | EPA 1681 (MPN) | BAC-MPN* | 2 | 8 | MPN/100 mL
Microbiological | Coliform Fecal | SM 9221 C,E (MPN) | BAC-MPN* | 2 | 8 | MPN/100 mL
Microbiological | Coliform Fecal Presence of Cl | EPA P. 124 (MF) | BAC-MF* | 1 | 4 | Col/100 mL
Microbiological | Coliform Fecal Presence of Cl | SM 9222 D-2006 (MF) | BAC-MF* | 1 | 4 | Col/100 mL
Microbiological | Coliform Fecal Presence of Cl | EPA P. 132 (MPN) | BAC-MPN* | 2 | 8 | MPN/100 mL
Microbiological | Coliform Fecal Presence of Cl | SM 9221 C,E (MPN) | BAC-MPN* | 2 | 8 | MPN/100 mL
Microbiological | Coliform Total | EPA P. 108 | BAC-MF* | 1 | 4 | Col/100 mL
Microbiological | Coliform Total | SM 9222 B-2006 (MF) | BAC-MF* | 1 | 4 | Col/100 mL
Microbiological | Coliform Total | USGS B-0025-85 (Total) (MF) | BAC-MF* | 1 | 4 | Col/100 mL
Microbiological | Coliform Total | EPA P. 114 (MPN) | BAC-MPN* | 2 | 8 | MPN/100 mL
Microbiological | Coliform Total | SM 9221 B-2006 (MPN) | BAC-MPN* | 2 | 8 | MPN/100 mL
Microbiological | Coliform Total Presence of Cl | EPA P. 111 (MF) | BAC-MF* | 1 | 4 | Col/100 mL
Microbiological | Coliform Total Presence of Cl | SM 9222 B-2006 (MF) | BAC-MF* | 1 | 4 | Col/100 mL
Microbiological | Coliform Total Presence of Cl | EPA P. 114 (MPN) | BAC-MPN* | 2 | 8 | MPN/100 mL
Microbiological | Coliform Total Presence of Cl | SM 9221 B (MPN) | BAC-MPN* | 2 | 8 | MPN/100 mL
Inorganic | Color | SM 2120 B-2011 (Visual) | COLOR* | 5 | 20 | PCU
Inorganic | Color | USGS I-1250-85 | COLOR* | 5 | 20 | PCU
Inorganic | Color | SM 2120 B-2011 (Spectrophotometric) | SPEC* | 5 | 20 | PCU
Inorganic | Color | NCASI Technical Bulletin no. 253 (1971) | SPEC* | 5 | 20 | PCU
Inorganic | Color | NCASI Technical Bulletin no. 803 (Method 71.01) | SPEC* | 5 | 20 | PCU
Inorganic | Color | SM 2120F-2011 | SPEC* | N/A | N/A | ADMI
Inorganic | Color | SM 2120 C-2011 (Spectrophotometric) | SPEC* | 5 | 20 | PCU
Inorganic | Copper | HACH 8506 (1979) | COLOR | 20 | 80 | ug/L
Inorganic | Copper | SM 3500-Cu B-11 | COLOR | 5 | 20 | mg/L
Inorganic | Copper | SM 3500-Cu C-11 | COLOR | 5 | 20 | mg/L
Inorganic | Copper | ASTM D4190-08 | DCP |  | 0
Inorganic | Copper | AES0029, Rev. 1991, Thermo Jerrell Ash Corp | DCP | 50 | 200 | ug/L
Inorganic | Copper | SM 3111 B-11 | FLAAS | 0.1 | 0.4 | mg/L
Inorganic | Copper | SM 3111 C-11 | FLAAS | 0.1 | 0.4 | mg/L
Inorganic | Copper | ANSI, P. 37 (Copper) | FLAAS | 20 | 80 | ug/L
Inorganic | Copper | AOAC 974.27 (Copper) | FLAAS | 20 | 80 | ug/L
Inorganic | Copper | ASTM D1688-12 (A or B) | FLAAS | 0.7 | 3 | ug/L
Inorganic | Copper | USGS I-3270-85 | FLAAS | 20 | 80 | ug/L
Inorganic | Copper | USGS I-3271-85 | FLAAS | 20 | 80 | ug/L
Inorganic | Copper | SM 3113 B-10 | GFAAS | 0.7 | 3 | ug/L
Inorganic | Copper | ASTM D1688­12 (C) | GFAAS | 0.7 | 3 | ug/L
Inorganic | Copper | USGSI­4274­89 | GFAAS | 0.7 | 3
Inorganic | Copper | USGS I­4471­97 | ICP/AES | 0.3 | 1.2 | ug/L
Inorganic | Copper | SM 3120 B-11 | ICP-ES | 0.3 | 1.2 | ug/L
Inorganic | Copper | ASTM D1976(12) | ICP-ES | 0.3 | 1.2 | ug/L
Inorganic | Copper | 200.5, Rev. 4.2 (2003) | ICP-ES | 0.3 | 1.2 | ug/L
Inorganic | Copper | EPA 200.7, Rev. 4.4 (1994) | ICP-ES | 0.003 | 0.012 | mg/L
Inorganic | Copper | SM 3125 B-11 | ICP-MS | 0.5 | 2 | ug/L
Inorganic | Copper | ASTM D5673-10 | ICP-MS | 0.5 | 2 | ug/L
Inorganic | Copper | AOAC 993.14 | ICP-MS | 0.5 | 2 | ug/L
Inorganic | Copper | USGS I-4020-05 | ICP-MS | 0.5 | 2 | ug/L
Inorganic | Copper | EPA 200.8, Rev. 5.4 (1994) | ICP-MS | 0.5 | 2 | ug/L
Inorganic | Copper | EPA 200.9, Rev. 2.2 (1994) | STGFAA | 0.7 | 2.8 | ug/L
Microbiological | Cryptosporidium | EPA 1623 | IMS-FA | 1 | 4 | Oocysts/100L
Inorganic | Cyanide- Total | Kelada-01 (EPA 821–B–01–009, Rev. 1.2, August 2001) | COLOR-Auto | 5 | 20 | ug/L
Inorganic | Cyanide- Total | ASTM D7511-12 | DIFFUSION-AMP-Auto | 2 | 8 | ug/L
Inorganic | Cyanide- Total | ASTM D7284-13 | DIFFUSION-AMP-Auto | 2 | 8 | ug/L
Inorganic | Cyanide- Total | ASTM D2036-09 (A) | DIFFUSION-AMP-Auto | 2 | 8 | mg/L
Inorganic | Cyanide- Total | ASTM D2036-09 (A) | IC | 1 | 4 | mg/L
Inorganic | Cyanide- Total | ASTM D2036­09 (A) | ISE | 0.05 | 0.2 | mg/L
Inorganic | Cyanide- Total | SM 4500-CN­F-11 | ISE | 0.05 | 0.2 | mg/L
Inorganic | Cyanide- Total | ASTM D2036-09 (A) | SPEC | 5 | 20 | ug/L
Inorganic | Cyanide- Total | USGS I-3300-85 | SPEC | 5 | 20 | ug/L
Inorganic | Cyanide- Total | SM 4500-CN E-11 | SPEC | 5 | 20 | ugL
Inorganic | Cyanide- Total | EPA 335.4, Rev. 1.0 (1993) | SPEC-Auto | 5 | 20 | ug/L
Inorganic | Cyanide- Total | USGS 10­204­00-1­X | SPEC-Auto | 5 | 20 | ug/L
Inorganic | Cyanide- Total | USGS I­4302­852 | SPEC-Auto | 5 | 20 | ug/L
Inorganic | Cyanide- Total | ASTM D2036-09 (A) | TITR* | 0.2 | 1 | mg/L
Inorganic | Cyanide- Total | ANSI, P. 22 | TITR* | 0.2 | 1 | mg/L
Inorganic | Cyanide- Total | SM 4500-CN D-11 | TITR* | 0.2 | 1 | mg/L
Pesticide | d-BHC | ASTM D5812-96 (02) | GC | 0.01 | 0.04 | ug/L
Pesticide | d-BHC | EPA 617 | GC | 0.01 | 0.04 | ug/L
Pesticide | d-BHC | ASTM D3086-90 | GC-ECD | 0.01 | 0.04 | ug/L
Pesticide | d-BHC | EPA 608.3 | GC-ECD | 0.01 | 0.04 | ug/L
Pesticide | d-BHC | SM 6630 B & C-.2007 | GC-ECD | 0.01 | 0.04 | ug/L
Pesticide | d-BHC | 3M Corp 3M0222 (10/28/94) | GC-ECD | 0.011 | 0.044 | ug/L
Pesticide | d-BHC | EPA 625.1 | GC-MS | 2 | 8 | ug/L
Pesticide | d-BHC | SM 6410 B-2000 | GC-MS | 2 | 8 | ug/L
Pesticide | Demeton | EPA 622 | GC | 0.12 | 0.48 | ug/L
Pesticide | Demeton | EPA 614 | GC | 0.12 | 0.48 | ug/L
Pesticide | Demeton | EPA, 1978, p. 25 | GC | 0.12 | 0.48 | ug/L
Pesticide | Demeton | SM, 15th ed., Supplement, 1981, p. S51 | GC | 0.12 | 0.48 | ug/L
Pesticide | Demeton | EPA 625.1 | GC-MS | 2 | 8 | ug/L
Non-Pesticide Organics | Dibenzo(a,h)anthracene | EPA 610-GC | GC-FID | 1 | 4 | ug/L
Non-Pesticide Organics | Dibenzo(a,h)anthracene | EPA 1625B | GC-MS | 0.05 | 0.2 | ug/L
Non-Pesticide Organics | Dibenzo(a,h)anthracene | EPA 625.1 | GC-MS | 0.05 | 0.2 | ug/L
Non-Pesticide Organics | Dibenzo(a,h)anthracene | SM 6410 B-2000 | GC-MS | 0.05 | 0.2 | ug/L
Non-Pesticide Organics | Dibenzo(a,h)anthracene | USGS 0-3116-87 p.27 | GC-MS | 0.05 | 0.2 | ug/L
Non-Pesticide Organics | Dibenzo(a,h)anthracene | ASTM D4657-92 (98) | HPLC-UV | 0.05 | 0.2 | ug/L
Non-Pesticide Organics | Dibenzo(a,h)anthracene | EPA 610-HPLC | HPLC-UV-FLUOR | 0.05 | 0.2 | ug/L
Non-Pesticide Organics | Dibenzo(a,h)anthracene | SM 6440 B-2005HPLC | HPLC-UV-FLUOR | 0.05 | 0.2 | ug/L
Non-Pesticide Organics | Dibromochloromethane | EPA 601 | GC-ELCD | 0.2 | 0.8 | ug/L
Non-Pesticide Organics | Dibromochloromethane | SM 62200 C-2011 | GC-ELCD | 0.2 | 0.8 | ug/L
Non-Pesticide Organics | Dibromochloromethane | EPA 1624B | GC-MS | 0.2 | 0.8 | ug/L
Non-Pesticide Organics | Dibromochloromethane | EPA 624.1 | GC-MS | 0.2 | 0.8 | ug/L
Non-Pesticide Organics | Dibromochloromethane | SM 6200 B-2011 | GC-MS | 0.2 | 0.8 | ug/L
Pesticide | Dieldrin | EPA 617 | GC | 0.01 | 0.04 | ug/L
Pesticide | Dieldrin | ASTM D3086-90 | GC | 0.01 | 0.04 | ug/L
Pesticide | Dieldrin | ASTM D5812-96(02) | GC | 0.01 | 0.04 | ug/L
Pesticide | Dieldrin | EPA, 1978 p. 7 | GC | 0.01 | 0.04 | ug/L
Pesticide | Dieldrin | USGS O-3104-83 | GC | 0.01 | 0.04 | ug/L
Pesticide | Dieldrin | 3M Corp 3M0222 (10/28/94) | GC | 0.01 | 0.04 | ug/L
Pesticide | Dieldrin | EPA 608.3 | GC-ECD | 0.01 | 0.04 | ug/L
Pesticide | Dieldrin | SM 6630 B-2007 | GC-ECD | 0.01 | 0.04 | ug/L
Pesticide | Dieldrin | SM 6630 C-2007 | GC-ECD | 0.01 | 0.04 | ug/L
Pesticide | Dieldrin | EPA 625.1 | GC-MS | 2 | 8 | ug/L
Pesticide | Dieldrin | SM 6410 B-2000 | GC-MS | 2 | 8 | ug/L
Pesticide | Dieldrin | USGS O-1126-95 | GC-MS | 2 | 8 | ug/L
Non-Pesticide Organics | Diethyl phthalate | EPA 606 | GC-ECD | 1 | 4 | ug/L
Non-Pesticide Organics | Diethyl phthalate | EPA 1625B | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | Diethyl phthalate | EPA 625.1 | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | Diethyl phthalate | SM 6410 B-2000 | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | Diethyl phthalate | USGS 0-3116-87 p.27 | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | Dimethyl phthalate | EPA 606 | GC-ECD | 5 | 20 | ug/L
Non-Pesticide Organics | Dimethyl phthalate | EPA 1625B | GC-MS | 10 | 40 | ug/L
Non-Pesticide Organics | Dimethyl phthalate | EPA 625.1 | GC-MS | 10 | 40 | ug/L
Non-Pesticide Organics | Dimethyl phthalate | SM 6410 B-2000 | GC-MS | 10 | 40 | ug/L
Non-Pesticide Organics | Dimethyl phthalate | USGS 0-3116-87 p.27 | GC-MS | 10 | 40 | ug/L
Non-Pesticide Organics | Di-n-Butyl phthalate | EPA 606 | GC-ECD | 5 | 20 | ug/L
Non-Pesticide Organics | Di-n-Butyl phthalate | EPA 1625B | GC-MS | 5 | 20 | ug/L
Non-Pesticide Organics | Di-n-Butyl phthalate | EPA 625.1 | GC-MS | 5 | 20 | ug/L
Non-Pesticide Organics | Di-n-Butyl phthalate | SM 6410 B-2000 | GC-MS | 5 | 20 | ug/L
Non-Pesticide Organics | Di-n-Butyl phthalate | USGS 0-3116-87 p.27 | GC-MS | 5 | 20 | ug/L
Non-Pesticide Organics | Di-n-Octyl phthalate | EPA 606 | GC-ECD | 1 | 4 | ug/L
Non-Pesticide Organics | Di-n-Octyl phthalate | EPA 1625B | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | Di-n-Octyl phthalate | EPA 625.1 | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | Di-n-Octyl phthalate | SM 6410 B-2000 | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | Di-n-Octyl phthalate | USGS 0-3116-87 p.27 | GC-MS | 1 | 4 | ug/L
Inorganic | Dissolved Oxygen | ASTM D888-09 (C) | LDO* | 0.1 | 0.4 | mg/L
Inorganic | Dissolved Oxygen | Hach Method 10360, Rev. 1.2 (October 2011) | LDO* | 0.1 | 0.4 | mg/L
Inorganic | Dissolved Oxygen | Method 1002-8-2009, In-Situ, Inc. | LDO* | 0.1 | 0.4 | mg/L
Inorganic | Dissolved Oxygen | USGS I-1576-78 | POT* | 0.1 | 0.4 | mg/L
Inorganic | Dissolved Oxygen | ASTM D888-09 (B) | POT* | 0.1 | 0.4 | mg/L
Inorganic | Dissolved Oxygen | SM 4500-O G-.2011 | POT* | 0.1 | 0.4 | mg/L
Inorganic | Dissolved Oxygen | DEP-SOP FT 1500 | POT* or LDO* | 0.1 | 0.4 | mg/L
Inorganic | Dissolved Oxygen | SM 4500-O (B-F)-2011 | TITR* | 0.1 | 0.4 | mg/L
Inorganic | Dissolved Oxygen | ASTM D888-09 (A) | TITR* | 0.1 | 0.4 | mg/L
Inorganic | Dissolved Oxygen | AOAC 973.45B | TITR* | 0.1 | 0.4 | mg/L
Inorganic | Dissolved Oxygen | USGS I-1575-78 | TITR* | 0.1 | 0.4 | mg/L
Pesticide | Endosulfan I | ASTM D5812-96(02) | GC | 0.01 | 0.04 | ug/L
Pesticide | Endosulfan I | EPA 617 | GC | 0.01 | 0.04 | ug/L
Pesticide | Endosulfan I | EPA, 1978 p.7 | GC | 0.01 | 0.04 | ug/L
Pesticide | Endosulfan I | USGS O-3104-83 | GC | 0.01 | 0.04 | ug/L
Pesticide | Endosulfan I | 3M Corp 3M0222 (10/28/94) | GC | 0.01 | 0.04 | ug/L
Pesticide | Endosulfan I | ASTM D3086 -90 | GC-ECD | 0.01 | 0.04 | ug/L
Pesticide | Endosulfan I | EPA 608.3 | GC-ECD | 0.01 | 0.04 | ug/L
Pesticide | Endosulfan I | SM 6630 B-2007 | GC-ECD | 0.01 | 0.04 | ug/L
Pesticide | Endosulfan I | SM 6630 C-2007 | GC-ECD | 0.01 | 0.04 | ug/L
Pesticide | Endosulfan I | EPA 625.1 | GC-MS | 4 | 16 | ug/L
Pesticide | Endosulfan I | SM 6410 B-2000 | GC-MS | 4 | 16 | ug/L
Pesticide | Endosulfan I | USGS O-2002-01 | GC-MS | 4 | 16 | ug/L
Pesticide | Endosulfan II | ASTM D5812-96(02) | GC | 0.01 | 0.04 | ug/L
Pesticide | Endosulfan II | EPA 617 | GC | 0.01 | 0.04 | ug/L
Pesticide | Endosulfan II | EPA, 1978 p.7 | GC | 0.01 | 0.04 | ug/L
Pesticide | Endosulfan II | 3M Corp 3M022 (10/28/94) | GC | 0.01 | 0.04 | ug/L
Pesticide | Endosulfan II | ASTM D3086 -90 | GC-ECD | 0.01 | 0.04 | ug/L
Pesticide | Endosulfan II | EPA 608.3 | GC-ECD | 0.01 | 0.04 | ug/L
Pesticide | Endosulfan II | SM 6630 B-2007 | GC-ECD | 0.01 | 0.04 | ug/L
Pesticide | Endosulfan II | SM 6630 C-2007 | GC-ECD | 4 | 16 | ug/L
Pesticide | Endosulfan II | EPA 625.1 | GC-MS | 4 | 16 | ug/L
Pesticide | Endosulfan II | SM 6410 B-2000 | GC-MS | 4 | 16 | ug/L
Pesticide | Endosulfan II | USGS O-2002-01 | GC-MS | 4 | 16 | ug/L
Pesticide | Endosulfan sulfate | EPA 617 | GC | 0.02 | 0.08 | ug/L
Pesticide | Endosulfan sulfate | SM 6630 C-2007 | GC-ECD | 0.02 | 0.08 | ug/L
Pesticide | Endosulfan sulfate | EPA 625.1 | GC-MS | 2 | 8 | ug/L
Pesticide | Endosulfan sulfate | SM18 6410 B-2000 | GC-MS | 2 | 8 | ug/L
Pesticide | Endosulfan, sulfate | 3M Corp 3M0222 (10/28/94) | GC | 0.02 | 0.08 | ug/L
Pesticide | Endosulfan, sulfate | EPA 608.3 | GC-ECD | 0.02 | 0.08 | ug/L
Pesticide | Endrin | EPA P. 7 | GC | 0.05 | 0.2 | ug/L
Pesticide | Endrin | EPA 505 | GC | 0.02 | 0.08 | ug/L
Pesticide | Endrin | EPA 508 | GC | 0.02 | 0.08 | ug/L
Pesticide | Endrin | EP 617 | GC | 0.02 | 0.08 | ug/L
Pesticide | Endrin | EPA 1656 | GC | 0.02 | 0.08 | ug/L
Pesticide | Endrin | ASTM D5812-96 (02) | GC | 0.02 | 0.08 | ug/L
Pesticide | Endrin | USGS O-3104-83 | GC | 0.02 | 0.08 | ug/L
Pesticide | Endrin | 3M Corp 3M0222 (10/28/94) | GC | 0.02 | 0.08 | ug/L
Pesticide | Endrin | ASTM D3086-90 | GC-ECD | 0.02 | 0.08 | ug/L
Pesticide | Endrin | EPA 608.3 | GC-ECD | 0.02 | 0.08 | ug/L
Pesticide | Endrin | SM 6630 B-2007 | GC-ECD | 0.02 | 0.08 | ug/L
Pesticide | Endrin | SM 6630 C-2007 | GC-ECD | 0.02 | 0.08 | ug/L
Pesticide | Endrin | EPA 625.1 | GC-MS | 2 | 8 | ug/L
Pesticide | Endrin | SM 6410 B-2000 | GC-MS | 2 | 8 | ug/L
Pesticide | Endrin | EPA 525.1 | GC-MS | 2 | 8 | ug/L
Pesticide | Endrin | EPA 525.2 | GC-MS | 2 | 8 | ug/L
Pesticide | Endrin aldehyde | EPA 617 | GC | 0.02 | 0.08 | ug/L
Pesticide | Endrin aldehyde | EPA 608.3 | GC-ECD | 0.02 | 0.08 | ug/L
Pesticide | Endrin aldehyde | SM 6630 C-2007 | GC-ECD | 0.02 | 0.08 | ug/L
Pesticide | Endrin aldehyde | 3M Corp 3M0222 (10/28/94) | GC-ECD | 0.015 | 0.06 | ug/L
Microbiological | Enterococci | SM 9230C-2007 | BAC-MF* | 1 | 4 | Col/100 mL
Microbiological | Enterococci | EPA 1600 | BAC-MF* | 1 | 4 | Col/100 mL
Microbiological | Enterococci | Enterolert® | BAC-MPN* | 1 | 4 | MPN/100 mL
Microbiological | Enterococci | EPA (1978), p.139 | BAC-MPN* | 1 | 4 | MPN/100 mL
Microbiological | Enterococci | SM 9230B--2007 | BAC-MPN* | 1 | 4 | MPN/100 mL
Microbiological | Enterococci | SM 9230D--2008 | BAC-MPN* | 1 | 4 | MPN/100 mL
Microbiological | Enterococci | ASTM D6503-99 | BAC-MPN* | 1 | 4 | MPN/100 mL
Microbiological | Enterococci | EPA p. 143 | Pour Plate | 1 | 4 | Col/100 mL
Microbiological | Escherichia coli (E. coli) | EPA 1603 | BAC-MF* | 1 | 4 | MPN/100 mL
Microbiological | Escherichia coli (E. coli) | Hach mColiBlue®-24 | BAC-MF* | 1 | 4 | MPN/100 mL
Microbiological | Escherichia coli (E. coli) | SM 9221 B.2-2006/ 9221 F-2006 | BAC-MPN* | 1 | 4 | MPN/100 mL
Microbiological | Escherichia coli (E. coli) | SM 9223 B-2004 | BAC-MPN* | 1 | 4 | MPN/100 mL
Microbiological | Escherichia coli (E. coli) | AOAC 991.15 | BAC-MPN* | 1 | 4 | MPN/100 mL
Microbiological | Escherichia coli (E. coli) | Colilert®/Colilert-18® | BAC-MPN* | 1 | 4 | MPN/100 mL
Non-Pesticide Organics | Ethylbenzene | EPA 1624B | GC-MS | 0.5 | 2 | ug/L
Non-Pesticide Organics | Ethylbenzene | EPA 624.1 | GC-MS | 0.5 | 2 | ug/L
Non-Pesticide Organics | Ethylbenzene | SM 6200 B-2011 | GC-MS | 0.5 | 2 | ug/L
Non-Pesticide Organics | Ethylbenzene | EPA 602 | GC-PID | 0.2 | 0.8 | ug/L
Non-Pesticide Organics | Ethylbenzene | SM 6200 C-2011 | GC-PID | 0.5 | 2 | ug/L
Non-Pesticide Organics | Fluoranthene | EPA 610-GC | GC-FID | 1 | 4 | ug/L
Non-Pesticide Organics | Fluoranthene | SM 6440 B-GC | GC-FID | 1 | 4 | ug/L
Non-Pesticide Organics | Fluoranthene | EPA 1625B | GC-MS | 0.05 | 0.2 | ug/L
Non-Pesticide Organics | Fluoranthene | EPA 625.1 | GC-MS | 0.05 | 0.2 | ug/L
Non-Pesticide Organics | Fluoranthene | SM 6410 B-2000 | GC-MS | 0.05 | 0.2 | ug/L
Non-Pesticide Organics | Fluoranthene | USGS 0-3116-87 p.27 | GC-MS | 0.05 | 0.2 | ug/L
Non-Pesticide Organics | Fluoranthene | ASTM D4657-92 (98) | HPLC-UV | 0.3 | 1.2 | ug/L
Non-Pesticide Organics | Fluoranthene | EPA 610-HPLC | HPLC-UV-FLUOR | 0.3 | 1.2 | ug/L
Non-Pesticide Organics | Fluoranthene | SM 6440 B2005-HPLC | HPLC-UV-FLUOR | 0.3 | 1.2 | ug/L
Non-Pesticide Organics | Fluorene | EPA 610-GC | GC-FID | 1 | 4 | ug/L
Non-Pesticide Organics | Fluorene | SM 6440 B-GC | GC-FID | 1 | 4 | ug/L
Non-Pesticide Organics | Fluorene | EPA 1625B | GC-MS | 0.05 | 0.2 | ug/L
Non-Pesticide Organics | Fluorene | EPA 625.1 | GC-MS | 0.05 | 0.2 | ug/L
Non-Pesticide Organics | Fluorene | SM 6410 B-2000 | GC-MS | 0.05 | 0.2 | ug/L
Non-Pesticide Organics | Fluorene | USGS 0-3116-87 p.27 | GC-MS | 0.05 | 0.2 | ug/L
Non-Pesticide Organics | Fluorene | ASTM D4657-92 (98) | HPLC-UV | 0.3 | 1.2 | ug/L
Non-Pesticide Organics | Fluorene | EPA 610-HPLC | HPLC-UV-FLUOR | 0.3 | 1.2 | ug/L
Non-Pesticide Organics | Fluorene | SM 6440 B-2005-HPLC | HPLC-UV-FLUOR | 0.3 | 1.2 | ug/L
Inorganic | Fluoride | SM 4140 B-11 | CIE/UV | 0.1 | 0.4 | mg/L
Inorganic | Fluoride | ASTM D6508-10 | CIE/UV | 0.1 | 0.4 | mg/L
Inorganic | Fluoride | ASTM D6508, Rev. 2 | CIE/UV | 0.1 | 0.4 | mg/L
Inorganic | Fluoride | ASTM D1179-10(A) | COLOR | 0.1 | 0.4 | mg/L
Inorganic | Fluoride | SM 4500-F-D-11 | COLOR | 0.1 | 0.4 | mg/L
Inorganic | Fluoride | SM 4500-F-E-11 | COLOR-Auto | 0.1 | 0.4 | mg/L
Inorganic | Fluoride | EPA 300.0, Rev 2.1 (1993) and | IC | 0.009 | 0.036 | mg/L
Inorganic | Fluoride | EPA 300.1, Rev 1.0 (1997) | IC | 0.009 | 0.04 | mg/L
Inorganic | Fluoride | AOAC 993.30 | IC | 0.009 | 0.04 | mg/L
Inorganic | Fluoride | ASTM D4327­03 | IC | 0.009 | 0.04 | mg/L
Inorganic | Fluoride | SM 4110 B-11 or C-11 | IC | 0.009 | 0.04 | mg/L
Inorganic | Fluoride | EPA 300.0, Rev 2.1 (1993) | IC | 0.009 | 0.04 | mg/L
Inorganic | Fluoride | ASTM D1179­10(B) | ISE | 0.1 | 0.4 | mg/L
Inorganic | Fluoride | USGS I-4327-85 | ISE-Auto | 0.05 | 0.2 | mg/L
Pesticides | g-BHC | EPA 617 | GC | 0.01 | 0.04 | ug/L
Pesticides | g-BHC | ASTM D5812-96 (02) | GC | 0.01 | 0.04 | ug/L
Pesticides | g-BHC | EPA, 1978, p. 7 | GC | 0.01 | 0.04 | ug/L
Pesticides | g-BHC | USGS O-3104-83 | GC | 0.01 | 0.04 | ug/L
Pesticides | g-BHC | 3M Corp 3M0222 (10/28/94) | GC | 0.01 | 0.04 | ug/L
Pesticides | g-BHC | ASTM D3086-90 | GC-ECD | 0.03 | 0.12 | ug/L
Pesticides | g-BHC | EPA 608.3 | GC-ECD | 0.01 | 0.04 | ug/L
Pesticides | g-BHC | SM 6630 B | GC-ECD | 0.01 | 0.04 | ug/L
Pesticides | g-BHC | SM 6630 C | GC-ECD | 0.01 | 0.04 | ug/L
Pesticides | g-BHC | EPA 625.1 | GC-MS | 2 | 8 | ug/L
Pesticides | g-BHC | SM 6410 B-2000 | GC-MS | 2 | 8 | ug/L
Pesticides | g-BHC | USGS O-1126-95 | GC-MS | 2 | 8 | ug/L
Microbiological | Giardia | EPA 1623 | IMS-FA | 1 | 4 | Cysts/100L
Inorganic | Gold | AES0029, Rev. 1991 Thermo Jarrell Ash Corp. | DCP | 50 | 200 | ug/L
Inorganic | Gold | SM 3111 B-11 | FLAAS | 0.25 | 1 | mg/L
Inorganic | Gold | EPA 231.2(Issued 1978) | GFAA | 1 | 4 | ug/L
Inorganic | Gold | SM 3113 B-10 | GFAAS | 1 | 4 | ug/L
Inorganic | Gold | EPA 200.8 Rev 5.4 (1998) | ICP/MS | 2 | 8 | ug/L
Inorganic | Gold | SM 3125 B-11 | ICP/MS | 2 | 8 | ug/L
Inorganic | Gold | ASTM D5673-10 | ICP/MS | 2 | 8 | ug/L
Inorganic | Gold | AOAC 993.14 (3) | ICP/MS | 2 | 8 | ug/L
Pesticides | Guthion (Azinphos Methyl) | EPA 614 | GC | 1.5 | 6.0 | ug/L
Pesticides | Guthion (Azinphos Methyl) | EPA 622 | GC | 1.5 | 6.0 | ug/L
Pesticides | Guthion (Azinphos Methyl) | Standard Methods, 15th ed. Supplement, 1981, p. S51 | GC | 1.5 | 6.0 | ug/L
Pesticides | Guthion (Azinphos Methyl) | EPA, 1978, p. 25 | GC-FPD | 0.009 | 0.04 | ug/L
Pesticides | Guthion (Azinphos Methyl) | EPA 1657 | GC-FPD | 0.009 | 0.04 | ug/L
Pesticides | Guthion (Azinphos Methyl) | EPA 625.1 | GC-MS | 0.038 | 0.15 | ug/L
Pesticides | Guthion (Azinphos Methyl) | USGS O–1126–95 (1995) | GC-MS | 0.038 | 0.15 | ug/L
Non-Pesticide Organics | Halides- Adsorbable Organic | EPA 1650, Rev. C (1997) | COUL | 6.6 | 20 | ug/L
Inorganic | Halides-Total Organic | EPA 450.1 | COUL | 0.1 | 0.4 | mg/L
Inorganic | Hardness | EPA 130.1 (Issued 1971) | SPEC-Auto | 5 | 20 | mg/L
Inorganic | Hardness | AOAC 973.52B | TITR* | 5 | 20 | mg/L
Inorganic | Hardness | ASTM D1126­02 | TITR* | 5 | 20 | mg/L
Inorganic | Hardness | SM 2340 C-2011 | TITR* | 5 | 20 | mg/L
Inorganic | Hardness | USGS I-1338-85 | TITR* | 5 | 20 | mg/L
Inorganic | Hardness - Calculated | SM 2340 B-11 | CALC* | 0.2 | 0.8 | mg/L
Pesticide | Heptachlor | EPA 617 | GC | 0.01 | 0.04 | ug/L
Pesticide | Heptachlor | EPA 505 | GC | 0.01 | 0.04 | ug/L
Pesticide | Heptachlor | EPA 508 | GC | 0.01 | 0.04 | ug/L
Pesticide | Heptachlor | EPA 1656 | GC | 0.01 | 0.04 | ug/L
Pesticide | Heptachlor | ASTM D5812-96(02) | GC | 0.01 | 0.04 | ug/L
Pesticide | Heptachlor | EPA,1978, p.7 | GC | 0.01 | 0.04 | ug/L
Pesticide | Heptachlor | USGS O-3104-83 | GC | 0.01 | 0.04 | ug/L
Pesticide | Heptachlor | 3M Corp 3M0222 (10/28/94) | GC | 0.01 | 0.04 | ug/L
Pesticide | Heptachlor | EPA 608.3 | GC-ECD | 0.01 | 0.04 | ug/L
Pesticide | Heptachlor | SM 6630 B-2007 | GC-ECD | 0.01 | 0.04 | ug/L
Pesticide | Heptachlor | SM 6630 C-2007 | GC-ECD | 0.01 | 0.04 | ug/L
Pesticide | Heptachlor | ASTM D3086-90 | GC-ECD | 0.01 | 0.04 | ug/L
Pesticide | Heptachlor | EPA 625.1 | GC-MS | 2 | 8 | ug/L
Pesticide | Heptachlor | SM 6410 B-2000 | GC-MS | 1.9 | 7.6 | ug/L
Pesticide | Heptachlor | EPA 525.1 | GC-MS | 2 | 8 | ug/L
Pesticide | Heptachlor | EPA 525.2 | GC-MS | 2 | 8 | ug/L
Pesticide | Heptachlor epoxide | ASTM D5812-96(02) | GC | 0.01 | 0.04 | ug/L
Pesticide | Heptachlor epoxide | EPA 617 | GC | 0.01 | 0.04 | ug/L
Pesticide | Heptachlor epoxide | EPA, 1978, p.7 | GC | 0.01 | 0.04 | ug/L
Pesticide | Heptachlor epoxide | USGS O-3104-83 | GC | 0.01 | 0.04 | ug/L
Pesticide | Heptachlor epoxide | 3M Corp 3M0222 (10/28/94) | GC | 0.01 | 0.04 | ug/L
Pesticide | Heptachlor epoxide | SM, 15th ed., Supplement, p.S73 | GC | 0.01 | 0.04 | ug/L
Pesticide | Heptachlor epoxide | ASTM D3086-90 | GC-ECD | 0.01 | 0.04 | ug/L
Pesticide | Heptachlor epoxide | EPA 608.3 | GC-ECD | 0.01 | 0.04 | ug/L
Pesticide | Heptachlor epoxide | SM 6630 B-2007 | GC-ECD | 0.01 | 0.04 | ug/L
Pesticide | Heptachlor epoxide | SM 6630 C-2007 | GC-ECD | 0.01 | 0.04 | ug/L
Pesticide | Heptachlor epoxide | EPA 625.1 | GC-MS | 2 | 8 | ug/L
Pesticide | Heptachlor epoxide | SM 6410 B-2000 | GC-MS | 2.2 | 8.8 | ug/L
Non-Pesticide Organics | Hexachlorobenzene | EPA 612 | GC-ECD | 0.05 | 0.2 | ug/L
Non-Pesticide Organics | Hexachlorobenzene | EPA 1625B | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | Hexachlorobenzene | EPA 625.1 | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | Hexachlorobenzene | SM 6410 B-2000 | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | Hexachlorobenzene | USGS 0-3116-87 p.27 | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | Hexachlorobutadiene | EPA 612 | GC-ECD | 3 | 12 | ug/L
Non-Pesticide Organics | Hexachlorobutadiene | EPA 1625B | GC-MS | 3 | 12 | ug/L
Non-Pesticide Organics | Hexachlorobutadiene | EPA 625.1 | GC-MS | 3 | 12 | ug/L
Non-Pesticide Organics | Hexachlorobutadiene | USGS 0-3116-87 p.27 | GC-MS | 3 | 12 | ug/L
Non-Pesticide Organics | Hexachlorobutadiene | SM 6410 B-2000 | GC-MS | 3 | 12 | ug/L
Non-Pesticide Organics | Hexachlorocyclopentadiene | EPA 612 | GC-ECD | 0.04 | 0.16 | ug/L
Non-Pesticide Organics | Hexachlorocyclopentadiene | EPA 1625B | GC-MS | 3 | 12 | ug/L
Non-Pesticide Organics | Hexachlorocyclopentadiene | USGS 0-3116-87 p.27 | GC-MS | 3 | 12 | ug/L
Non-Pesticide Organics | Hexachlorocyclopentadiene | SM 6410 B-2000 | GC-MS | 3 | 12 | ug/L
Non-Pesticide Organics | Hexachloroethane | EPA 1625B | GC-MS | 10 | 40 | ug/L
Non-Pesticide Organics | Hexachloroethane | EPA 625.1 | GC-MS | 1.6 | 6.4 | ug/L
Non-Pesticide Organics | Hexachloroethane | SM 6410 B-2000 | GC-MS | 1.6 | 6.4 | ug/L
Non-Pesticide Organics | Hexachloroethane | USGS 0-3116-87 p.27 | GC-MS | 1.6 | 6.4 | ug/L
Inorganic | Hydrogen Sulfide | SM 4500-S2-H. | CALC* | 0.03 | 0.12 | mg/L
Non-Pesticide Organics | Indeno(1,2,3-cd)pyrene | EPA 610-GC | GC-FID | 1 | 4 | ug/L
Non-Pesticide Organics | Indeno(1,2,3-cd)pyrene | SM 6440 B-GC | GC-FID | 1 | 4 | ug/L
Non-Pesticide Organics | Indeno(1,2,3-cd)pyrene | EPA 1625B | GC-MS | 0.05 | 0.2 | ug/L
Non-Pesticide Organics | Indeno(1,2,3-cd)pyrene | EPA 625.1 | GC-MS | 0.05 | 0.2 | ug/L
Non-Pesticide Organics | Indeno(1,2,3-cd)pyrene | SM 6410 B-2000 | GC-MS | 0.05 | 0.2 | ug/L
Non-Pesticide Organics | Indeno(1,2,3-cd)pyrene | USGS 0-3116-87 p.27 | GC-MS | 0.05 | 0.2 | ug/L
Non-Pesticide Organics | Indeno(1,2,3-cd)pyrene | EPA 610-HPLC | HPLC-UV-FLUOR | 0.05 | 0.2 | ug/L
Non-Pesticide Organics | Indeno(1,2,3-cd)pyrene | SM 6440 B-2005-HPLC | HPLC-UV-FLUOR | 0.05 | 0.2 | ug/L
Non-Pesticide Organics | Indeno(1,2,3-cd)pyrene | ASTM D4657-92 (98) | HPLC-UV-FLUOR | 0.05 | 0.2 | ug/L
Inorganic | Iron | SM 3500-Fe-2011 | COLOR | 10 | 40 | ug/L
Inorganic | Iron | SM 3500-Fe-2011 | COLOR | 1 | 4 | ug/L
Inorganic | Iron | SM 3500-Fe-2011 | COLOR | 1 | 4 | ug/L
Inorganic | Iron | ASTM D1068-10 (C) | COLOR | 1 | 4 | ug/L
Inorganic | Iron | Hach Method 8008 | COLOR | 1 | 4 | ug/L
Inorganic | Iron | ASTM D4190­08 | DCP | 50 | 200 | ug/L
Inorganic | Iron | AES0029, Rev. 1991 Thermo Jarrell Ash Corp. | DCP | 50 | 200 | ug/L
Inorganic | Iron | SM 3111 B or C­11 | FLAAS | 0.12 | 0.48 | mg/L
Inorganic | Iron | SM 3113 B-10 | FLAAS | 1 | 4 | ug/L
Inorganic | Iron | ASTM D1068­ 10(A) | FLAAS | 0.12 | 0.48 | mg/L
Inorganic | Iron | AOAC 974.27 | FLAAS | 0.12 | 0.48 | mg/L
Inorganic | Iron | USGS I­3381­85 | FLAAS | 0.12 | 0.48 | mg/L
Inorganic | Iron | ASTM D1068­ 10(B) | GFAAS | 0.12 | 0.48 | mg/L
Inorganic | Iron | EPA 200.7, Rev. 4.4 (1994) | ICP | 0.03 | 0.12 | mg/L
Inorganic | Iron | SM 3125-B-11 | ICP/MS | 10 | 40 | ug/L
Inorganic | Iron | SM 3120-B-11 | ICP-ES | 3.3 | 13.2 | ug/L
Inorganic | Iron | EPA 200.5, Rev. 4.2 (2003) (68) | ICP-ES | 3.3 | 13.2 | ug/L
Inorganic | Iron | ASTM 1976-12 | ICP-ES | 3.3 | 13.2 | ug/L
Inorganic | Iron | USGS I-4471­97 | ICP-ES | 3.3 | 13.2 | ug/L
Inorganic | Iron | ASTM 5673-10 | ICP-MS | 10 | 40 | ug/L
Inorganic | Iron | EPA 200.8, Rev 5.4 (1998) | ICP-MS | 10 | 40 | ug/L
Inorganic | Iron | AOAC 993.14 | ICP-MS | 10 | 40 | ug/L
Inorganic | Iron | EPA 200.9, Rev. 2.2 (1994) | STGFAA | 0.12 | 0.48 | ug/L
Inorganic | Iron, Ferrous | SM 3500-Fe B.4c. | COLOR | 10 | 40 | ug/L
Non-Pesticide Organics | Isophorone | EPA 609 | GC | 1 | 4 | ug/L
Non-Pesticide Organics | Isophorone | EPA 1625B | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | Isophorone | EPA 625.1 | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | Isophorone | SM 6410 B-2000 | GC-MS | 2.2 | 8.8 | ug/L
Non-Pesticide Organics | Isophorone | USGS 0-3116-87, p.27 | GC-MS | 1 | 4 | ug/L
Inorganic | Lead | USGS I­3399-85 | AA | 0.5 | 2 | mg/L
Inorganic | Lead | USGS I-4403-89 | AA | 0.7 | 2.8 | ug/L
Inorganic | Lead | AOAC 974.27 | AA | 0.5 | 2 | mg/L
Inorganic | Lead | ASTM D3559­08 (D) | AA | 0.7 | 2.8 | ug/L
Inorganic | Lead | SM 3500-Pb D | COLOR | 1 | 4 | ug/L
Inorganic | Lead | SM 3500-Pb D-97 | COLOR | 1 | 4 | ug/L
Inorganic | Lead | AES0029, Rev. 1991 Thermo Jarrell Ash Corp. | DCP | 50 | 200 | ug/L
Inorganic | Lead | ASTM D4190­08 | DCP | 500 | 2000 | ug/L
Inorganic | Lead | SM 3111B OR C-99 | FLAAS | 0.5 | 2 | mg/L
Inorganic | Lead | SM 3113 B-10 | FLAAS | 0.7 | 2.8 | ug/L
Inorganic | Lead | ASTM D3559­08 (A or B) | FLAAS | 0.5 | 2 | mg/L
Inorganic | Lead | EPA 200.5, Rev. 4.2 (2003) (68) | ICP/AES | 1.1 | 4.4 | ug/L
Inorganic | Lead | EPA 200.7, Rev. 4.4 (1994) | ICP/AES | 10 | 40 | ug/L
Inorganic | Lead | ASTM 1976-12 | ICP/AES | 1.1 | 4.4 | ug/L
Inorganic | Lead | AOAC 993.14 | ICP/MS | 0.6 | 2.4 | ug/L
Inorganic | Lead | SM 3125 B-11 | ICP/MS | 0.6 | 2.4 | ug/L
Inorganic | Lead | ASTM D5673­10 | ICP/MS | 0.6 | 2.4 | ug/L
Inorganic | Lead | AOAC 993.14 | ICP/MS | 0.6 | 2.4 | ug/L
Inorganic | Lead | USGS I-4471-97 | ICP/MS | 1.1 | 4.4 | ug/L
Inorganic | Lead | SM 3120 B-11 | ICP-ES | 10 | 40 | ug/L
Inorganic | Lead | SM 3125 B-11 | ICP-MS | 1.1 | 4.4 | ug/L
Inorganic | Lead | EPA 200.8, Rev. 5.4 (1994) | ICP-MS | 0.6 | 2.4 | ug/L
Inorganic | Lead | USGS I­4471­97 | ICP-MS/ASE | 1.1 | 4.4 | ug/L
Inorganic | Lead | EPA 200.9, Rev. 2.2 (1994) | STGFAA | 0.7 | 2.8 | ug/L
Inorganic | Lead | ASTM D3559­08 (C) | Voltammetry | 1 | 4 | ug/L
Inorganic | Magnesium | AES0029, Rev. 1991 Thermo Jarrell Ash Corp. | DCP | 50 | 200 | ug/L
Inorganic | Magnesium | AOAC 974.27 (Magnesium) | FAAS | 7 | 28 | ug/L
Inorganic | Magnesium | ASTM D511­ 09(B) | FAAS | 7 | 28 | ug/L
Inorganic | Magnesium | SM 3111 B-11(Mg) | FAAS | 7 | 28 | ug/L
Inorganic | Magnesium | USGS I-3447-85 | FAAS | 20 | 80 | ug/L
Inorganic | Magnesium | ASTM D6919­09 | IC | 2.4 | 9.6 | ug/L
Inorganic | Magnesium | EPA 200.8, Rev. 5.4 (1998) | ICP | 10 | 40 | ug/L
Inorganic | Magnesium | USGS I-4471-97 | ICP | 20 | 80 | ug/L
Inorganic | Magnesium | SM 3120 B-11 | ICP-ES | 20 | 80 | ug/L
Inorganic | Magnesium | EPA 200.7, Rev. 4.4 (1994) | ICP/AES | 20 | 80 | ug/L
Inorganic | Magnesium | EPA 200.5, Rev. 4.2 (2003) (68) | ICP/AES | 20 | 80 | ug/L
Inorganic | Magnesium | ASTM D1976-12 | ICP/AES | 30 | 120 | ug/L
Inorganic | Magnesium | SM 3125 B-11 | ICP/MS | 10 | 40 | ug/L
Inorganic | Magnesium | ASTM D5673-10 | ICP/MS | 10 | 40 | ug/L
Inorganic | Magnesium | AOAC 993.14(3) | ICP/MS | 10 | 40 | ug/L
Pesticides | Malathion | SM, 15th ed., Supplement, 1981, p. S51 | GC | 0.05 | 0.1 | ug/L
Pesticides | Malathion | EPA 625.1 | GC/MS | 0.05 | 0.1 | ug/L
Pesticides | Malathion | USGS 0-1126-95 | GC/MS-SIM | 0.005 | 0.02 | ug/L
Pesticides | Malathion | SM 6630 B 2007 | GC-ECD | 0.2 | 0.8 | ug/L
Pesticides | Malathion | EPA 614 | GC-FPD | 0.05 | 0.1 | ug/L
Pesticides | Malathion | EPA 1657 | GC-FPD | 11 | 44 | ug/L
Pesticides | Malathion | EPA, 1978, p. 25 | GC-FPD | 0.05 | 0.1 | ug/L
Pesticides | Malathion | EPA 614 | GC-NPD | 0.05 | 0.2 | ug/L
Inorganic | Manganese | AOAC 974.27 | AA | 0.1 | 0.4 | mg/L
Inorganic | Manganese | SM 3500-Mn B-2011 | COLOR | 50 | 200 | ug/L
Inorganic | Manganese | AOAC 920.203 | COLOR | 0.2 | 0.8 | mg/L
Inorganic | Manganese | Hach method 8034 | COLOR | 0.2 | 0.8 | mg/L
Inorganic | Manganese | ASTM D4190­08 | DCP | 50 | 200 | ug/L
Inorganic | Manganese | AES0029, Rev. 1991 Thermo Jarrell Ash Corp. | DCP-AES | 50 | 200 | ug/L
Inorganic | Manganese | SM 3111 B-11 | FLAAS | 0.1 | 0.4 | mg/L
Inorganic | Manganese | SM 3113 B-10 | FLAAS | 0.2 | 0.8 | ug/L
Inorganic | Manganese | ASTM D858­12 (A or B) | FLAAS | 0.1 | 0.4 | ug/L
Inorganic | Manganese | USGS I­3454­85 | FLAAS | 0.1 | 0.4 | mg/L
Inorganic | Manganese | ASTM D858­12 (C) | GFAAS | 0.2 | 0.8 | ug/L
Inorganic | Manganese | AOAC 993.14 | ICP/MS | 0.1 | 0.4 | ug/L
Inorganic | Manganese | SM 3125 B-2011 | ICP/MS | 0.1 | 0.4 | ug/L
Inorganic | Manganese | EPA 200.5, Rev. 4.2 (2003) | ICP/MS | 0.06 | 0.24 | ug/L
Inorganic | Manganese | ASTM D1976-12 | ICP-AES | 2 | 8 | ug/L
Inorganic | Manganese | EPA 200.7, Rev. 4.4 (1994) | ICP-AES | 1 | 4 | ug/L
Inorganic | Manganese | USGS I­4471­97 | ICP-AES or MS | 0.1 | 0.4 | ug/L
Inorganic | Manganese | SM 3120 B-11 | ICP-ES | 0.2 | 0.8 | ug/L
Inorganic | Manganese | ASTM D5673­10 | ICP-MS | 0.1 | 0.4 | ug/L
Inorganic | Manganese | 200.8, Rev. 5.4 (1994) | ICP-MS | 0.1 | 0.4 | ug/L
Inorganic | Manganese | 200.9, Rev. 2.2 (1994) | STGFAA | 0.3 | 1.2 | ug/L
Inorganic | Mercury | SM 3112 B-2011 | CVAAS | 0.2 | 0.8 | ug/L
Inorganic | Mercury | EPA 245.1, Rev. 3.0 (1994) | CVAAS | 0.2 | 0.8 | ug/L
Inorganic | Mercury | ASTM D3223­12 | CVAAS | 0.2 | 0.8 | ug/L
Inorganic | Mercury | USGS I­3462­85 | CVAAS | 0.2 | 0.8 | ug/L
Inorganic | Mercury | AOAC 977.22 | CVAAS | 0.2 | 0.8 | ug/L
Inorganic | Mercury | EPA 245.2 (Issued 1974) | CVAAS-Auto | 0.2 | 0.8 | ug/L
Inorganic | Mercury | EPA 1631E | CVAFS | 0.0002 | 0.0008 | ug/L
Inorganic | Mercury | EPA 245.7 Rev. 2.0 (2005) (ng/L) | CVAFS | 0.0018 | 0.0050 | ug/L
Inorganic | Mercury | USGS I-4464-01 | CVAFS | 0.0018 | 0.0050 | ug/L
Pesticides | Methoxychlor | EPA 505 | GC | 0.05 | 0.2 | ug/L
Pesticides | Methoxychlor | EPA 508 | GC | 0.05 | 0.2 | ug/L
Pesticides | Methoxychlor | EPA 608.2 | GC | 0.05 | 0.2 | ug/L
Pesticides | Methoxychlor | EPA 617 | GC | 0.05 | 0.2 | ug/L
Pesticides | Methoxychlor | EPA 1656 | GC | 0.05 | 0.2 | ug/L
Pesticides | Methoxychlor | EPA 608.3 | GC | 0.05 | 0.2 | ug/L
Pesticides | Methoxychlor | ASTM D5812-96(02) | GC | 0.05 | 0.2 | ug/L
Pesticides | Methoxychlor | EPA, 1978, p.7 | GC | 0.05 | 0.2 | ug/L
Pesticides | Methoxychlor | USGS O-3104-83 | GC | 0.05 | 0.2 | ug/L
Pesticides | Methoxychlor | 3M Corp 3M0222 (10/28/94) | GC | 0.05 | 0.2 | ug/L
Pesticides | Methoxychlor | ASTM D3086-90 | GC-ECD | 0.05 | 0.2 | ug/L
Pesticides | Methoxychlor | SM 6630 B-2007 | GC-ECD | 0.05 | 0.2 | ug/L
Pesticides | Methoxychlor | SM 6630 C-2007 | GC-ECD | 0.05 | 0.2 | ug/L
Pesticides | Methoxychlor | EPA 525.1 | GC-MS | 0.84 | 3.4 | ug/L
Pesticides | Methoxychlor | EPA 525.2 | GC-MS | 0.13 | 0.52 | ug/L
Pesticides | Methoxychlor | EPA 625.1 | GC-MS | 0.84 | 3.4 | ug/L
Pesticides | Methoxychlor | USGS O-1126-95 | GC-MS-SIM | 0.01 | 0.04 | ug/L
Non-Pesticide Organics | Methylene Chloride (Dichloromethane) | SM 6200 C-2011 | GC | 0.5 | 2 | ug/L
Non-Pesticide Organics | Methylene Chloride (Dichloromethane) | EPA 601 | GC-ELCD | 0.5 | 2 | ug/L
Non-Pesticide Organics | Methylene Chloride (Dichloromethane) | EPA 1624B | GC-MS | 0.5 | 2 | ug/L
Non-Pesticide Organics | Methylene Chloride (Dichloromethane) | EPA 624.1 | GC-MS | 0.5 | 2 | ug/L
Non-Pesticide Organics | Methylene Chloride (Dichloromethane) | SM 6200 B-2011 | GC-MS | 0.5 | 2 | ug/L
Pesticides | Mirex | EPA P. 7 | GC | 0.02 | 0.08 | ug/L
Pesticides | Mirex | ASTM D3086-90 | GC | 0.02 | 0.08 | ug/L
Pesticides | Mirex | ASTM D5812-96(02) | GC | 0.02 | 0.08 | ug/L
Pesticides | Mirex | USGS O-3104-83 | GC | 0.02 | 0.08 | ug/L
Pesticides | Mirex | EPA 617 | GC | 0.02 | 0.08 | ug/L
Pesticides | Mirex | EPA 608.3 | GC | 0.02 | 0.08 | ug/L
Pesticides | Mirex | SM 6630 B-2007 | GC-ECD | 0.02 | 0.08 | ug/L
Pesticides | Mirex | SM 6630 C-2007 | GC-ECD | 0.02 | 0.08 | ug/L
Inorganic | Molybdenum | SM 3111 D-11 | AA | 0.5 | 2 | mg/L
Inorganic | Molybdenum | USGS I­3490­85 | AA | 0.5 | 2 | mg/L
Inorganic | Molybdenum | USGS I­3492­96 | AA | 1 | 4 | ug/L
Inorganic | Molybdenum | AES0029, Rev. 1991 Thermo Jarrell Ash Corp. | DCP | 0.3 | 1.2 | ug/L
Inorganic | Molybdenum | SM 3113 B-10 | FLAAS | 1 | 4 | ug/L
Inorganic | Molybdenum | EPA 200.7, Rev. 4.4 (1994) | ICP/AES | 0.004 | 0.016 | mg/L
Inorganic | Molybdenum | ASTM D1976-12 | ICP/AES | 8 | 32 | ug/L
Inorganic | Molybdenum | SM 3125 B-11 | ICP/MS | 0.3 | 1.2 | ug/L
Inorganic | Molybdenum | USGS I­4471­97 | ICP-AES or ICP-MS | 0.004 | 0.016 | mg/L
Inorganic | Molybdenum | SM 3120 B-11 | ICP-ES | 0.3 | 1.2 | ug/L
Inorganic | Molybdenum | EPA 200.8, Rev. 5.4 (1994) | ICP-MS | 0.3 | 1.2 | ug/L
Inorganic | Molybdenum | ASTM D5673­10 | ICP-MS | 0.3 | 1.2 | ug/L
Inorganic | Molybdenum | AOAC 993.14 | ICP-MS | 0.3 | 1.2 | ug/L
Non-Pesticide Organics | Naphthalene | EPA 610-GC | GC | 0.05 | 0.2 | ug/L
Non-Pesticide Organics | Naphthalene | EPA 1625B | GC-MS | 0.05 | 0.2 | ug/L
Non-Pesticide Organics | Naphthalene | EPA 625.1 | GC-MS | 0.05 | 0.2 | ug/L
Non-Pesticide Organics | Naphthalene | SM 6410 B-2000 | GC-MS | 0.05 | 0.2 | ug/L
Non-Pesticide Organics | Naphthalene | USGS 0-3116-87 p.27 | GC-MS | 0.05 | 0.2 | ug/L
Non-Pesticide Organics | Naphthalene | EPA 610-HPLC | HPLC | 2 | 8 | ug/L
Non-Pesticide Organics | Naphthalene | SM 6440 B-HPLC | HPLC-UV-FLUOR | 2 | 8 | ug/L
inorganic | Nickel | SM 3111 B or C -2011 | AA | 0.3 | 1.2 | mg/L
inorganic | Nickel | SM 3113 B-10 | AA | 0.6 | 2.4 | ug/L
inorganic | Nickel | ASTM D1886­08 (A or B) | AA | 0.3 | 1.2 | mg/L
inorganic | Nickel | ASTM D1886­08 (C) | AA | 0.6 | 2.4 | ug/L
Inorganic | Nickel | USGS I­3499­85 | AA | 1 | 4 | ug/L
Inorganic | Nickel | USGS I­4503­89 | AA | 0.6 | 2.4 | ug/L
inorganic | Nickel | ASTM D4190­08 | DCP | 50 | 200 | ug/L
Inorganic | Nickel | AES0029, Rev. 1991 Thermo Jarrell Ash Corp. | DCP | 50 | 200 | ug/L
inorganic | Nickel | EPA 200.5, Rev. 4.2 (2003) | ICP | 0.6 | 2.4 | ug/L
inorganic | Nickel | ASTM D5673-10 | ICP | 0.5 | 2 | ug/L
inorganic | Nickel | EPA 200.7, Rev. 4.4 (1994) | ICP-ES | 0.005 | 0.02 | mg/L
inorganic | Nickel | ASTM D1976-12 | ICP-ES | 15 | 60 | ug/L
inorganic | Nickel | SM 3120 B-11 | ICP-ES | 15 | 60 | ug/L
Inorganic | Nickel | USGS I­4471­97 | ICP-ES | 0.6 | 2.4 | ug/L
inorganic | Nickel | SM 3125 B-2011 | ICP-MS | 0.5 | 2 | ug/L
inorganic | Nickel | EPA 200.8, Rev. 5.4 (1994) | ICP-MS | 0.5 | 2 | ug/L
Inorganic | Nickel | AOAC 993.14 | ICP-MS | 0.8 | 3.2 | ug/L
Inorganic | Nickel | USGS I-4020-05 | ICP-MS | 0.8 | 3.2 | ug/L
inorganic | Nickel | EPA 200.9, Rev. 2.2 (1994) | STGFAA | 0.6 | 2.4 | ug/L
Inorganic | Nitrate (N) | SM 4140 B-2011 | CIE/UV | 0.082 | 0.4 | mg/L
Inorganic | Nitrate (N) | ASTM D6508-10 (D6508, Rev. 2) | CIE/UV | 0.082 | 0.4 | mg/L
Inorganic | Nitrate (N) | EPA 352.1 (Issued 1971) | COLOR | 0.1 | 0.4 | mg/L
Inorganic | Nitrate (N) | SM 419 D | COLOR | 0.1 | 0.4 | mg/L
Inorganic | Nitrate (N) | ANSI, 1975, P. 28 | COLOR | 0.1 | 0.4 | mg/L
Inorganic | Nitrate (N) | AOAC 973.50 | COLOR | 0.1 | 0.4 | mg/L
Inorganic | Nitrate (N) | EPA 353.2 | COLOR-Auto (Calculated Difference)* | 0.1 | 0.4 | mg/L
Inorganic | Nitrate (N) | Chinchilla Easy (1-Reagent) Nitrate Method, Rev. 11/12/11 | COLOR-Auto (Calculated Difference)* | 0.011 | 0.044 | mg/L
Inorganic | Nitrate (N) | USGS I-2547-11 | COLOR-Auto (Calculated Difference)* | 0.1 | 0.4 | mg/L
Inorganic | Nitrate (N) | USGS I-2548-11 | COLOR-Auto (Calculated Difference)* | 0.1 | 0.4 | mg/L
Inorganic | Nitrate (N) | NECi Method N07-0003, Rev. 9.0, March 2014 | COLOR-Auto (Calculated Difference)* | 0.1 | 0.4 | mg/L
Inorganic | Nitrate (N) | SM 4500-NO3 F-2011 | COLOR-Auto (Calculated Difference)* | 0.1 | 0.4 | mg/L
Inorganic | Nitrate (N) | ASTM D3867-04 (A) | COLOR-Auto (Calculated Difference)* | 0.1 | 0.4 | mg/L
Inorganic | Nitrate (N) | USGS I-4545-85 | COLOR-Auto (Calculated Difference)* | 0.1 | 0.4 | mg/L
Inorganic | Nitrate (N) | EPA 300.0, Rev 2.1 (1993) | IC | 0.1 | 0.4 | mg/L
Inorganic | Nitrate (N) | EPA 300.1, Rev 1.0 (1997) | IC | 0.017 | 0.068 | mg/L
Inorganic | Nitrate (N) | SM 4110 B-2011 | IC | 0.1 | 0.4 | mg/L
Inorganic | Nitrate (N) | SM 4110 C-2011 | IC | 0.017 | 0.068 | mg/L
Inorganic | Nitrate (N) | ASTM D4327-03 | IC | 0.1 | 0.4 | mg/L
Inorganic | Nitrate (N) | AOAC 993.30 | IC | 0.1 | 0.4 | mg/L
Inorganic | Nitrate (N) | SM 4500- NO3 -D -2011 | ISE | 0.14 | 0.56 | mg/L
Inorganic | Nitrate (N) | Hach Method 10206, Rev. 2.1, January 2013 | SPEC | 0.05 | 0.2 | mg/L
Inorganic | Nitrate (N) | SM 4500-NO3 E-2011 | SPEC (Calculated Difference)* | 0.002 | 0.01 | mg/L
Inorganic | Nitrate (N) | ASTM D3867-04 (B) | SPEC (Calculated Difference)* | 0.002 | 0.01 | mg/L
Inorganic | Nitrate-Nitrite (N) | SM 4140 B-2011 | CIE-UV | 0.084 | 0.336 | mg/L
Inorganic | Nitrate-Nitrite (N) | ASTM D6508-10 (D6508, Rev. 2) | CIE-UV | 0.084 | 0.34 | mg/L
Inorganic | Nitrate-Nitrite (N) | ASTM D3867 -04 (B) | COLOR | 0.01 | 0.05 | mg/L
Inorganic | Nitrate-Nitrite (N) | SM 4500-NO3 E -2011 (Nitrate-Nitrite (N)) | COLOR | 0.01 | 0.05 | mg/L
Inorganic | Nitrate-Nitrite (N) | SM 4500-NO3 F-2011(Nitrate-Nitrite (N)) | COLOR-Auto | 0.01 | 0.05 | mg/L
Inorganic | Nitrate-Nitrite (N) | SM 4500-NO3 H-2011 | COLOR-Auto | 0.01 | 0.05 | mg/L
Inorganic | Nitrate-Nitrite (N) | USGS I-4545-85 | COLOR-Auto | 0.01 | 0.05 | mg/L
Inorganic | Nitrate-Nitrite (N) | USGS I-2545-90 | COLOR-Auto | 0.01 | 0.05 | mg/L
Inorganic | Nitrate-Nitrite (N) | Chinchilla Easy (1-Reagent) Nitrate Method, Rev. 11/12/11 | COLOR-Auto | 0.01 | 0.04 | mg/L
Inorganic | Nitrate-Nitrite (N) | USGS I-2547-11 | COLOR-Auto | 0.01 | 0.04 | mg/L
Inorganic | Nitrate-Nitrite (N) | USGS I-2548-11 | COLOR-Auto | 0.002 | 0.01 | mg/L
Inorganic | Nitrate-Nitrite (N) | NECi Method N07-0003, Rev. 9.0, March 2014 | COLOR-Auto | 0.002 | 0.01 | mg/L
Inorganic | Nitrate-Nitrite (N) | ASTM D4327-03 | IC | 0.004 | 0.016 | mg/L
Inorganic | Nitrate-Nitrite (N) | EPA 300.0, Rev 2.1 (1993) | IC | 0.004 | 0.016 | mg/L
Inorganic | Nitrate-Nitrite (N) | 300.1, Rev 1.0 (1997) | IC | 0.017 | 0.068 | mg/L
Inorganic | Nitrate-Nitrite (N) | SM 4110 B-2011 | IC | 0.004 | 0.016 | mg/L
Inorganic | Nitrate-Nitrite (N) | SM 4110 C-2011 | IC | 0.017 | 0.068 | mg/L
Inorganic | Nitrate-Nitrite (N) | AOAC 993.30 | IC | 0.004 | 0.016 | mg/L
Inorganic | Nitrate-Nitrite (N) | Hach Method 10206, Rev. 2.1, January 2013 | SPEC | 0.05 | 0.2 | mg/L
Inorganic | Nitrate-Nitrite (N) | ASTM D3867-04 (A) | SPEC-Auto | 0.01 | 0.05 | mg/L
Inorganic | Nitrate-Nitrite (N) | EPA 353.2, Rev. 2.0 (1993) | SPEC-Auto | 0.01 | 0.05 | mg/L
inorganic | Nitrite (N) | SM 4140 B-2011 | CIE-UV | 0.102 | 0.408 | mg/L
inorganic | Nitrite (N) | ASTM D6508-10 (D6508, Rev. 2) | CIE-UV | 0.102 | 0.408 | mg/L
Inorganic | Nitrite (N) | EPA 353.2, Rev. 2.0 (1993) | COLOR-Auto | 0.005 | 0.02 | mg/L
Inorganic | Nitrite (N) | ASTM D3867­04 (A) | COLOR-Auto | 0.005 | 0.02 | mg/L
inorganic | Nitrite (N) | SM 4500-NO3 -F2011 | COLOR-Auto | 0.005 | 0.02 | mg/L
Inorganic | Nitrite (N) | USGS I-4540-85 | COLOR-Auto | 0.005 | 0.02 | mg/L
Inorganic | Nitrite (N) | USGS I­4545­85 | COLOR-Auto | 0.005 | 0.02 | mg/L
Inorganic | Nitrite (N) | Chinchilla Easy (1-Reagent) Nitrate Method, Rev. 11/12/11 | COLOR-Auto | 0.005 | 0.02 | mg/L
Inorganic | Nitrite (N) | USGS I-2547-11 | COLOR-Auto | 0.01 | 0.04 | mg/L
Inorganic | Nitrite (N) | USGS I-2458-11 | COLOR-Auto | 0.01 | 0.04 | mg/L
Inorganic | Nitrite (N) | NECi Method N07-0003, Rev. 9.0, March 2014 | COLOR-Auto | 0.01 | 0.04 | mg/L
Inorganic | Nitrite (N) | EPA 300.0, Rev 2.1, 1993 | IC | 0.004 | 0.016 | mg/L
Inorganic | Nitrite (N) | EPA 300.1, Rev 1.0 (1997) | IC | 0.015 | .060 | mg/L
inorganic | Nitrite (N) | ASTM 4327-03 | IC | 0.004 | 0.016 | mg/L
Inorganic | Nitrite (N) | SM 4110 C-2011 | IC | 0.015 | .060 | mg/L
Inorganic | Nitrite (N) | AOAC 993.30 | IC | 0.004 | 0.016 | mg/L
Inorganic | Nitrite (N) | SM 4110 B-2011 | IC | 0.004 | 0.016 | mg/L
Inorganic | Nitrite (N) | ASTM D3867­04 (B) | SPEC | 0.01 | 0.04 | mg/L
inorganic | Nitrite (N) | SM 4500-NO2 B-2011 | SPEC | 0.005 | 0.02 | mg/L
Inorganic | Nitrite (N) | SM 4500-NO3 -E -2011 | SPEC | 0.01 | 0.04 | mg/L
Inorganic | Nitrite (N) | Hach Method 8507 (1979) | SPEC | 0.005 | 0.02 | mg/L
Non-Pesticide Organics | Nitrobenzene | EPA 609 | GC-FID | 4 | 20 | ug/L
Non-Pesticide Organics | Nitrobenzene | EPA 1625B | GC-MS | 2 | 8 | ug/L
Non-Pesticide Organics | Nitrobenzene | EPA 625.1 | GC-MS | 2 | 8 | ug/L
Non-Pesticide Organics | Nitrobenzene | SM 6410 B-2000 | GC-MS | 2 | 8 | ug/L
Non-Pesticide Organics | Nitrobenzene | USGS 0-3116-87 p.27 | GC-MS | 2 | 8 | ug/L
Non-Pesticide Organics | Nitrobenzene | ASTM D4657-92 (98) | HPLC | 2 | 8 | ug/L
Inorganic | Nitrogen- Total (Direct Measurement) | ASTM D5176-08 (2015) | Pyrolysis-Chemiluninescence | 0.1 | 0.5 | mg/L
Inorganic | Nitrogen- Total (Direct Measurement) | ASTM D8083-16 | Combustion-Chemiluminescence | 0.05 | 0.2 | mg/L
Inorganic | Nitrogen- Total (Direct Measurement) | ASTM D8001-16e1 | Persulfate-IC | 0.1 | 0.5 | mg/L
Inorganic | Nitrogen- Total (Direct Measurement) | SM 4500-N B-17 | Persulfate/UV-Detection Various | 0.1 | 0.5 | mg/L
Inorganic | Nitrogen- Total (Direct Measurement) | SM 4500-N C-17 | Persulfate-Detection Various | 0.1 | 0.5 | mg/L | SFWMD MDL=0.02 mg/L (colorimetric)
Inorganic | Nitrogen- Total (Direct Measurement) | USGS I-4650-03 | Persulfate-COLOR-Auto | 0.015 | 0.03 | mg/L
Inorganic | Nitrogen- Total Kjeldahl | EPA 350.1 Rev 2.0 (1993) | COLOR-Auto | 0.01 | 0.04 | mg/L
Inorganic | Nitrogen- Total Kjeldahl | SM 4500-NH3 G-11 or H-11 | COLOR-Auto | 0.01 | 0.04 | mg/L
Inorganic | Nitrogen- Total Kjeldahl | EPA 351.1, Rev. 1978 | COLOR-Auto | 0.05 | 0.2 | mg/L
Inorganic | Nitrogen- Total Kjeldahl | USGS I-4551-78 | COLOR-Auto | 0.05 | 0.2 | mg/L
Inorganic | Nitrogen- Total Kjeldahl | ASTM D3590­ 11 (B) | COLOR-Auto | 0.1 | 0.4 | mg/L
Inorganic | Nitrogen- Total Kjeldahl | EPA 351.2, Rev. 2.0 (1993) | COLOR-Auto | 0.1 | 0.4 | mg/L
Inorganic | Nitrogen- Total Kjeldahl | USGS I-4515-91 | COLOR-Auto | 0.034 | 0.14 | mg/L
Inorganic | Nitrogen- Total Kjeldahl | SM 4500-Norg D-11 | COLOR-Auto | 0.034 | 0.14 | mg/L
Inorganic | Nitrogen- Total Kjeldahl | OI Analytical, Method PAI-DK02 (1994) | COLOR-Auto | 0.1 | 0.4 | mg/L
Inorganic | Nitrogen- Total Kjeldahl | NCASI TNTP-W10900, June 2011 | COLOR-Auto or Manual | 0.05 | 0.2 | mg/L
Inorganic | Nitrogen- Total Kjeldahl | OI Analytical, Method PAI-DK03 (1994) | DIFFUSION-COLOR-Auto | 0.02 | 0.08 | mg/L
Inorganic | Nitrogen- Total Kjeldahl | Timberline Instruments, Method Ammonia-001 (2011) | DIFFUSION-COND-Auto | 0.02 | 0.08 | mg/L
Inorganic | Nitrogen- Total Kjeldahl | SM 4500-NH3 D-11 or E-11 | ISE | 0.5 | 2 | mg/L
Inorganic | Nitrogen- Total Kjeldahl | ASTM 1426-08 (B) | ISE | 0.5 | 2 | mg/L
Inorganic | Nitrogen- Total Kjeldahl | ASTM 1426-08 (A) | NESSLER | 0.05 | 0.2 | mg/L
Inorganic | Nitrogen- Total Kjeldahl | SM 4500-NH3 F-11 | SPEC | 0.01 | 0.04 | mg/L
Inorganic | Nitrogen- Total Kjeldahl | Hach 10242, Rev. 1.1, January 2013 | SPEC-Auto | 0.43 | 1.7 | mg/L
Inorganic | Nitrogen- Total Kjeldahl | AOAC 973.48 (1998) | TITR* | 0.05 | 0.2 | mg/L
Inorganic | Nitrogen- Total Kjeldahl | SM 4500-NH3 C-11 | TITR* | 0.05 | 0.2 | mg/L
Inorganic | Nitrogen- Total Kjeldahl | OI Analytical, Method PAI-DK01 (1994) | TITR-Auto | 0.05 | 0.2 | mg/L
Non-Pesticide Organics | N-Nitrosodimethylamine | EPA 625.1 | GC-MS | 3 | 12 | ug/L
Non-Pesticide Organics | N-Nitrosodimethylamine | SM 6410 B-2000 | GC-MS | 2 | 8 | ug/L
Non-Pesticide Organics | N-Nitrosodimethylamine | USGS 0-3116-87 p.27 | GC-MS | 2 | 8 | ug/L
Non-Pesticide Organics | N-Nitrosodimethylamine | EPA 607 | GC-NPD | 3 | 10 | ug/L
Non-Pesticide Organics | N-Nitrosodi-n-propylamine | EPA 1625B | GC-MS | 2 | 8 | ug/L
Non-Pesticide Organics | N-Nitrosodi-n-propylamine | EPA 625.1 | GC-MS | 2 | 8 | ug/L
Non-Pesticide Organics | N-Nitrosodi-n-propylamine | USGS 0-3116-87 p.27 | GC-MS | 2 | 8 | ug/L
Non-Pesticide Organics | N-Nitrosodi-n-propylamine | SM 6410 B-2000 | GC-MS | 2 | 8 | ug/L
Non-Pesticide Organics | N-Nitrosodi-n-propylamine | EPA 607 | GC-NPD | 1.9 | 7.6 | ug/L
Non-Pesticide Organics | N-Nitrosodiphenylamine | EPA 625.1 | GC-MS | 1.9 | 7.6 | ug/L
Non-Pesticide Organics | N-Nitrosodiphenylamine | EPA 1625B | GC-MS | 1.9 | 7.6 | ug/L
Non-Pesticide Organics | N-Nitrosodiphenylamine | SM 6410 B-2000 | GC-MS | 3 | 10 | ug/L
Non-Pesticide Organics | N-Nitrosodiphenylamine | USGS 0-3116-87 p.27 | GC-MS | 3 | 12 | ug/L
Non-Pesticide Organics | Nonylphenol Diethoxylate (NP2EO) | ASTM D7065-11 | GC-MS | 1.8 | 7.2 | ug/L
Non-Pesticide Organics | Nonylphenol Monoethoxylate (NP1EO) | ASTM D7065-11 | GC-MS | 1.2 | 4.8 | ug/L
Inorganic | Odor | SM 2150 B | GM* | 1 | 4 | TON
Inorganic | Oil & Grease | SM 5520 B-2011 and 5520 F-2011 | GRAV | 1 | 4 | mg/L
Inorganic | Oil & Grease | SM 5520 B 2011 | GRAV | 1 | 4 | mg/L
Inorganic | Oil & Grease (HEM) | EPA 1664 Rev. A | GRAV | 1.4 | 5.6 | mg/L
Inorganic | Oil & Grease (HEM) | EPA 1664 Rev. B | GRAV | 1.4 | 5.6 | mg/L
Inorganic | Orthophosphate (P) | ASTM D6508-10 (D6508, Rev.2) | CIE-UV | 0.1 | 0.4 | mg/L
Inorganic | Orthophosphate (P) | SM 4140- B-2011 | CIE-UV | 0.1 | 0.4 | mg/L
Inorganic | Orthophosphate (P) | AOAC 973.56 | COLOR-Auto | 0.01 | 0.04 | mg/L
Inorganic | Orthophosphate (P) | EPA 365.1, Rev. 2.0 (1993) | COLOR-Auto | 0.01 | 0.04 | mg/L
Inorganic | Orthophosphate (P) | SM 4500-P F-2011 or G-2011 | COLOR-Auto | 0.01 | 0.04 | mg/L
Inorganic | Orthophosphate (P) | USGS I-4601-85 | COLOR-Auto | 0.01 | 0.04 | mg/L
Inorganic | Orthophosphate (P) | AOAC 993.30 | IC | 0.003 | 0.01 | mg/L
Inorganic | Orthophosphate (P) | ASTM D4327-03 | IC | 0.003 | 0.01 | mg/L
Inorganic | Orthophosphate (P) | EPA 300.0, Rev 2.1 (1993) | IC | 0.003 | 0.1 | mg/L
Inorganic | Orthophosphate (P) | EPA 300.1, Rev 1.0 (1997) | IC | 0.019 | 0.076 | mg/L
Inorganic | Orthophosphate (P) | SM 4110 B-2011 | IC | 0.003 | 0.01 | mg/L
Inorganic | Orthophosphate (P) | SM 4110 C-2011 | IC | 0.004 | 0.016 | mg/L
Inorganic | Orthophosphate (P) | AOAC 973.55 | SPEC | 0.01 | 0.04 | mg/L
Inorganic | Orthophosphate (P) | ASTM D515-88(A) | SPEC | 0.01 | 0.04 | mg/L
Inorganic | Orthophosphate (P) | EPA 365.3 (Issued 1978) | SPEC | 0.01 | 0.04 | mg/L
Inorganic | Orthophosphate (P) | SM 4500-P E -2011 | SPEC | 0.01 | 0.04 | mg/L
Inorganic | Palladium | AES0029, Rev. 1991 Thermo Jarrell Ash Corp. | DCP | 5 | 20 | ug/L
Inorganic | Palladium | SM 3111B-2011 | FLAAS | 100 | 400 | ug/L
Inorganic | Palladium | EPA 253.2 | GFAA | 5 | 20 | ug/L
Inorganic | Palladium | SM 3125 B-2011 | ICP/MS | 5 | 20 | ug/L
Pesticide | Parathion ethyl | USGS, TWRI, Book 5, Ch. A3, P. 27 | GC | 0.3 | 1.2 | ug/L
Pesticide | Parathion ethyl | SM 6630 B-2007 | GC-ECD | 0.1 | 0.4 | ug/L
Pesticide | Parathion ethyl | EPA 614 | GC-FPD | 0.012 | 0.048 | ug/L
Pesticide | Parathion ethyl | EPA, 1978, P. 25 | GC-FPD | 0.05 | 0.2 | ug/L
Pesticide | Parathion ethyl | USGS O-1126-95 | GC-MS-SIM | 0.022 | 0.088 | ug/L
Pesticide | Parathion methyl | USGS, TWRI, Book 5, Ch. A3, P. 27 | GC | 0.3 | 1.2 | ug/L
Pesticide | Parathion methyl | SM 6630 B-2007 | GC-ECD | 0.1 | 0.4 | ug/L
Pesticide | Parathion methyl | EPA 614 | GC-FPD | 0.012 | 0.048 | ug/L
Pesticide | Parathion methyl | EPA 1657 | GC-FPD | 0.018 | 0.072 | ug/L
Pesticide | Parathion methyl | EPA, 1978, P. 25 | GC-FPD | 0.05 | 0.2 | ug/L
Pesticide | Parathion methyl | EPA 622 | GC-FPD or TBD | 0.3 | 1.2 | ug/L
Pesticide | Parathion methyl | EPA 625.1 | GC-MS | 0.1 | 0.4 | ug/L
pesticide | Parathion methyl | USGS O-1126-95 | GC-MS-SIM | 0.035 | 0.14 | ug/L
Pesticide | Parathion methyl | EPA 614 | GC-NPD | 0.1 | 0.4 | ug/L
Non-Pesticide Organics | PCB- 1016 | 3M Corp 3M0222 (10/28/94) | Empore Disk/ GC-ECD | 0.26 | 1.04 | ug/L
Non-Pesticide Organics | PCB- 1016 | EPA, 1978, p. 43 | GC | 0.1 | 0.4 | ug/L
Non-Pesticide Organics | PCB- 1016 | EPA 608.3 | GC-ECD | 0.1 | 0.4 | ug/L
Non-Pesticide Organics | PCB- 1016 | EPA 625.1 | GC-MS | 40 | 160 | ug/L
Non-Pesticide Organics | PCB- 1016 | SM 6410 B-2000 | GC-MS | 40 | 160 | ug/L
Non-Pesticide Organics | PCB- 1221 | 3M Corp 3M0222 (10/28/94) | Empore Disk/ GC-ECD | 0.26 | 1.04 | ug/L
Non-Pesticide Organics | PCB- 1221 | EPA, 1978, p. 43 | GC | 0.1 | 0.4 | ug/L
Non-Pesticide Organics | PCB- 1221 | EPA 608.3 | GC-ECD | 0.1 | 0.4 | ug/L
Non-Pesticide Organics | PCB- 1221 | EPA 625.1 | GC-MS | 40 | 160 | ug/L
Non-Pesticide Organics | PCB- 1221 | SM 6410 B-2000 | GC-MS | 40 | 160 | ug/L
Non-Pesticide Organics | PCB- 1232 | 3M Corp 3M0222 (10/28/94) | Empore Disk/ GC-ECD | 0.26 | 1.04 | ug/L
Non-Pesticide Organics | PCB- 1232 | EPA, 1978, p. 43 | GC | 0.1 | 0.4 | ug/L
Non-Pesticide Organics | PCB- 1232 | EPA 608.3 | GC-ECD | 0.1 | 0.4 | ug/L
Non-Pesticide Organics | PCB- 1232 | EPA 625.1 | GC-MS | 40 | 160 | ug/L
Non-Pesticide Organics | PCB- 1232 | SM 6410 B-2000 | GC-MS | 40 | 160 | ug/L
Non-Pesticide Organics | PCB- 1242 | 3M Corp 3M0222 (10/28/94) | Empore Disk/ GC-ECD | 0.26 | 1.04 | ug/L
Non-Pesticide Organics | PCB- 1242 | EPA, 1978, p. 43 | GC | 0.1 | 0.4 | ug/L
Non-Pesticide Organics | PCB- 1242 | EPA 608.3 | GC-ECD | 0.1 | 0.4 | ug/L
Non-Pesticide Organics | PCB- 1242 | EPA 625.1 | GC-MS | 40 | 160 | ug/L
Non-Pesticide Organics | PCB- 1242 | SM 6410 B-2000 | GC-MS | 40 | 160 | ug/L
Non-Pesticide Organics | PCB- 1248 | 3M Corp 3M0222 (10/28/94) | Empore Disk/ GC-ECD | 0.26 | 1.04 | ug/L
Non-Pesticide Organics | PCB- 1248 | EPA, 1978, p. 43 | GC | 0.1 | 0.4 | ug/L
Non-Pesticide Organics | PCB- 1248 | EPA 608.3 | GC-ECD | 0.1 | 0.4 | ug/L
Non-Pesticide Organics | PCB- 1248 | EPA 625.1 | GC-MS | 40 | 160 | ug/L
Non-Pesticide Organics | PCB- 1248 | SM 6410 B-2000 | GC-MS | 40 | 160 | ug/L
Non-Pesticide Organics | PCB- 1254 | 3M Corp 3M0222 (10/28/94) | Empore Disk/ GC-ECD | 0.26 | 1.04 | ug/L
Non-Pesticide Organics | PCB- 1254 | EPA, 1978, p. 43 | GC | 0.1 | 0.4 | ug/L
Non-Pesticide Organics | PCB- 1254 | EPA 608.3 | GC-ECD | 0.1 | 0.4 | ug/L
Non-Pesticide Organics | PCB- 1254 | EPA 625.1 | GC-MS | 40 | 160 | ug/L
Non-Pesticide Organics | PCB- 1254 | SM 6410 B-2000 | GC-MS | 40 | 160 | ug/L
Non-Pesticide Organics | PCB- 1260 | 3M Corp 3M0222 (10/28/94) | Empore Disk/ GC-ECD | 0.26 | 1.04 | ug/L
Non-Pesticide Organics | PCB- 1260 | EPA, 1978, p. 43 | GC | 0.1 | 0.4 | ug/L
Non-Pesticide Organics | PCB- 1260 | EPA 608.3 | GC-ECD | 0.1 | 0.4 | ug/L
Non-Pesticide Organics | PCB- 1260 | EPA 625.1 | GC-MS | 40 | 160 | ug/L
Non-Pesticide Organics | PCB- 1260 | SM 6410 B-2000 | GC-MS | 40 | 160 | ug/L
Non-Pesticide Organics | Pentachlorophenol | EPA, 1978, p.140 | GC | 3 | 10 | ug/L
Non-Pesticide Organics | Pentachlorophenol | EPA 604 | GC-FID | 3 | 10 | ug/L
Non-Pesticide Organics | Pentachlorophenol | SM 6420 B-2000 | GC-FID | 3 | 10 | ug/L
Non-Pesticide Organics | Pentachlorophenol | EPA 8041 A | GC-FID or ECD | 3 | 12 | ug/L
Non-Pesticide Organics | Pentachlorophenol | EPA 1625B | GC-MS | 3 | 10 | ug/L
Non-Pesticide Organics | Pentachlorophenol | EPA 625.1 | GC-MS | 3.6 | 14 | ug/L
Non-Pesticide Organics | Pentachlorophenol | SM 6410 B-2000 | GC-MS | 3 | 10 | ug/L
Non-Pesticide Organics | Pentachlorophenol | USGS 0-3116-87 p.27 | GC-MS | 3 | 10 | ug/L
Inorganic | pH | USGS I-2587-85 | Automated electrode* | 0.1 | 0.1 | SU
Inorganic | pH | AOAC 973.41 | GM* | 0.1 | 0.1 | SU
Inorganic | pH | ASTM D1293-99 (A or B) | GM* | 0.1 | 0.1 | SU
Inorganic | pH | EPA 150.2 | GM* | 0.1 | 0.1 | SU
Inorganic | pH | SM 4500-H+B 2011 | GM* | 0.1 | 0.1 | SU
Inorganic | pH | USGS I-1586-85 | GM* | 0.1 | 0.1 | SU
Inorganic | pH | DEP-SOP FT 1100 | ISE* | 0.1 | 0.1 | SU
Inorganic | pH | IM21 378-75 WA (Bran Luebbe) | ISE* | 0.1 | 0.1 | SU
Non-Pesticide Organics | Phenanthrene | EPA 610-GC | GC-FID | 1 | 4 | ug/L
Non-Pesticide Organics | Phenanthrene | EPA 1625B | GC-MS | 0.05 | 0.2 | ug/L
Non-Pesticide Organics | Phenanthrene | EPA 625.1 | GC-MS | 0.05 | 0.2 | ug/L
Non-Pesticide Organics | Phenanthrene | SM 6410 B-2000 | GC-MS | 0.05 | 0.2 | ug/L
Non-Pesticide Organics | Phenanthrene | USGS 0-3116-87 p.27 | GC-MS | 0.05 | 0.2 | ug/L
Non-Pesticide Organics | Phenanthrene | ASTM D4657-92 (98) | HPLC-UV | 0.05 | 0.2 | ug/L
Non-Pesticide Organics | Phenanthrene | EPA 610-HPLC | HPLC-UV-FLUOR | 0.05 | 0.2 | ug/L
Non-Pesticide Organics | Phenanthrene | SM 6440 B-2005-HPLC | HPLC-UV-FLUOR | 0.05 | 0.2 | ug/L
Non-Pesticide Organics | Phenol | EPA 604 | GC-FID | 1 | 4 | ug/L
Non-Pesticide Organics | Phenol | SM 6420 B-2000 | GC-FID | 1 | 4 | ug/L
Non-Pesticide Organics | Phenol | EPA 1625B | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | Phenol | EPA 625.1 | GC-MS | 1.5 | 6.0 | ug/L
Non-Pesticide Organics | Phenol | SM 6410 B-2000 | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | Phenol | USGS 0-3116-87 p.27 | GC-MS | 1 | 4 | ug/L
Non-Pesticide Organics | Phenolics, Chlorinated | EPA 1653 | GC-MS | 1 | 4 | ug/L
Inorganic | Phenols | SM 5530D-2010 | COLOR | 5 | 20 | ug/L
Inorganic | Phenols | ASTM D1783-01 (A or B) | COLOR | 5 | 20 | ug/L
Inorganic | Phenols | EPA 420.1, Rev. 1978 | COLOR | 5 | 20 | ug/L
Inorganic | Phenols | EPA 420.4, Rev. 1.0 (1993) | COLOR-Auto | 5 | 20 | ug/L
Inorganic | Phosphorus (elemental) mg/L | Addison and Ackman 1970 | GC/LC | 1 | 4 | ug/L
Inorganic | Phosphorus- Total | AOAC 973.56 (Phosphorous -Total) | COLOR-Auto | 0.01 | 0.04 | mg/L
Inorganic | Phosphorus- Total | ASTM D515-88 (B) | COLOR-Auto | 0.01 | 0.04 | mg/L
Inorganic | Phosphorus- Total | EPA 365.1 Rev. 2.0 (1993) | COLOR-Auto | 0.01 | 0.04 | mg/L
Inorganic | Phosphorus- Total | EPA 365.4 (Issued 1974) | COLOR-Auto | 0.01 | 0.04 | mg/L
Inorganic | Phosphorus- Total | SM 4500-P F, G, or H (2011) | COLOR-Auto | 0.01 | 0.04 | mg/L
Inorganic | Phosphorus- Total | USGS I-4600-85 | COLOR-Auto | 0.01 | 0.04 | mg/L
Inorganic | Phosphorus- Total | USGS I-4610-91 | COLOR-Auto | 0.01 | 0.04 | mg/L
Inorganic | Phosphorus- Total | NCASI TNTP-W10900, June 2011 | COLOR-Auto or Manual | 0.01 | 0.04 | mg/L
Inorganic | Phosphorus- Total | EPA 200.7 Rev 4.4 (1994) | ICP-AES | 0.06 | 0.2 | mg/L
Inorganic | Phosphorus- Total | USGS I-4471-97 | ICP-AES | 0.06 | 0.2 | mg/L
Inorganic | Phosphorus- Total | SM 3120 B-11 | ICP-ES | 0.06 | 0.2 | mg/L
Inorganic | Phosphorus- Total | ASTM D515-88 (A) (Phosphorous -Total) | SPEC | 0.01 | 0.04 | mg/L
Inorganic | Phosphorus- Total | EPA 365.3 (Issued 1978) | SPEC | 0.01 | 0.04 | mg/L
Inorganic | Phosphorus- Total | SM 4500-P E -2011 | SPEC | 0.01 | 0.04 | mg/L
Inorganic | Platinum | AES0029, Rev. 1991 Thermo Jarrell Ash Corp. | DCP-AES | 0.2 | 0.8 | mg/L
Inorganic | Platinum | SM 3111 B-2011 | FLAAS | 2 | 8 | mg/L
Inorganic | Platinum | EPA 255.2 (Issued 1978) | GFAAS | 0.02 | 0.08 | mg/L
Inorganic | Platinum | SM 3125 B-2011 | ICP/MS | 0.02 | 0.08 | mg/L
Non-Pesticide Organics | p-tert-Octylphenol (OP) | ASTM D7065-11 | GC-MS | 0.2 | 0.8 | ug/L
Non-Pesticide Organics | Pyrene | EPA 610 | GC-FID | 1 | 4 | ug/L
Non-Pesticide Organics | Pyrene | SM 6440 B 2005 | GC-FID | 1 | 4 | ug/L
Non-Pesticide Organics | Pyrene | EPA 1625B | GC-MS | 0.05 | 0.2 | ug/L
Non-Pesticide Organics | Pyrene | EPA 625.1 | GC-MS | 0.05 | 0.2 | ug/L
Non-Pesticide Organics | Pyrene | SM 6410 B-2000 | GC-MS | 0.05 | 0.2 | ug/L
Non-Pesticide Organics | Pyrene | USGS 0-3116-87 p.27 | GC-MS | 0.05 | 0.2 | ug/L
Non-Pesticide Organics | Pyrene | ASTM D4657-92 | HPLC-UV | 0.5 | 2 | ug/L
Non-Pesticide Organics | Pyrene | EPA 610 | HPLC-UV-FLUOR | 0.5 | 2 | ug/L
Non-Pesticide Organics | Pyrene | SM 6440 B-2005 | HPLC-UV-FLUOR | 0.5 | 2 | ug/L
Radiologic | Radium | EPA 903.1 | A-SCINT* | 0.5 | 2 | pCi/L
Radiologic | Radium | SM 7500-Ra C | A-SCINT* | 0.5 | 2 | pCi/L
Radiologic | Radium | SM 7500-Ra C-01 | A-SCINT* | 0.5 | 2 | pCi/L
Radiologic | Radium | ASTM D3454-91, 97 | A-SCINT* | 0.5 | 2 | pCi/L
Radiologic | Radium | USGS 76-177 (1976), p.81 | A-SCINT* | 0.5 | 2 | pCi/L
Radiologic | Radium - Total | EPA 903.0 | A-GPC* | 1 | 4 | pCi/L
Radiologic | Radium - Total | SM 7500-Ra B | A-GPC* | 1 | 4 | pCi/L
Radiologic | Radium - Total | SM 7500-Ra B-01 | A-GPC* | 1 | 4 | pCi/L
Radiologic | Radium - Total | ASTM D2460-90, 97 | A-GPC* | 1 | 4 | pCi/L
Radiologic | Radium-226 | ASTM D3454-91 | A-SCINT* | 0.5 | 2 | pCi/L
Radiologic | Radium-226 | DOE Ra-05 | A-SCINT* | 1 | 4 | pCi/L
Radiologic | Radium-226 | EPA 903.0 (Alpha) | A-SCINT* | 1 | 4 | pCi/L
Radiologic | Radium-226 | EPA 903.1 | A-SCINT* | 0.5 | 2 | pCi/L
Radiologic | Radium-226 | EPA Ra-04 (Ra-226) | A-SCINT* | 1 | 4 | pCi/L
Radiologic | Radium-228 | EPA Ra-05 (Ra-228) | A-GPC* | 1 | 4 | pCi/L
Radiologic | Radium-228 | Brooks-Blanchard | GAS-FLOW-P* | 1 | 4 | pCi/L
Radiologic | Radium-228 | EPA 904 (Beta) | GAS-FLOW-P* | 1 | 4 | pCi/L
Radiologic | Radium-228 | EPA 9320 (Beta) | GAS-FLOW-P* | 1 | 4 | pCi/L
Inorganic | Residue ­Volatile, mg/L | EPA 160.4 (Issued 1971) | GRAV* | 1 | 4 | mg/L
Inorganic | Residue ­Volatile, mg/L | USGS I-3753-85 | GRAV* | 1 | 4 | mg/L
Inorganic | Residues- Filterable (TDS) | USGS I-1750-85 | GRAV* | 10 | 40 | mg/L
Inorganic | Residues- Filterable (TDS) | ASTM 5907-13 | GRAV* | 10 | 40 | mg/L
Inorganic | Residues- Filterable (TDS) | SM 2540 C-2011 | GRAV* | 10 | 40 | mg/L
Inorganic | Residues- Nonfilterable (TSS) | ASTM5907-13 | GRAV* | 1 | 4 | mg/L
Inorganic | Residues- Nonfilterable (TSS) | SM 2540 D-2011 | GRAV* | 1 | 4 | mg/L
Inorganic | Residues- Nonfilterable (TSS) | USGS I-3765-85 | GRAV* | 1 | 4 | mg/L
Inorganic | Residues- settleable | SM 2540 F-2011 | GRAV* | 1 | 4 | mg/L
Inorganic | Residues- total mg/L | SM2540B-2011 | GRAV* | 1 | 4 | mg/L
Inorganic | Residues- total mg/L | USGS I-3750-85 | GRAV* | 1 | 4 | mg/L
Inorganic | Residues- volatile, mg/L | SM 2540 E-2011 | GRAV* | 1 | 4 | mg/L
Inorganic | Selenium | ASTM D3859-08(B) | AA | 0.6 | 2 | ug/L
Inorganic | Selenium | ASTM D3859-08(B) | AA | 0.6 | 2.4 | ug/L
Inorganic | Selenium | USGS I­4668-98 | AA | 0.6 | 2.4 | ug/L
Inorganic | Selenium | ASTM D3859-08(B) | AA | 0.6 | 2.4 | ug/L
Inorganic | Selenium | SM 3113 B-10 | FLAAS | 0.6 | 2.4 | ug/L
Inorganic | Selenium | USGS I­3667­85 | FLAAS-HYDRIDE | 2 | 8 | ug/L
Inorganic | Selenium | SM 3114 B -2011 or C-2011 (Se) | FLAAS-HYDRIDE | 2 | 8 | ug/L
Inorganic | Selenium | ASTM D3859­-08 (A) | Hydride | 2 | 8 | ug/L
Inorganic | Selenium | ASTM D3859-08(A) | Hydride | 2 | 8 | ug/L
Inorganic | Selenium | EPA 200.7, Rev. 4.4 (1994) | ICP | 0.02 | 0.08 | mg/L
Inorganic | Selenium | EPA 200.8, Rev. 5.4 (1994) | ICP /MS | 7.9 | 31.6 | ug/L
Inorganic | Selenium | EPA 200.5, Rev. 4.2 (2003) (68) | ICP/AES | 1.3 | 5.2 | ug/L
Inorganic | Selenium | ASTM D1976-12 | ICP/AES | 75 | 300 | ug/L
Inorganic | Selenium | EPA 200.7 | ICP/AES | 0.02 | 0.08 | mg/L
Inorganic | Selenium | AOAC 993.14 | ICP/MS | 0.8 | 3.2 | ug/L
Inorganic | Selenium | SM 3125 B-2011 | ICP/MS | 7.9 | 31.6 | ug/L
Inorganic | Selenium | ASTM D5673­10 | ICP/MS | 7.9 | 32 | ug/L
Inorganic | Selenium | SM 3120 B-11 | ICP-ES | 8 | 32 | ug/L
Inorganic | Selenium | SM 3120 B-11 | ICP-ES | 2 | 8 | ug/L
Inorganic | Selenium | EPA 200.9, Rev. 2.2 (1994) | STGFAA | 0.6 | 2.4 | ug/L
Inorganic | Selenium | EPA 200.9 | TSGF | 2 | 8 | ug/L
Inorganic | Silver | USGS I­4724­89 | AA | 0.2 | 0.8 | ug/L
Inorganic | Silver | AES0029, Rev. 1991 Thermo Jarrell Ash Corp. | DCP | 0.2 | 0.8 | ug/L
Inorganic | Silver | ANS P. 37 (Silver) | FLAAS | 10 | 40 | ug/L
Inorganic | Silver | AOAC 974.27 (Silver) | FLAAS | 0.06 | 0.2 | ug/L
Inorganic | Silver | SM 3111 B-2011 (Ag) | FLAAS | 0.06 | 0.2 | ug/L
Inorganic | Silver | SM 3111 C-2011 (Ag) | FLAAS | 0.06 | 0.2 | ug/L
Inorganic | Silver | SM 3113 B-10 | FLAAS | 0.2 | 0.8 | ug/L
Inorganic | Silver | USGS I-3720-85 | FLAAS | 0.06 | 0.2 | ug/L
Inorganic | Silver | ASTM D5673­10 | ICP//MS | 0.1 | 0.4 | ug/L
Inorganic | Silver | ASTM D1976-12 | ICP/AES | 7 | 30 | ug/L
Inorganic | Silver | EPA 200.7, Rev. 4.4 (1994) | ICP/AES | 0.2 | 0.8 | ug/L
Inorganic | Silver | EPA 200.5, Rev 4.2 (2003) | ICP/AES | 0.2 | 0.8 | ug/L
Inorganic | Silver | USGS I-4471-97 | ICP/AES or MS | 0.2 | 0.8 | ug/L
Inorganic | Silver | AOAC 993.14 | ICP/MS | 0.1 | 0.4 | ug/L
Inorganic | Silver | SM 3125 B-2011 | ICP/MS | 0.1 | 0.4 | ug/L
Inorganic | Silver | SM 3120 B-11 | ICP-ES | 1 | 4 | ug/L
Inorganic | Silver | EPA 200.8, Rev. 5.4 (1994) | ICP-MS | 0.1 | 0.4 | ug/L
Inorganic | Silver | 200.9, Rev. 2.2 (1994) | STGFAA | 0.5 | 2 | ug/L
Inorganic | Specific Conductance | AOAC 973.40 | COND* | 1 | 4 | umhos/cm
Inorganic | Specific Conductance | ASTM D1125­95 (99) (A) | COND* | 1 | 1 | umhos/cm
Inorganic | Specific Conductance | DEP-SOP FT 1200 | COND* | 1 | 1 | umhos/cm
Inorganic | Specific Conductance | EPA 120.1 (Rev. 1982) | COND* | 1 | 1 | umhos/cm
Inorganic | Specific Conductance | SM 2510 B-2011 | COND* | 1 | 1 | umhos/cm
Inorganic | Sulfate | AOAC D6508, Rev 2 | CIE | 0.1 | 0.4 | mg/L
Inorganic | Sulfate | SM 4140 B-2011 | CIE/UV | 0.1 | 0.4 | mg/L
Inorganic | Sulfate | ASTM D6508-1010 | CIE/UV | 0.1 | 0.4 | mg/L
Inorganic | Sulfate | EPA 375.2, Rev. 2.0 (1993) | COLOR | 0.5 | 2 | mg/L
Inorganic | Sulfate | SM 4500-SO4 F-2011 or G-2011 | COLOR-Auto | 0.5 | 2 | mg/L
Inorganic | Sulfate | SM 4500-SO4 C-2011 or D-2011 | GRAV | 10 | 40 | mg/L
Inorganic | Sulfate | ASTM 925.54 | GRAV | 10 | 40 | mg/L
Inorganic | Sulfate | SM 4110B-2011 or C-2011 | IC | 0.35 | 1.4 | mg/L
Inorganic | Sulfate | EPA 300.0, Rev 2.1 (1993) | IC | 0.02 | 0.08 | mg/L
inorganic | Sulfate | 300.1-1, Rev 1.0 (1997) | IC | 0.019 | 0.076 | mg/L
inorganic | Sulfate | ASTM D4327 -03 | IC | 0.02 | 0.08 | mg/L
Inorganic | Sulfate | AOAC 993.30 | IC | 0.02 | 0.08 | mg/L
Inorganic | Sulfate | USGS I-4020-05 | IC | 0.02 | 0.08 | mg/L
Inorganic | Sulfate | SM 4500-SO4 E-2011 | TURBID | 1 | 4 | mg/L
inorganic | Sulfate | ASTM D516-11 | TURBID | 1 | 4 | mg/L
Inorganic | Sulfide | SM 4500-S-2 D-2011 | COLOR | 0.1 | 0.4 | mg/L
Inorganic | Sulfide | SM 4500-S2-G-2011 | ISE | 0.04 | 0.16 | mg/L
Inorganic | Sulfide | ASTM D4658-09 | ISE | 0.04 | 0.16 | mgL
Inorganic | Sulfide | SM 4500-S-2 F-2011 | TITR* | 1 | 4 | mg/L
Inorganic | Sulfide | USGS I-3840-85 | TITR* | 1 | 4 | mg/L
Inorganic | Sulfite | SM 4500-SO3 ² B-2011 | TITR* | 1 | 4 | mg/L
Inorganic | Surfactants | ASTM D2330-02 | COLOR | 25 | 100 | ug/L
Inorganic | Surfactants | SM 5540 C-2011 | COLOR | 25 | 100 | ug/L
Inorganic | Temperature | DEP-SOP FT 1400 | TEMP* | 0.1 | 0.1 | degree C
Inorganic | Temperature | USGS Bk 1, Ch. D1, 1975 | TEMP* | 0.1 | 0.1 | degree C
Inorganic | Temperature | SM 2550 B-2010 | TEMP* | 0.1 | 0.1 | degree C
Non-Pesticide Organics | Tetrachloroethene | SM 6200 C-2011 | GC | 0.2 | 0.8 | ug/L
Non-Pesticide Organics | Tetrachloroethene | EPA, 1978, p. 130 | GC | 0.2 | 0.8 | ug/L
Non-Pesticide Organics | Tetrachloroethene | EPA 601 | GC-ELCD | 0.2 | 0.8 | ug/L
Non-Pesticide Organics | Tetrachloroethene | EPA 1624B | GC-MS | 0.2 | 0.8 | ug/L
Non-Pesticide Organics | Tetrachloroethene | EPA 624.1 | GC-MS | 0.2 | 0.8 | ug/L
Non-Pesticide Organics | Tetrachloroethene | SM 6200 B-2011 | GC-MS | 0.2 | 0.8 | ug/L
Inorganic | Thallium | SM 3111 B (Tl)-2011 | FLAAS | 100 | 400 | ug/L
Inorganic | Thallium | SM 3113 B-10 | GFAAS | 1 | 4 | ug/L
Inorganic | Thallium | EPA 279.2 (issued 1978) | GFAAS | 1 | 4 | ug/L
Inorganic | Thallium | EPA 200.7, Rev. 4.4 (1994) | ICP-AES | 0.001 | 0.004 | mg/L
Inorganic | Thallium | ASTM D1976-12 | ICP-AES | 40 | 160 | ug/L
Inorganic | Thallium | SM 3120 B-11 | ICP-ES | 10 | 40 | ug/L
Inorganic | Thallium | SM 3125 B-2011 | ICP-MS | 0.3 | 1.2 | ug/L
Inorganic | Thallium | 200.8, Rev. 5.4 (1994) | ICP-MS | 0.3 | 1.2 | ug/L
Inorganic | Thallium | ASTM D5673-10 | ICP-MS | 0.3 | 1.2 | ug/L
Inorganic | Thallium | AOAC 993.14 | ICP-MS | 0.8 | 3.2 | ug/L
Inorganic | Thallium | USGS 1-4471-97 | ICP-MS | 0.2 | 1.2 | ug/L
Inorganic | Thallium | 200.9, Rev. 2.2 (1994) | TSGF | 0.7 | 2.8 | ug/L
Non-Pesticide Organics | Toluene | EPA 1624B | GC-MS | 0.5 | 2 | ug/L
Non-Pesticide Organics | Toluene | EPA 624.1 | GC-MS | 0.5 | 2 | ug/L
Non-Pesticide Organics | Toluene | SM 6200 B-2011 | GC-MS | 0.5 | 2 | ug/L
Non-Pesticide Organics | Toluene | EPA 602 | GC-PID | 0.2 | 0.8 | ug/L
Non-Pesticide Organics | Toluene | SM 6200 C-2011 | GC-PID | 0.5 | 2 | ug/L
Pesticide | Toxaphene | ASTM D3086-90 | GC-ECD | 0.8 | 3.2 | ug/L
Pesticide | Toxaphene | EPA 608.3 | GC-ECD | 0.5 | 2 | ug/L
Pesticide | Toxaphene | SM 6630 B-2007 | GC-ECD | 0.8 | 3.2 | ug/L
Pesticide | Toxaphene | SM 6630 C-2007 | GC-ECD | 0.8 | 3.2 | ug/L
Pesticide | Toxaphene | EPA 505 | GC-ECD | 0.8 | 3 | ug/L
Pesticide | Toxaphene | EPA 508 | GC-ECD | 0.8 | 3 | ug/L
Pesticide | Toxaphene | EPA 617 | GC-ECD | 0.8 | 3 | ug/L
Pesticide | Toxaphene | EPA 1656 | GC-ECD | 0.8 | 3 | ug/L
Pesticide | Toxaphene | ASTM D5812-96 | GC-ECD | 0.8 | 3 | ug/L
Pesticide | Toxaphene | EPA, 1978, p. 7 | GC-ECD | 0.8 | 3 | ug/L
Pesticide | Toxaphene | 3M Corp 3M0222 (10/28/94) | GC-ECD | 0.8 | 3 | ug/L
Pesticide | Toxaphene | USGS 0-3105-83 | GC-ECD | 0.8 | 3 | ug/L
Pesticide | Toxaphene | EPA 625.1 | GC-MS | 500 | 2000 | ug/L
Pesticide | Toxaphene | SM 6410 B-2000 | GC-MS | 500 | 2000 | ug/L
Pesticide | Toxaphene | EPA 525.1 | GC-MS | 500 | 2000 | ug/L
Pesticide | Toxaphene | EPA 525.2 | GC-MS | 500 | 2000 | ug/L
Non-Pesticide Organics | trans-1,2-Dichloroethene | SB 6200 C-2011 | GC | 0.2 | 0.8 | ug/L
Non-Pesticide Organics | trans-1,2-Dichloroethene | EPA 601 | GC-ELCD | 0.2 | 0.8 | ug/L
Non-Pesticide Organics | trans-1,2-Dichloroethene | EPA 1624B | GC-MS | 0.5 | 2 | ug/L
Non-Pesticide Organics | trans-1,2-Dichloroethene | EPA 624.1 | GC-MS | 0.5 | 2 | ug/L
Non-Pesticide Organics | trans-1,2-Dichloroethene | SM 6200 B-2011 | GC-MS | 0.2 | 0.8 | ug/L
Non-Pesticide Organics | Trichloroethene | SM 6200 C-2011 | GC-ELCD | 0.2 | 0.8 | ug/L
Non-Pesticide Organics | Trichloroethene | EPA 1624B | GC-MS | 0.2 | 0.8 | ug/L
Non-Pesticide Organics | Trichloroethene | EPA 624.1 | GC-MS | 0.2 | 0.8 | ug/L
Non-Pesticide Organics | Trichloroethene | SM 6200 B-2011 | GC-MS | 0.2 | 0.8 | ug/L
Non-Pesticide Organics | Trichloroethylene (trichloroethene) | EPA 601 | GC-ELCD | 0.2 | 0.8 | ug/L
Non-Pesticide Organics | Trichlorofluoromethane | SM 6200 C-2011 | GC-ELCD | 0.2 | 0.8 | ug/L
Non-Pesticide Organics | Trichlorofluoromethane | EPA 624.1 | GC-MS | 0.5 | 2 | ug/L
Non-Pesticide Organics | Trichlorofluoromethane | SM 6200 B-2011 | GC-MS | 0.2 | 0.8 | ug/L
Inorganic | Turbidity | ASTM D1889-00 | TURBID* | 0.1 | 0.1 | NTU
Inorganic | Turbidity | USGS I-3860-85 | TURBID* | 0.1 | 0.1 | NTU
Inorganic | Turbidity | DEP-SOP FT 1600 | TURBID* | 0.1 | 0.1 | NTU
Inorganic | Turbidity | EPA 180.1, Rev. 2.0 (1993) | TURBID* | 0.1 | 0.1 | NTU
Inorganic | Turbidity | SM 2130 B-2011 | TURBID* | 0.1 | 0.1 | NTU
Inorganic | Turbidity | Mitchell Method M5331 | TURBID* | 0.1 | 0.1 | NTU
Inorganic | Turbidity | Mitchell Method M5271 | TURBID* | 0.1 | 0.1 | NTU
Inorganic | Turbidity | Orion Method AQ4500 | TURBID* | 0.1 | 0.1 | NTU
Non-Pesticide Organics | Vinyl chloride | EPA 601 | GC-ELCD | 0.5 | 2 | ug/L
Non-Pesticide Organics | Vinyl chloride | SM 6200 C-2011 | GC-ELCD | 0.12 | 0.48 | ug/L
Non-Pesticide Organics | Vinyl chloride | SM 6200 B-2011 | GCMS | 0.5 | 2 | ug/L
Non-Pesticide Organics | Vinyl chloride | EPA 1624b | GC-MS | 0.5 | 2 | ug/L
Non-Pesticide Organics | Vinyl chloride | EPA 624.1 | GC-MS | 0.5 | 2 | ug/L
Inorganic | Zinc | ASTM D1691-12 (A or B) | AA | 5 | 20 | ug/L
Inorganic | Zinc | SM 3500-Zn B-2011 | COLOR | 250 | 1000 | ug/L
Inorganic | Zinc | Method 8009 Zincon | COLOR | 250 | 1000 | ug/L
Inorganic | Zinc | ASTM D4190­08 | DCP | 50 | 200 | ug/L
Inorganic | Zinc | AES0029, Rev. 1991 Thermo Jarrell Ash Corp. | DCP | 50 | 200 | ug/L
Inorganic | Zinc | ASTM 4190-08 | DCP | 50 | 200 | ug/L
Inorganic | Zinc | AOAC 974.27 (Zinc) | FLAAS | 5 | 20 | ug/L
Inorganic | Zinc | SM 3111 B -2011(Zn) | FLAAS | 5 | 20 | ug/L
Inorganic | Zinc | SM 3111 C -2011 (Zn) | FLAAS | 5 | 20 | ug/L
Inorganic | Zinc | USGS I-3900-85 | FLAAS | 0.2 | 0.8 | ug/L
Inorganic | Zinc | ANS P. 37 (Zinc) | FLAAS | 5 | 20 | ug/L
Inorganic | Zinc | EPA 289.2 (Issued 1978) | GFAAS | 5 | 20 | ug/L
Inorganic | Zinc | EPA 200.5, Rev 4.2 (2003) (68) | ICP | 0.4 | 1.6 | ug/L
Inorganic | Zinc | EPA 200.7, Rev. 4.4 (1994) | ICP | 0.002 | 0.008 | mg/L
Inorganic | Zinc | ASTM D1976-12 | ICP/AES | 2 | 8 | ug/L
Inorganic | Zinc | USGS I­4471­97 | ICP/AES | 0.4 | 1.6 | ug/L
Inorganic | Zinc | AOAC 993.14 | ICP/MS | 1.8 | 7.2 | ug/L
Inorganic | Zinc | ASTM D5673­10 | ICP/MS | 1.8 | 7.2 | ug/L
Inorganic | Zinc | USGS I-4020-05 | ICP/MS | 1.8 | 7.2 | ug/L
Inorganic | Zinc | SM 3120 B-11 | ICP-ES | 2 | 8 | ug/L
Inorganic | Zinc | SM 3125 B-2011 | ICP-MS | 1.8 | 7.2 | ug/L
Inorganic | Zinc | EPA 200.8, Rev. 5.4 (1994) | ICP-MS | 1.8 | 7.2 | ug/L

## MDL Note
An asterisk (*) in the TechniqueID column indicates an analytical technique for which the EPA procedure for MDL determination may not be appropriate (see 40 CFR, Part 136, Appendix B). The laboratory may instead indicate the PQL or laboratory Reporting Limit as an equivalent detection level for a specified analyte, as determined by alternative procedures. Field organizations may report field test results as a minimum increment based on the field sensor specifications. Per 62-160.340(3)(c), F.A.C., analytical results below the laboratory Reporting Limit shall be indicated as the reporting limit value followed by the data qualifier code "U" (see Table 1, rule 62-160.700, F.A.C.).