## Slide 1: ER and AR Transactivation Assay
Tests for the ability of the activated receptor to turn on gene transcription
Consists of mammalian (HEK293T) cells stably transfected with a reporter gene containing a genetically engineered DNA binding domain and loaded with a fluorescent substrate (Invitrogen, 2010)
When an estrogenic/androgenic substance, such as a water sample containing these compounds, is introduced to the transfected mammalian cells, the compound binds to the receptor activating transcription of the reporter gene (Abraham, et al. 1940; Zlokarnick 2000)
The information relayed from the activated gene initiates a series of cellular processes to produce the enzyme, β-lactamase (bla), and the gene is fully expressed (Abraham, et al. 1940; Zlokarnick 2000)
Gene expression is measured using fluorescence resonance energy transfer between the two fluorophores in the presence of bla production (Zlokarnik, et al 1998)

## Slide 2
Green Light
Estrogen receptor
Estrogen
Compound
Gal4 transcription factor
Fluoroscein
molecule
Coumarin
molecule
β-lactamase
Blue Light
HEK293T
β-lactamase reporter gene
UAS
Transactivation Assay:
Tests for the ability of the activated receptor to turn on gene transcription
Notes:   Gal4/UAS system
  Is a biochemical method used to study gene expression, and it has also been adapted to study receptor chemical-binding functions in vitro in cell culture. Wikipedia
    Gal4 – transcription factor, protein found in fungus
      Transcription factor is a protein that controls the rate of transcription of genetic information from DNA to messenger RNA, by binding to a specific DNA sequence. Wikipedia
    UAS – upstream activation sequence
      It is distinct from the promoter and increases the expression of a neighboring gene. Due to its essential role in activating transcription, the upstream activating sequence is often considered to be analogous to the function of the enhancer in multicellular eukaryotes.Wikipedia

## Slide 3: ER and AR Transactivation Assay Results
Estrogen Receptor
Grab
  No estrogenic activity at the Alafia, Withlacoochee, and field blank
  Interference at Apalachicola
  Estrogenic activity found at Ochlockonee
POCIS & SPMD
  Variable results indicated interference with the assay
  POCIS extract from Apalachicola showed no estrogenic activity
Androgen Receptor
Grab
  Low level of androgenic activity found in the field blank.
    May be due to androsterone contamination in the sample (analyzed by SGS-Axys)
  Low level of androgenic activity found in the grab sample for Ochlockonee
  Interferences were found in the Apalachicola sample
POCIS & SPMD
  No activity

## Slide 4: Estrogen and Androgen Receptor (ER & AR) Binding Assay Using Fluorescence Polarization
Tests the ability of estrogen or an estrogen mimic to bind directly to the estrogen receptor, or an androgen mimic to bind directly to the androgen receptor.
The ER ligand binding assay uses a recombinant human estrogen receptor (hER) and a fluorescein-labeled estrone derivative as the ligand (EDSTAC)
The androgen receptor binding assay employs the rat androgen receptor ligand binding domain (EDSTAC)
The competitor is identified by the hormone (e.g. estradiol) or hormone mimicking compound’s (e.g. Tamoxifin) ability to disrupt the binding of the fluorescent inhibitor molecule (EDSTAC)

## Slide 5
+
+
Low binding affinity ligand
High binding affinity ligand
Recombinant hER
Human estrogen receptor (hER)
Fluorescent ligand
Ligand Binding Assay with Fluorescent Polarization

## Slide 6: ER & AR Binding Assay Results
Estrogen Receptor
Androgen Receptor

## Slide 7: Bioluminescent Yeast Estrogen & Androgen Screen
Used to detect water’s estrogenic and androgenic activity by promoting a light-generating or bioluminescent response in the yeast strain (Sanserivino et al., 2005)
The yeast cell containing a human estrogen receptor – α gene and constructed bacterial DNA fused with reporter genes required for the bioluminescent reaction. (Eldridge, et. al 2011)
Luminescence is generated when the yeast cell is dosed with a sample containing estrogen/estrogen-mimicking compounds (Eldridge, et. al 2011)
Source: http://www.490biotech.com.

## Slide 8: Schematic representation of S. cerevisiae BLYES (Eldridge, et. al 2011)
Estrogenic compounds cross the cell membrane and bind to the human estrogen receptor (hER)
This complex interacts with estrogen response elements (RE) initiating transcription of luxA and luxB.
S. cerevisiae BLYES contains the human estrogen receptor in its genome, while S. cerevisiae BLYAS has the human androgen receptor in the genome

## Slide 9: Bioluminescence YES/YAS Results
BLYES Results
Grab
  Alafia, Apalachicola, Withlacoochee Rivers had estrogenic samples
POCIS
  Alafia, Apalachicola,Ochlockonee Withlacoochee Rivers had estrogenic samples
BLYAS results
No androgenic activity

## Slide 10: SOS Chromotest
The SOS Chromotest is a colorimetric test that measures the reaction of E. coli bacteria to DNA damaging compounds in water samples (Quillardet, et. Al 1982)
Wikipedia.org

## Slide 11: SOS Chromotest Results
POCIS & SPMD
  Presence of DNA damaging compounds at the Alafia River site.