YSI Suggestions

- Make sure calibration cup, sensor guard and all sensors are clean

- Use a clean probe guard during calibration to prevent contamination

- Never re-use standards, always use fresh standards for calibration

(it was mentioned that turbidity standards could be reused but ONLY as a rinse)

- the sensor being calibrated should be completely submerged in standard

(temperature sensor should also be submerged except when performing a DO% saturation)

- pH and Conductivity can be calibrated standing upright or facing down in the ring-stand as well as on its side as long as the correct probe is submerged in the standard

- pH and Conductivity calibrations can also be completed in a beaker or other container if the temperature sensor and sensor being calibrated are completely submerged

- an air-saturated water calibration was suggested for calibrating DO %

- all sensors can be stored long-term in deionized water EXCEPT the pH probe (which can be stored in pH buffer 4)

- pH buffer 4 works also works as a long-term storing solution for all sensors

- restore to factory calibration when necessary

- recondition sensors if they have been allowed to dry out (pH and ODO)

- make sure ports are free of water

- Calibration Order: Temp, Conductivity, pH, Depth, DO (order matters)

- Do not accept a “bad” calibration, if you accept a questionable calibration, your readings will be erroneous

Conductivity Calibration

- before calibrating, start by cleaning and drying the sensor, it should read near zero at this point

- always rinse the sensor first with DI water followed by a 3x rinse with the calibration standard

- gently rotate and/or move the sensor up and down to remove any air bubbles from the conductivity cell

- allow at least one minute for temperature equilibrium

(the white line on graph should show no significant changes for at least 40 seconds)

*if the probe is not fully submerged, the conductivity will read 1/2 of the expected calibration value*

Barometer

- the Barometer is factory calibrated and should rarely need to be recalibrated

Dissolved Oxygen

- ODO calibrations requires the current “true” barometric pressure

- wait 5 to 15 minutes for chamber to fully saturate

(the white line on graph should show no significant changes for at least 40 seconds)

PH

(most susceptible to improper handling)

- always rinse the sensor first with DI water followed by a 3x rinse with the calibration standard

- allow at least one minute for temperature equilibrium

*rehydrate sensor overnight if it has been allowed to dry out*

(the white line on graph should show no significant changes for at least 40 seconds)

Cleaning Guard

- Light cleaning: clean sensor guard with a cloth soaked in a mild solution of dish soap and water

- Deep cleaning: soak guard in a mild solution of dish soap and water or white vinegar for hard growths

- use plastic scrub brush for remaining bio-fouling

Cleaning Depth Sensor

- flush depth sensor with provided syringe and clean water, do no insert objects into ports

Cleaning Conductivity Sensor

- use provided cleaning brush to sweep away dust and debris

- clean sensor in a mild solution of dish soap and water

- soak sensor in white vinegar for heavy deposits

Cleaning Dissolved Oxygen Sensor

(typical shelf life is only 18-24 months)

- Clean sensor with a lens cleaning tissue (lint free)

- DO NOT use organic solvents to clean the sensor cap

*If ODO sensor accidentally goes dry, soak in water for 24 hours*

Cleaning pH Sensor

(typical shelf life is 12-18 months)

- soak sensor in in a mild solution of dish soap and water for 10-15 minutes

- for heavy deposits, HCL and 1:1 bleach/water solution can be used (follow instructions per YSI manual)

*If pH sensor accidentally goes dry, soak in pH buffer 4 for 24 hours*

*can store long-term in the provided shipping bottle of KCL or an airtight solution of pH buffer 4*

ODO Sensor rehydration

-if ODO sensor has been left dry for longer than 8 hours, it must be rehydrated

-soak the ODO sensor in warm (room temperature water) for 24 hours

pH Sensor reconditioning

-if pH sensor has been left dry or accidently stored in distilled/deionized water for an extended period of time, it must be reconditioned

-soak the pH sensor in buffer 4 overnight

These are only suggestions provided by YSI, always check with your SOP before making any changes to the calibration procedure.