## NEW CM TRACKER FORMAT
KEY
Presently working on
Working
In Review
Approved
Deployment Date | Retrieval Date | Project | Site Name | WBID | Station | Data: Downloaded from Sonde, Trimmed, & Saved in ROC folder | Meter Passed CCV? Yes / No (if no, specify parameter) | Aquarius: Trimmed Data Load (Working) | Aquarius QC: Data Review & Adding Qualifiers (In Review) | Aquarius QC Notes | Aquarius Approval: Review Qualifiers and Approve Data (Approved) | Aquarius Approval Notes | Save in ROC folder: CM Fieldsheet, Calibration Log & Photos | Additional Notes
2017
2017-10-19 | 2017-10-26 | SMP | McCoys Creek at Leland Street | 2257 | 20030725 | 3/16/2023KV | No | 3/16/2023KV | 3/28/24 KV | *Depth J- qualified for not being calibrated or having quarterly verification. | 5/16/2024 KV | *Necessary qualifiers present. Data properly trimmed. KV | Y
2017-11-09 | 2017-11-16 | SMP | Craig Cr @ Hendricks Ave | 2297 | 20030793 | 3/16/2023KV | No | 3/16/2023KV | 3/26/2024 KV | *Temp not quarterly verified. All but depth J- qualified. | 5/15/2024 KV | *Necessary qualifiers present. Data properly trimmed. KV | Y
2017-12-15 | 2017-12-21 | SMP | McCoys Creek at Leland Street | 2257 | 20030725 | 3/16/2023KV | No | 3/16/2023KV | 4/3/2024 KV | *Depth J- qualified, not quarterly verified | 5/16/2024 KV | *Necessary qualifiers present. Data properly trimmed. KV | Y
2018
2018-01-12 | 2018-01-18 | SMP | Moncrief Cr @ Moncrief Rd | 2228B | 20030726 | 3/16/2023KV | No | 3/16/2023KV | 4/17/2024 KV | *Depth J- qualified: no quarterly verification. pH J- qualified, did not pass mV slopes. | 5/16/2024 KV | *Necessary qualifiers present. Data properly trimmed. KV | Y
2018-01-12 | 2018-01-18 | SMP | Craig Cr @ Hendricks Ave | 2297 | 20030793 | 3/16/2023KV | No | 3/16/2023KV | 3/27/24 KV | *J- qualified: no temp or depth quarterly verifications. | 5/15/2024 KV | *Necessary qualifiers present. Data properly trimmed. KV | Y
2018-02-09 | 2018-02-14 | SMP | McCoys Creek at Leland Street | 2257 | 20030725 | 3/16/2023KV | No | 3/16/2023KV | 4/3/2024 KV | *J- qualified: no temp or depth quarterly verifications. | 5/16/2024 KV | *Necessary qualifiers present. Data properly trimmed. KV | Y
2018-02-22 | 2018-02-27 | SMP | Goodbys Creek at Baymeadows Rd | 2326B | 20030593 | 3/15/2023KV | No | 3/15/2023KV | 4/17/2024 KV | * Depth J- qualified, no quarterly verification | 5/16/2024 KV | *Necessary qualifiers present. Data properly trimmed. KV | Y
2018-03-14 | 2018-03-20 | SMP | Craig Cr @ Hendricks Ave | 2297 | 20030793 | 3/16/2023KV | No | 3/16/2023KV | 3/27/24 KV | J- qualified: no temp or depth quarterly verifications. | 5/15/2024 KV | *Necessary qualifiers present. Data properly trimmed. KV | Y
2018-04-12 | 2018-04-19 | SMP | McCoys Creek at Leland Street | 2257 | 20030725 | 3/16/2023KV | No | 3/16/2023KV | 4/4/2024 KV | All parameters J-qualified: no temp or depth quarterly verifications. | 5/16/2024 KV | *Necessary qualifiers present. Data properly trimmed. KV | Y
2018-05-10 | 2018-05-16 | SMP | Craig Cr @ Hendricks Ave | 2297 | 20030793 | 3/16/2023KV | No | 3/16/2023KV | 3/28/24 KV | J- qualified: no temp or depth quarterly verifications. | 5/15/2024 KV | *Necessary qualifiers present. Data properly trimmed. KV | Y
2018-05-10 | 2018-05-16 | SMP | Goodbys Creek at Baymeadows Rd | 2326B | 20030593 | NA- sonde did not collect data | [6 empty] | Y
2018-06-14 | 2018-06-20 | SMP | McCoys Creek at Leland Street | 2257 | 20030725 | 3/16/2023KV | No | 3/16/2023KV | 4/10/24 KV | * No pH, mV, or salinity. J- qualified. No quarterly temp or depth verifications. No lot numbers/ expiration dates for spec. con. Cal or ICV. | 5/16/2024 KV | *Necessary qualifiers present. Data properly trimmed. KV | Y
2018-06-14 | 2018-06-20 | SMP | Goodbys Creek at Baymeadows Rd | 2326B | 20030593 | 3/15/2023KV | No | 3/15/2023KV | 4/17/2024 KV | *All parameters J- qualified. No quarterly verifications | 5/16/2024 KV | *Necessary qualifiers present. Data properly trimmed. KV | Y
2018-08-16 | 2018-08-23 | SMP | McCoys Creek at Leland Street | 2257 | 20030725 | 3/16/2023KV | No | 3/16/2023KV | 4/16/24 KV | *No pH. No Salinity. J- qualified all parameters, no quarterly depth/temp verifications. No CCV post- deploy. | 5/16/2024 KV | *Necessary qualifiers present. Data properly trimmed. KV | Y
2018-08-16 | 2018-08-23 | SMP | Goodbys Creek at Baymeadows Rd | 2326B | 20030593 | 3/15/2023KV | No | 3/15/2023KV | 4/17/2024 KV | *J- qualified, no CCVs, no quarterly verifications | 5/16/2024 KV | *Necessary qualifiers present. Data properly trimmed. KV | Y
2018-09-19 | 2018-09-27 | SMP | Craig Cr @ Hendricks Ave | 2297 | 20030793 | 3/16/2023KV | No | 3/16/2023KV | 3/28/24 KV | *NO SAL or PH. All parameters J- qualified for no temperaure or depth verifications. | 5/15/2024 KV | *Necessary qualifiers present. Data properly trimmed. KV | Y
2018-10-15 | 2018-10-22 | SMP | McCoys Creek at Leland Street | 2257 | 20030725 | 3/16/2023KV | No | 3/16/2023KV | 4/17/2024 KV | J- qualified: no quarterly temperature or depth verifications, fouling | 5/16/2024 KV | *Necessary qualifiers present. Data properly trimmed. KV | Y
2018-10-15 | 2018-10-22 | SMP | Goodbys Creek at Baymeadows Rd | 2326B | 20030593 | 3/15/2023KV | No | 3/15/2023KV | 4/17/2024 KV | *No Sal or pH. No quarterly temp or depth verifications | 5/16/2024 KV | *Necessary qualifiers present. Data properly trimmed. KV | Y
2018-11-29 | 2018-12-06 | SMP | Goodbys Creek at Baymeadows Rd | 2326B | 20030593 | 3/15/2023KV | No | 3/15/2023KV | 4/17/2024 KV | *J- qualified: no quarterly temp or depth verifications, no pH CCV standard lot/ expiration | 5/16/2024 KV | *Necessary qualifiers present. Data properly trimmed. KV | Y
2018-12-21 | 2018-12-27 | SMP | Craig Cr @ Hendricks Ave | 2297 | 20030793 | 3/16/2023KV | No | 3/16/2023KV | 3/28/24 KV | * J- qualified: no temp or depth quarterly verifications. | 5/15/2024 KV | *Necessary qualifiers present. Data properly trimmed. KV | Y
2019
2019-01-17 | 2019-01-24 | SMP | Big Fishweir Creek at End of Merrimac Ave | 2280B | G2NE1175 | 3/16/2023KV | No | 3/16/2023KV | 3/18/2024KV | *J- Qualifier for all because of no quarterly verifications | 5/15/2024 KV | *Necessary qualifiers present. Data properly trimmed. KV | Y, (no pictures)
2019-02-25 | 2019-03-04 | SMP | Hopkins Cr at Kings Rd | 2266 | 20030758 | 3/16/2023KV | No | 3/16/2023KV | 3/21/2024KV | *no sal, pH, pH mV. J- qualified, depth not calibrated and no quarterly verifications. Small amount of fouling | 5/15/2024 KV | *Necessary qualifiers present. Data properly trimmed. KV | Y, (no pictures)
2019-03-18 | 2019-04-03 | SMP | Sherman Creek @ conf san pablo r | 2227 | 20030802 | 3/16/2023KV | No | 3/16/2023KV | 3/25/2024 KV | *No pH or Salinity. J- qualify all parameters for equipment malfunction, no post - deployment info or quarterly verifications. | 5/15/2024 KV | *Necessary qualifiers present. Data properly trimmed. KV | Y (no pictures)
2019-04-03 | 2019-04-09 | SMP | Hopkins Cr at Kings Rd | 2266 | 20030758 | 3/16/2023KV | No | 3/16/2023KV | 3/21/2024KV | *J- Qualifier for all because of no quarterly verifications. Some data ? Qual due to sonde out of water | 5/15/2024 KV | *Necessary qualifiers present. Data properly trimmed. KV | Y
2019-05-09 | 2019-05-16 | SMP | Big Fishweir Creek at End of Merrimac Ave | 2280B | G2NE1175 | 3/16/2023KV | No | 3/16/2023KV | 3/19/2024KV | *J- Qualifier for all because of no quarterly verifications | 5/15/2024 KV | *Necessary qualifiers present. Data properly trimmed. KV | Y
2019-05-22 | 2019-06-03 | SMP | Sherman Creek @ conf san pablo r | 2227 | 20030802 | 3/16/2023KV | No | 3/16/2023KV | 3/25/2024 KV | *No pH. J- qualified: no quarterly verifications. Smothering. | 5/15/2024 KV | *Necessary qualifiers present. Data properly trimmed. KV | Y
2019-06-06 | 2019-06-27 | SMP | Hopkins Cr at Kings Rd | 2266 | 20030758 | 3/16/2023KV | No | 3/16/2023KV | 3/21/2024KV | *no pH or mV, J- qual: no quarterly verifications depth and temp. No lot # or expir. for conductivity standards | 5/15/2024 KV | *Necessary qualifiers present. Data properly trimmed. KV | Y
2019-08-08 | 2019-08-19 | SMP | Sherman Creek @ conf san pablo r | 2227 | 20030802 | 3/16/2023KV | No | 3/16/2023KV | 3/25/2024 KV | J- qualified: depth not calibrated, no quarterly verifications, fouling during deployment. KV | 5/15/2024 KV | *Necessary qualifiers present. Data properly trimmed. KV | Y
2019-09-10 | 2019-09-16 | SMP | Big Fishweir Creek at End of Merrimac Ave | 2280B | G2NE1175 | 3/16/2023KV | No | 3/16/2023KV | 3/19/2024KV | *J- Qualifier for all because of no quarterly verifications. pH also failed CCV | 5/15/2024 KV | *Necessary qualifiers present. Data properly trimmed. KV | Y
2019-09-16 | 2019-09-25 | SMP | Hopkins Cr at Kings Rd | 2266 | 20030758 | 3/16/2023KV | No | 3/16/2023KV | 3/21/2024KV | *no pH or mV. J- qualified: no quarterly verifications | 5/15/2024 KV | *Necessary qualifiers present. Data properly trimmed. KV | Y
2019-09-30 | 2019-10-14 | SMP | Big Fishweir Creek at End of Merrimac Ave | 2280B | G2NE1175 | 3/16/2023KV | No | 3/16/2023KV | 3/19/2024KV | *J- Qualifier for all because of no quarterly verifications | 5/15/2024 KV | *Necessary qualifiers present. Data properly trimmed. KV | Y
2019-10-03 | 2019-10-10 | SMP | Sherman Creek @ conf san pablo r | 2227 | 20030802 | 3/16/2023KV | No | 3/16/2023KV | 3/25/2024 KV | J- qualified: pH failed CCV, no quarterly verifications for depth and temperature | 5/15/2024 KV | *Necessary qualifiers present. Data properly trimmed. KV | Y
2019-10-03 | 2019-10-10 | SMP | Hopkins Cr at Kings Rd | 2266 | 20030758 | 3/16/2023KV | No | 3/16/2023KV | 3/21/2024KV | * J- qualified: no quarterly verifications for depth or temperature, no CCV for DO. Some data ? Qualified: sonde out of water | 5/15/2024 KV | *Necessary qualifiers present. Data properly trimmed. KV | Y
2019-11-04 | 2019-11-12 | SMP | Sherman Creek @ conf san pablo r | 2227 | 20030802 | 3/16/2023KV | No | 3/16/2023KV | 3/26/2024 KV | *depth j- qualified. 2 points in DO ? Qualified for outliers. No pH | 5/15/2024 KV | *Necessary qualifiers present. Data properly trimmed. KV | Y
2019-11-18 | 2019-11-25 | SMP | Big Fishweir Creek at End of Merrimac Ave | 2280B | G2NE1175 | 3/16/2023KV | No | 3/16/2023KV | 3/19/2024KV | *DO ? Qualified due to erratic data. DRIFT spcon/sal?? | 5/15/2024 KV | *Necessary qualifiers present. Data properly trimmed. KV | Y
2020
2020-02-10 | 2020-02-17 | SMP | Plummer Creek at Second Island | 2130 | G4NE0301 | 3/15/2023KV | No | 3/15/2023KV | 3/12/2024KV | *J- qualified, guard not properly installed on sonde. NO pH | 5/15/2024 KV | *Necessary qualifiers present. Data properly trimmed. KV | Y, (no pictures)
2020-03-02 | 2020-03-10 | SMP | Unnamed Drain South Entrance | 2573 | G5NE0591 | 3/15/2023KV | No | 3/15/2023KV | 3/4/2024KV | *? Qualified- failed CCV DO and battery error | 5/15/2024 KV | *Necessary qualifiers present. Data properly trimmed. KV | Y, (no pictures)
2020-07-07 | 2020-07-14 | SMP | Plummer Creek at Second Island | 2130 | G4NE0301 | 3/15/2023KV | No | 3/15/2023KV | 3/12/2024KV | *J- qualified, fouling on probes and not quarterly verified. NO pH | 5/15/2024 KV | *Necessary qualifiers present. Data properly trimmed. KV | Y
2020-07-21 | 2020-07-29 | SMP | Unnamed Drain South Entrance | 2573 | G5NE0591 | 3/15/2023KV | Yes | 3/15/2023KV | 3/4/2024KV | *No pH readings during deployment. No qualifier needed KV | 5/15/2024 KV | No qualifiers needed, properly trimmed, first and last data points are within approval. KV | Y
2020-08-18 | 2020-08-25 | SMP | Unnamed Drain South Entrance | 2573 | G5NE0591 | 3/15/2023KV | Yes | 3/15/2023KV | 3/4/2024KV | *No pH readings during deployment. No qualifier needed KV | 5/15/2024 KV | No qualifiers needed, properly trimmed, first and last data points are within approval. KV | Y
2020-09-09 | 2020-09-17 | SMP | Plummer Creek at Second Island | 2130 | G4NE0301 | 3/15/2023KV | No | 3/15/2023KV | 3/12/2024KV | Depth J- qualified, not quarterly verified | 5/15/2024 KV | *Necessary qualifiers present. Data properly trimmed. KV | Y
2020-10-26 | 2020-11-03 | SMP | Unnamed Drain South Entrance | 2573 | G5NE0591 | 3/15/2023KV | Yes | 3/15/2023KV | 3/12/2024KV | *No qualifier needed | 5/15/2024 KV | No qualifiers needed, properly trimmed, first and last data points are within approval. KV | Y
2020-11-17 | 2020-11-24 | SMP | Plummer Creek at Second Island | 2130 | G4NE0301 | 3/15/2023KV | No | 3/15/2023KV | 3/12/2024KV | *All parameters J- qualified. No quarterly verifications. NO pH | 5/15/2024 KV | *Necessary qualifiers present. Data properly trimmed. KV | Y
2020-12-14 | 2020-12-22 | SMP | Plummer Creek at Second Island | 2130 | G4NE0301 | 3/15/2023KV | No | 3/15/2023KV | 3/12/2024KV | Spcon and sal J- qualified for no standard info on cal sheet, depth j- qual for not having quarterly verification. NO pH | 5/15/2024 KV | *Necessary qualifiers present. Data properly trimmed. KV | Y
2021
2021-01-05 | 2021-01-11 | SMP | Unnamed Drain South Entrance | 2573 | G5NE0591 | 3/15/2023KV | Yes | 3/15/2023KV | 3/4/2024KV | *no qualifiers needed | 5/15/2024 KV | No qualifiers needed, properly trimmed, first and last data points are within approval. KV | Y
2021-02-16 | 2021-02-23 | SMP | Pellicer Cr at US 1 | 2580B | 27010016 | 3/15/2023KV | No | 3/15/2023KV | 3/5/2024KV | *J qualified: Sonde housing sat above water level, not at proper deployment depth upon retrieval. Also quarterly temp and depth verifications cannot be verified | 5/15/2024 KV | *Necessary qualifiers present. Data properly trimmed. KV | Y, (no pictures)
2021-03-17 | 2021-03-24 | SMP | Tolomato River 800 m N of Bridge | 2363I1 | G5NE0601 | 3/15/2023KV | No | 3/15/2023KV | 3/11/2024KV | ? qualifier added to data where sonde was stranded due to tides. J- qualifier added to pH for not passing CCV. | 5/15/2024 KV | *Necessary qualifiers present. Data properly trimmed. KV | Y
2021-04-15 | 2021-04-22 | SMP | Pellicer Cr at US 1 | 2580B | 27010016 | 3/15/2023KV | No | 3/15/2023KV | 3/7/2024KV | *J Qualified- Could not verify if sonde met quarterly verification requirements | 5/15/2024 KV | *Necessary qualifiers present. Data properly trimmed. KV | Y
2021-05-17 | 2021-05-24 | SMP | Tolomato River 800 m N of Bridge | 2363I1 | G5NE0601 | 3/15/2023KV | No | 3/15/2023KV | 3/11/2024KV | *J Qualified- Could not verify if sonde met quarterly verification requirements | 5/15/2024 KV | *Necessary qualifiers present. Data properly trimmed. KV | Y
2021-06-15 | 2021-06-22 | SMP | Pellicer Cr at US 1 | 2580B | 27010016 | 3/15/2023KV | Yes | 3/15/2023KV | 3/11/2024KV | *Depth only J- qualified | 5/15/2024 KV | *Necessary qualifiers present. Data properly trimmed. KV | Y
2021-07-13 | 2021-07-20 | SMP | Tolomato River 800 m N of Bridge | 2363I1 | G5NE0601 | 3/15/2023KV | Yes | 3/15/2023KV | 3/12/2024KV | *J- Qualified due to fouling of sonde | 5/15/2024 KV | *Necessary qualifiers present. Data properly trimmed. KV | Y
2021-08-25 | 2021-09-02 | SMP | Pellicer Cr at US 1 | 2580B | 27010016 | 3/15/2023KV | Yes | 3/15/2023KV | 3/11/2024KV | *no qualifiers needed. No pH | 5/15/2024 KV | No qualifiers needed, properly trimmed, first and last data points are within approval. KV | Y
2021-10-13 | 2021-10-20 | SMP | Tolomato River 800 m N of Bridge | 2363I1 | G5NE0601 | 3/15/2023KV | Yes | 3/15/2023KV | 3/12/2024KV | *No qualifier needed. DRIFT??? (Sal/spcon) | 5/15/2024 KV | No qualifiers needed, properly trimmed, first and last data points are within approval. KV | Y
2021-11-22 | 2021-12-01 | SMP | Pellicer Cr at US 1 | 2580B | 27010016 | 3/15/2023KV | Yes | 3/15/2023KV | 3/11/2024KV | *no qualifiers needed. No pH | 5/15/2024 KV | No qualifiers needed, properly trimmed, first and last data points are within approval. KV | Y
2021-12-08 | 2021-12-15 | SMP | Tolomato River 800 m N of Bridge | 2363I1 | G5NE0601 | 3/15/2023KV | Yes | 3/15/2023KV | 3/12/2024KV | *J- qualified, bolt on sonde housing broke allowing it to sink deeper into water column. | 5/15/2024 KV | *Necessary qualifiers present. Data properly trimmed. KV | Y
2022
2022-02-16 | 2022-02-22 | SMP | Little Trout River at Old Kings Rd. | 2206 | 20030739 | 3/15/2023KV | Yes | 3/15/2023KV | 2/28/2024KV | *added note and qualifier for each parameter when sonde was out of water | 5/14/2024 KV | *Necessary qualifiers present. Data properly trimmed. KV | Y
2022-02-23 | 2022-02-28 | SMP | San Sebastian River @ SR 16 | 2491 | 27010058 | 3/15/2023KV | Yes | 3/15/2023KV | 2/19/2024KV | *J- qual all paramaters from final day of sampling from stabilization line getting stuck below surface | 5/14/2024 KV | *Necessary qualifiers present. Data properly trimmed. KV | Y
2022-03-21 | 2022-05-02 | SMP | Little Trout River at Old Kings Rd. | 2206 | 20030739 | 3/15/2023KV | No | 3/15/2023KV | 2/29/2024KV | *? Qualified, note added for each parameter (sonde battery died during deployment, no CCV) | 5/14/2024 KV | *Necessary qualifiers present. Data properly trimmed. KV | Y
2022-05-09 | 2022-05-16 | SMP | San Sebastian River @ SR 16 | 2491 | 27010058 | 3/15/2023KV | Yes | 3/15/2023KV | 2/19/2024KV | *added note and ? qualifier for each parameter when sonde was out of water. J- qualify 2nf half of deployment for equipment failure. Housing came loose from stabilizing poles | 5/14/2024 KV | *Necessary qualifiers present. Data properly trimmed. KV | Y
2022-06-06 | 2022-06-27 | SMP | Little Trout River at Old Kings Rd. | 2206 | 20030739 | 3/15/2023KV | Yes | 3/15/2023KV | 2/29/2024KV | *no qualifiers needed | 5/14/2024 KV | No qualifiers needed, properly trimmed, first and last data points are within approval. KV | Y
2022-08-11 | 2022-08-24 | SMP | San Sebastian River @ US 1 | 2491 | 27010141 | 3/15/2023KV | Yes | 3/15/2023KV | 3/19/2024KV | *no qualifiers needed | 5/14/2024 KV | No qualifiers needed, properly trimmed, first and last data points are within approval. KV | Y
2022-08-31 | 2022-09-07 | SMP | Little Trout River at Old Kings Rd. | 2206 | 20030739 | 3/15/2023KV | Yes | 3/15/2023KV | 2/29/2024KV | *no qualifiers needed | 5/14/2024 KV | No qualifiers needed, properly trimmed, first and last data points are within approval. KV | Y
2022-11-02 | 2022-11-09 | SMP | Little Trout River at Old Kings Rd. | 2206 | 20030739 | 3/15/2023KV | Yes | 3/15/2023KV | 2/29/2024KV | *no qualifiers needed | 5/14/2024 KV | No qualifiers needed, properly trimmed, first and last data points are within approval. KV | Y
2022-11-16 | 2022-11-29 | SMP | San Sebastian River @ US 1 | 2491 | 27010141 | 2/29/2024KV | Yes | 2/29/2024KV | 2/29/2024KV | *no qualifiers needed | 5/14/2024 KV | No qualifiers needed, properly trimmed, first and last data points are within approval. KV | Y
2023
2023-02-15 | 2023-03-01 | SMP | Thomas Creek 5 Miles Down From Ethel | 2161B | G4NE0298 | 4/26/2023KV | Yes | 4/26/2023KV | 2/5/2024KV | *no qualifiers needed | 5/14/2024 KV | No qualifiers needed, properly trimmed, first and last data points are within approval. KV | Y
2023-04-12 | 2023-04-26 | SMP | Thomas Creek 5 Miles Down From Ethel | 2161B | G4NE0298 | 4/26/2023KV | Yes | 4/26/2023KV | 2/5/2024KV | *no qualifiers needed | 5/14/2024 KV | No qualifiers needed, properly trimmed, first and last data points are within approval. KV | Y
2023-06-14 | 2023-06-28 | SMP | Thomas Creek 5 Miles Down From Ethel | 2161B | G4NE0298 | 6/28/2023KV | No- pH | 6/28/2023KV | 2/5/2024KV | *pH J- qualified. mV did not show on handheld for ICV or CCV. KV | 5/14/2024 KV | *Necessary qualifiers present. Data properly trimmed. KV | Y | *pH noticed during approval. Changed then. KV
2023-08-09 | 2023-08-23 | SMP | Thomas Creek 5 Miles Down From Ethel | 2161B | G4NE0298 | 9/5/2023KV | No- pH | 9/5/2023KV | 2/5/2024KV | *pH J- qualified. mV did not show on handheld for ICV or CCV. KV | 5/14/2024 KV | *Necessary qualifiers present. Data properly trimmed. KV | Y | *pH noticed during approval. Changed then. KV
2023-10-11 | 2023-10-25 | SMP | Thomas Creek 5 Miles Down From Ethel | 2161B | G4NE0298 | 10/25/2023KV | No- pH | 10/25/2023KV | 2/5/2024KV | *J- qualified, equiptment housing failure. pH J- qualified. mV did not show on handheld for ICV or CCV. KV | 5/14/2024 KV | *Necessary qualifiers present. Data properly trimmed. KV | Y | *pH noticed during approval. Changed then. KV
2023-12-06 | 2023-12-20 | SMP | Thomas Creek 5 Miles Down From Ethel | 2161B | G4NE0298 | 12/6/2023KV | Yes | 12/6/2023KV | 2/29/2024KV | *Added rain qualifier 17th-19th | 5/14/2024 KV | *Necessary qualifiers present. Data properly trimmed. KV | Y
2024
2024-01-16 | 2024-01-31 | SMP | Sisters Creek NW of Shell Island | 2205A | G2NE1232 | 2/29/2024KV | Yes | 2/29/2024KV | 2/29/2024KV | *no qualifiers needed | 4/18/2024 KV | No qualifiers needed, properly trimmed, first and last data points are within approval. KV | Y
2024-02-14 | 2024-02-28 | SMP | Plummer Creek at Second Island | 2130 | G4NE0301 | 2/29/2024KV | Yes | 2/29/2024KV | 2/29/2024KV | *no qualifiers needed | 4/18/2024 KV | No qualifiers needed, properly trimmed, first and last data points are within approval. KV | Y
2024-03-06 | 2024-03-20 | SMP | Sisters Creek NW of Shell Island | 2205A | G2NE1232 | 3/20/2024KV | Yes | 3/20/2024KV | 4/16/2024 KV | *J- qualified for fouling | 4/18/2024 KV | All qualifiers needed are present. properly trimmed, first and last data points are within approval range. KV | Y
2024-04-09 | 2024-04-25 | SMP | Plummer Creek at Second Island | 2130 | G4NE0301 | 4/25/2024 KV | Yes | 4/25/2024 KV | 4/25/2024 KV | *no qualifiers needed | 5/14/2024 KV | No qualifiers needed, properly trimmed, first and last data points are within approval. KV | Y
2024-04-30 | 2024-05-23 | SMP | Sisters Creek NW of Shell Island | 2205A | G2NE1232 | 4/23/2024 KV | Yes | 4/23/2024 KV | 4/28/2024 KV | *J- qualified first half of deployment for SMP requirements not being completed, entire deployment has J- qualifier for potential tampering. ?- qualified 2nd half for intermittent desiccation. | 6/19/2024 KV | Necessary qualifiers present. KV | Y | DEPLOYMENT DOES NOT COUNT TOWARDS SMP REQUIREMENT. Sonde ended up very close to shore half way through deployment. Went out to switch batteries during lab closure to put off taking samples upon retrieval and found float line missing and anchor with sonde housing close to shore. Emergency retrieved; does not meet SMP requirements. Data still uploaded and qualified. | Additional deployment attempted 6/5/24-6/24/24 however sonde did not deploy properly and data was not collected.
2024-07-10 | 2024-07-18 | SMP | Plummer Creek at Second Island | 2130 | G4NE0301 | 7/11/2024 KV | Yes | 8/27/2024 KV

## Sheet1
Planned Deployments | [5 empty] | *Minimum 5 deployments per site, min 1 deployment per quarter per site
Quarter | Date | Sisters Creek | Plummers Creek | Completed (Y/N) | Quarterly requirement met (Y/N)
1 | 2024-01-16 | 1 | _ | Y | Y
1 | 2024-02-14 | _ | 1 | Y | Y
1 | 2024-03-06 | 2 | _ | Y | _
2 | 2024-04-09 | _ | 2 | Y | Y
2 | 2024-04-30 | 3 | _ | N | _
2 | 2024-06-05 | 3 | _
3 | 2024-07-03 | _ | 3
3 | 2024-08-07 | 4 | _
3 | 2024-09-04 | _ | 4
4 | 2024-10-02 | 5 | _
4 | 2024-11-06 | _ | 5

## OLD CM AQUARIUS TRACKER
KEY
Presently working on
Working
In Review
Approved
Deployment Date | Retrieval Date | Project | Site Name | WBID | Station | Save: Data file/ FS/ Cal Log/ Pics | Aquarius: Trimming & Data Load (within 2 weeks of retrieval) | QC1: Data Review & Adding Qualifiers (In Review) | QC1 Notes | Photos (Taken during first site visit of year) Y/N | All Docs (see notes on sheet 2) Y/N | QC2: Data Approval (Approved) | Initials | Approval Notes
2017
2017-10-19 | 2017-10-26 | SMP | McCoys Creek at Leland Street | 2257 | 20030725 |  | 3/16/2023KV | 6/27/23 NB |  | Y
2017-11-09 | 2017-11-16 | SMP | Craig Cr @ Hendricks Ave | 2297 | 20030793 |  | 3/16/2023KV |  |  | Y
2017-12-15 | 2017-12-21 | SMP | McCoys Creek at Leland Street | 2257 | 20030725 |  | 3/16/2023KV | 6/28/23 NB | trimmed
2018
2018-01-12 | 2018-01-18 | SMP | Moncrief Cr @ Moncrief Rd | 2228B | 20030726 |  | 3/16/2023KV |  |  | Y
2018-01-12 | 2018-01-18 | SMP | Craig Cr @ Hendricks Ave | 2297 | 20030793 |  | 3/16/2023KV |  |  | Y
2018-02-09 | 2018-02-14 | SMP | McCoys Creek at Leland Street | 2257 | 20030725 |  | 3/16/2023KV | 6/27/23 NB |  | Y
2018-02-22 | 2018-02-27 | SMP | Goodbys Creek at Baymeadows Rd | 2326B | 20030593 |  | 3/15/2023KV |  | conflicting deployment / retrieval dates on cm sheet | Y
2008-03-14 | 2018-03-20 | SMP | Craig Cr @ Hendricks Ave | 2297 | 20030793 |  | 3/16/2023KV
2018-04-12 | 2018-04-19 | SMP | McCoys Creek at Leland Street | 2257 | 20030725 |  | 3/16/2023KV | 6/27/23 NB | trimmed, R qualifier
2018-05-10 | 2018-05-16 | SMP | Craig Cr @ Hendricks Ave | 2297 | 20030793 |  | 3/16/2023KV
2018-06-14 | 2018-06-20 | SMP | McCoys Creek at Leland Street | 2257 | 20030725 |  | 3/16/2023KV | 6/28/23 NB | no pH probe & R qualifier
2018-06-14 | 2018-06-20 | SMP | Goodbys Creek at Baymeadows Rd | 2326B | 20030593 |  | 3/15/2023KV
2018-08-16 | 2018-08-23 | SMP | McCoys Creek at Leland Street | 2257 | 20030725 |  | 3/16/2023KV | 6/28/23 NB | no pH probe & R qualifier
2018-08-16 | 2018-08-23 | SMP | Goodbys Creek at Baymeadows Rd | 2326B | 20030593 |  | 3/15/2023KV
2018-09-19 | 2018-09-27 | SMP | Craig Cr @ Hendricks Ave | 2297 | 20030793 |  | 3/16/2023KV |  | retrieval data not on cm sheet
2018-10-15 | 2018-10-22 | SMP | McCoys Creek at Leland Street | 2257 | 20030725 |  | 3/16/2023KV | 6/28/23 NB | trimmed
2018-10-15 | 2018-10-22 | SMP | Goodbys Creek at Baymeadows Rd | 2326B | 20030593 |  | 3/15/2023KV
2018-11-29 | 2018-12-06 | SMP | Goodbys Creek at Baymeadows Rd | 2326B | 20030593 |  | 3/15/2023KV
2018-12-21 | 2018-12-27 | SMP | Craig Cr @ Hendricks Ave | 2297 | 20030793 |  | 3/16/2023KV
2019
2019-01-17 | 2019-01-24 | SMP | Big Fishweir Creek at End of Merrimac Ave | 2280B | G2NE1175 |  | 3/16/2023KV | 3/18/2024KV | *R qualifier; J- Qualifier for all because of no quarterly verifications | N | Y
2019-02-25 | 2019-03-04 | SMP | Hopkins Cr at Kings Rd | 2266 | 20030758 |  | 3/16/2023KV | 3/21/2024KV | *no sal, pH, pH mV. J- qualified, depth not calibrated and no quarterly verifications. Small amount of fouling | N | Y
2019-03-18 | 2019-04-03 | SMP | Sherman Creek @ conf san pablo r | 2227 | 20030802 |  | 3/16/2023KV |  |  | N
2019-04-03 | 2019-04-09 | SMP | Hopkins Cr at Kings Rd | 2266 | 20030758 |  | 3/16/2023KV | 3/21/2024KV | *J- Qualifier for all because of no quarterly verifications. Some data ? Qual due to sonde out of water |  | Y
2019-05-09 | 2019-05-16 | SMP | Big Fishweir Creek at End of Merrimac Ave | 2280B | G2NE1175 |  | 3/16/2023KV | 3/19/2014KV | *J- Qualifier for all because of no quarterly verifications |  | Y
2019-05-22 | 2019-06-03 | SMP | Sherman Creek @ conf san pablo r | 2227 | 20030802 |  | 3/16/2023KV
2019-06-06 | 2019-06-27 | SMP | Hopkins Cr at Kings Rd | 2266 | 20030758 |  | 3/16/2023KV | 3/21/2024KV | *no pH or mV, J- qual: no quarterly verifications depth and temp. No lot # or expir. for conductivity standards |  | Y
2019-08-08 | 2019-08-19 | SMP | Sherman Creek @ conf san pablo r | 2227 | 20030802 |  | 3/16/2023KV
2019-09-10 | 2019-09-16 | SMP | Big Fishweir Creek at End of Merrimac Ave | 2280B | G2NE1175 |  | 3/16/2023KV | 3/19/2024KV | *J- Qualifier for all because of no quarterly verifications. pH also failed CCV |  | Y
2019-09-16 | 2019-09-25 | SMP | Hopkins Cr at Kings Rd | 2266 | 20030758 |  | 3/16/2023KV | 3/21/2024KV | *no pH or mV. J- qualified: no quarterly verifications |  | Y
2019-09-30 | 2019-10-14 | SMP | Big Fishweir Creek at End of Merrimac Ave | 2280B | G2NE1175 |  | 3/16/2023KV | 3/19/2024KV | *J- Qualifier for all because of no quarterly verifications |  | Y
2019-10-03 | 2019-10-10 | SMP | Sherman Creek @ conf san pablo r | 2227 | 20030802 |  | 3/16/2023KV
2019-10-03 | 2019-10-10 | SMP | Hopkins Cr at Kings Rd | 2266 | 20030758 |  | 3/16/2023KV | 3/21/2024KV | * J- qualified: no quarterly verifications for depth or temperature, no CCV for DO. Some data ? Qualified: sonde out of water |  | Y
2019-11-04 | 2019-11-12 | SMP | Sherman Creek @ conf san pablo r | 2227 | 20030802 |  | 3/16/2023KV
2019-11-18 | 2019-11-25 | SMP | Big Fishweir Creek at End of Merrimac Ave | 2280B | G2NE1175 |  | 3/16/2023KV | 3/19/2024KV | *DO ? Qualified due to erratic data. DRIFT spcon/sal?? |  | Y
2020
2020-02-10 | 2020-02-17 | SMP | Plummer Creek at Second Island | 2130 | G4NE0301 |  | 3/15/2023KV | 3/12/2024KV | *J- qualified, guard not properly installed on sonde. NO pH | N | Y
2020-03-02 | 2020-03-10 | SMP | Unnamed Drain South Entrance | 2573 | G5NE0591 |  | 3/15/2023KV | 3/4/2024KV | *? Qualified- failed CCV DO and battery error | N | Y
2020-07-07 | 2020-07-14 | SMP | Plummer Creek at Second Island | 2130 | G4NE0301 |  | 3/15/2023KV | 3/12/2024KV | *J- qualified, fouling on probes and not quarterly verified. NO pH |  | Y
2020-07-21 | 2020-07-29 | SMP | Unnamed Drain South Entrance | 2573 | G5NE0591 |  | 3/15/2023KV | 3/4/2024KV | *No pH readings during deployment. No qualifier needed KV |  | Y
2020-08-18 | 2020-08-25 | SMP | Unnamed Drain South Entrance | 2573 | G5NE0591 |  | 3/15/2023KV | 3/4/2024KV | *No pH readings during deployment. No qualifier needed KV |  | Y
2020-09-09 | 2020-09-17 | SMP | Plummer Creek at Second Island | 2130 | G4NE0301 |  | 3/15/2023KV | 3/12/2024KV | Depth J- qualified, not quarterly verified |  | Y
2020-10-26 | 2020-11-03 | SMP | Unnamed Drain South Entrance | 2573 | G5NE0591 |  | 3/15/2023KV | 3/12/2024KV | *No qualifier needed |  | Y
2020-11-17 | 2020-11-24 | SMP | Plummer Creek at Second Island | 2130 | G4NE0301 |  | 3/15/2023KV | 3/12/2024KV | *All parameters J- qualified. No quarterly verifications. NO pH |  | Y
2020-12-14 | 2020-12-22 | SMP | Plummer Creek at Second Island | 2130 | G4NE0301 |  | 3/15/2023KV | 3/12/2024KV | Spcon and sal J- qualified for no standard info on cal sheet, depth j- qual for not having quarterly verification. NO pH |  | Y
2021
2021-01-05 | 2021-01-11 | SMP | Unnamed Drain South Entrance | 2573 | G5NE0591 |  | 3/15/2023KV | 3/4/2024KV | *no qualifiers needed | Y | Y
2021-02-16 | 2021-02-23 | SMP | Pellicer Cr at US 1 | 2580B | 27010016 |  | 3/15/2023KV | 3/5/2024KV | *J qualified: Sonde housing sat above water level, not at proper deployment depth upon retrieval. Also quarterly temp and depth verifications cannot be verified | N | Y
2021-03-17 | 2021-03-24 | SMP | Tolomato River 800 m N of Bridge | 2363I1 | G5NE0601 |  | 3/15/2023KV | 3/11/2024KV | ? qualifier added to data where sonde was stranded due to tides. J- qualifier added to pH for not passing CCV. | Y | Y
2021-04-15 | 2021-04-22 | SMP | Pellicer Cr at US 1 | 2580B | 27010016 |  | 3/15/2023KV | 3/7/2024KV | *J Qualified- Could not verify if sonde met quarterly verification requirements |  | Y
2021-05-17 | 2021-05-24 | SMP | Tolomato River 800 m N of Bridge | 2363I1 | G5NE0601 |  | 3/15/2023KV | 3/11/2024KV | *J Qualified- Could not verify if sonde met quarterly verification requirements |  | Y
2021-06-15 | 2021-06-22 | SMP | Pellicer Cr at US 1 | 2580B | 27010016 |  | 3/15/2023KV | 3/11/2024KV | *Depth only J- qualified |  | Y
2021-07-13 | 2021-07-20 | SMP | Tolomato River 800 m N of Bridge | 2363I1 | G5NE0601 |  | 3/15/2023KV | 3/12/2024KV | *J- Qualified due to fouling of sonde |  | Y
2021-08-25 | 2021-09-02 | SMP | Pellicer Cr at US 1 | 2580B | 27010016 |  | 3/15/2023KV | 3/11/2024KV | *no qualifiers needed. No pH |  | Y
2021-10-13 | 2021-10-20 | SMP | Tolomato River 800 m N of Bridge | 2363I1 | G5NE0601 |  | 3/15/2023KV | 3/12/2024KV | *No qualifier needed. DRIFT??? (Sal/spcon) |  | Y
2021-11-22 | 2021-12-01 | SMP | Pellicer Cr at US 1 | 2580B | 27010016 |  | 3/15/2023KV | 3/11/2024KV | *no qualifiers needed. No pH |  | Y
2021-12-08 | 2021-12-15 | SMP | Tolomato River 800 m N of Bridge | 2363I1 | G5NE0601 |  | 3/15/2023KV | 3/12/2024KV | *J- qualified, bolt on sonde housing broke allowing it to sink deeper into water column. |  | Y
2022
2022-02-16 | 2022-02-22 | SMP | Little Trout River at Old Kings Rd. | 2206 | 20030739 |  | 3/15/2023KV | 2/28/2024KV | *added note and qualifier for each parameter when sonde was out of water | Y | Y
2022-02-23 | 2022-02-28 | SMP | San Sebastian River @ SR 16 | 2491 | 27010058 |  | 3/15/2023KV | 2/19/2024KV | *added note on depth from final day of sampling from stabilization line getting stuck below surface | Y | Y
2022-03-21 | 2022-05-02 | SMP | Little Trout River at Old Kings Rd. | 2206 | 20030739 |  | 3/15/2023KV | 2/29/2024KV | *? Qualified, note added for each parameter (sonde battery died during deployment, no CCV) |  | Y
2022-05-09 | 2022-05-16 | SMP | San Sebastian River @ SR 16 | 2491 | 27010058 |  | 3/15/2023KV | 2/19/2024KV | *added note and qualifier for each parameter when sonde was out of water |  | Y
2022-06-06 | 2022-06-27 | SMP | Little Trout River at Old Kings Rd. | 2206 | 20030739 |  | 3/15/2023KV | 2/29/2024KV | *no qualifiers needed |  | Y
2022-08-11 | 2022-08-24 | SMP | San Sebastian River @ US 1 | 2491 | 27010141 |  | 3/15/2023KV | 3/19/2024KV | *no qualifiers needed | Y | Y
2022-08-31 | 2022-09-07 | SMP | Little Trout River at Old Kings Rd. | 2206 | 20030739 |  | 3/15/2023KV | 2/29/2024KV | *no qualifiers needed |  | Y
2022-11-02 | 2022-11-09 | SMP | Little Trout River at Old Kings Rd. | 2206 | 20030739 |  | 3/15/2023KV | 2/29/2024KV | *no qualifiers needed |  | Y
2022-11-16 | 2022-11-29 | SMP | San Sebastian River @ US 1 | 2491 | 27010141 |  | 2/29/2024KV | 2/29/2024KV | *no qualifiers needed |  | Y
2023
2023-02-15 | 2023-03-01 | SMP | Thomas Creek 5 Miles Down From Ethel | 2161B | G4NE0298 | 4/12/23 KV | 4/26/2023KV | 2/5/2024KV | *no qualifiers needed | Y | Y
2023-04-12 | 2023-04-26 | SMP | Thomas Creek 5 Miles Down From Ethel | 2161B | G4NE0298 | 4/26/23 KV | 4/26/2023KV | 2/5/2024KV | *no qualifiers needed |  | Y
2023-06-14 | 2023-06-28 | SMP | Thomas Creek 5 Miles Down From Ethel | 2161B | G4NE0298 |  | 6/28/2023KV | 2/5/2024KV | *no qualifiers needed |  | Y
2023-08-09 | 2023-08-23 | SMP | Thomas Creek 5 Miles Down From Ethel | 2161B | G4NE0298 |  | 9/5/2023KV | 2/5/2024KV | *no qualifiers needed |  | Y
2023-10-11 | 2023-10-25 | SMP | Thomas Creek 5 Miles Down From Ethel | 2161B | G4NE0298 |  | 10/25/2023KV | 2/5/2024KV | *J- qualified, equiptment housing failure |  | Y
2023-12-06 | 2023-12-20 | SMP | Thomas Creek 5 Miles Down From Ethel | 2161B | G4NE0298 |  | 12/6/2023KV | 2/29/2024KV | *Added rain qualifier 17th-19th |  | Y
2024
2024-01-16 | 2024-01-31 | SMP | Sisters Creek NW of Shell Island | 2205A | G2NE1232 |  | 2/29/2024KV | 2/29/2024KV | *no qualifiers needed | Y | Y
2024-02-14 | 2024-02-28 | SMP | Plummer Creek at Second Island | 2130 | G4NE0301 |  | 2/29/2024KV | 2/29/2024KV | *no qualifiers needed | Y | Y
2024-03-06 | 2024-03-20 | SMP | Sisters Creek NW of Shell Island | 2205A | G2NE1232 |  | 3/20/2024KV |  |  |  | Y

## TIPS
Approval notes- examples ONLY FOR APPROVER
Depth/ temp | Data trimmed, no outliers. No qualifiers.
DO | Data trimmed, DO mg/l and % graphs correspond. Data not qualified.
Spc/Sal | Data trimmed, no outliers. Spc and sal graphs correspond. Data not Qualified.
pH | Data trimmed, no outliers. pH and pH volt graphs. Data not qualified.
What to look for in documents:
CM field sheet
Cal sheet for deployed sonde
Cal sheet for ambient measurements meter day of deployment
Field sheet day of deployment, contract lab COC if applicable from day of deployment
Cal sheet for ambient measurements meter day of retrieval
Field sheet from day of retrieval
Contract lab COC from day of retrieval
Photos:
If they are not in CM photos folder, look in: WQAPmonitoring> StrategicMonitoring> [YEAR SAMPLE OCCURRED]> 21FLA> WBID#_NAME_WATERBODY> 21FLA [SiteID#]
Copy and paste photos to CM photos folder