EcoWatch

Start Logging Sonde:

-plug in sonde

-“click” Ecowatch Lite Icon

-“click” File  New Connection

-“press” Enter until # symbol appears

-type Menu, press enter

7. Sensor

Check that all necessary sensors have been selected (make sure optical DO has been selected-ODO)

6. Report

Again check that all necessary sensors have been selected (this ensures that the correct sensors will be reported when data is exported)

5. System

(1.) Check Date and Time (time- drift usually occurs, sync sonde with computer time)

1. Run

(2.) Unattended Sample

1. Check Interval

2. Check Date

4. Check Duration (should be 365)

5. Enter File Name

6. enter site name

7. Check Battery Voltage

8. Check Battery life

9.Check Free Memory

10. Check Parameters to Log

C. Start logging

Escape all the way until Menu page appears (use Esc or 0)

4. Check Status

(If sonde is in Run-Mode, then Logging Active will be displayed)

-Escape all the way out until the exit menu appears and type “Y” for yes

-Close program and disconnect YSI

Stop Logging Sonde:

-plug in sonde

-“click” Ecowatch Lite Icon

-“click” File  New Connection

-“press” Enter until # symbol appears

-type Menu, press enter

1. Run

(2.) Unattended Sample

B. Stop logging

Escape all the way until Menu page appears (use Esc or 0)

3. File

(2.) Upload

-Select correct file (ex. 1)

(1.) Proceed

- download 1.) PC600 and 2.)Comma/Delimited

I usually download files as a PC600 and a Comma/Delimited file as well as an excel or CSV file.

To get excel version: FileOpenselect correct file  FileExportselect excel or CSV

Move all files over to Continuous Monitoring Folder located Here; trim any unnecessary data, scan data for errors, review pH and Sp/Cond ranges to determine what standards to bracket with.