## Instructions
Instructions
● | Do not move, edit, swap, or shift any columns, rows or cells in the calibration log.
● | In the Drop-downs tab, edit the white columns to match your station, probe, and sonde information. Gray columns should not be edited, only selected in the log drop downs.
● | Within the log, only in white boxes labeled notes/other, you may type in extra items you wish the calibration log to prompt (see example).
● | If a calibration or ICV fails, place a note with the failed value in the pre-deployment comments section. Passing recalibrated value should go in the actual log.
● | General maintenance section is for any maintenance (in-field maintenance, probe repair date/ reason, etc.).
● | Parameters should be calibrated in the order of the calibration log from top to bottom; this matters especially for temperature and SpCond.
● | Your NIST device should be at least as accurate as your sensor. The accuracy of the NIST device must be included in your temperature accuracy determination, meaning that the temperature probe specification should be added to the accuracy of your thermometer.
● | Submit calibration log in .xls format named by AP station code and year, for example, BBSCH_2021_CalLog. There should be one calibration log excel file for each station and year.
● | Tabs at the bottom should be named by the station code and deployment date, for example, BBSCH_010222. There should be approximately 12 tabs in each yearly calibration log file.
Other information
● | Optional parameters will not be uploaded to the AP Data Portal at this time but will remain in the dataset for possible future upload.
● | Check marks represent DEP SOP requirement for ICV (initial calibration verification) and CCV (continuing calibration verification) steps.
● | The DL next to a parameter's units indicates that the reading comes from the datalogger's probe.

## FT1500
Solubility of Oxygen in Water at Atmospheric Pressure
Temperature | Oxygen Solubility | Use this table to DEP verify that DO mg/L are reading within range based on the temperature reading.
°C | mg/L | DEP SOP FT 1500
0 | 14.621
1 | 14.216
2 | 13.829
3 | 13.46
4 | 13.107
5 | 12.77
6 | 12.447
7 | 12.139
8 | 11.843
9 | 11.559
10 | 11.288
11 | 11.027
12 | 10.777
13 | 10.537
14 | 10.306
15 | 10.084
16 | 9.87
17 | 9.665
18 | 9.467
19 | 9.276
20 | 9.092
21 | 8.915
22 | 8.743
23 | 8.578
24 | 8.418
25 | 8.263
26 | 8.113
27 | 7.968
28 | 7.827
29 | 7.691
30 | 7.559
31 | 7.43
32 | 7.305
33 | 7.183
34 | 7.065
35 | 9.65
36 | 6.837
37 | 6.727
38 | 6.62
39 | 6.515
40 | 6.412
41 | 6.312
42 | 6.213
43 | 6.116
44 | 6.021
45 | 5.927
46 | 8.835
47 | 5.744
48 | 5.654
49 | 5.565
50 | 5.477

## Drop-downs
AP | Station | Datasonde | Sonde No. | pH Serial No. | ODO Serial No. | Turb Serial No. | Cond Serial No. | Chl Serial No. | Wiper Serial No. | Y/N | Pass/Fail | pH Model | ODO Model | Turb Model | Cond Model | Chl Model | Wiper Model | Sonde Model | Lot No. | Technician | Cell Constant |  | Battery Voltage | DO Gain |  | Sp Cond (mS/cm) |  | Temperature |  | DO |  | Depth |  | Turbidity | pH
BBSAP | BBSCH (Chassahowitzka) | 156018 | 19E101880 | 20B000147 | 20A103173 | 20A103291 | 20A103493 | 19J104091 | 20A103351 | Y | P | 577602 | 599100-01 | 599101-01 | 599827 | 599103-01 | 599090-01 | EXO2 | 2908A10 | Trisha Green | EXO Wiped CT | 0.419-0.519 (w) | ~6V (EXO 2) | EXO Optical | 0.87-1.25 |  | ±5% |  | ±0.5(w)+0.1°C | Optical (0-200% air sat.) | ± 0.3 mg/L | EXO (10m) | ±5% | 0.00 ±0.3 | 7 ±0.2
SMMAP | BBSST (Steinhatchee) | 156019 | 19E101881 | 20B000148 | 20A103174 | 20A103293 | 20A103494 | 19J104090 | 20A103352 | N | F | 577602 | 599100-01 | 599101-01 | 599827 |  | 599090-01 | EXO3 | 2008B00 | Mary McMurray | EXO Non-Wiped CT | 5.05-5.95 (n) | ~3V (EXO 3) | 6600 | 0.7 - 1.40 | [6 empty] | EXO (10m) | ±0.05 m | 124 ±6.2 | 10 ±0.2
CHAP | EB01 (Tom Winter) | 156894 | 19K100551 | 19K101466 | 19J104036 | 19J104933 | 19J103598 |  | 19J105081 | Auto | [8 empty] | 2004922 | Aliza Karim | [11 empty] | EXO (10m) | whichever is greater |  | 4 ±0.2
MPAP | EB04 (Hendry & Mullock Creeks) | 156893 | 19K100552 | 19K101465 | 19J104035 | 19J104932 | 19J103597 |  | 19J105082 | [9 empty] | 2103C60 | Jon Brucker | [11 empty] | 6600 (10m) | ±0.06 ft (0.018 m)
NWFAP | MP1A (northern MPAP) | 1 | 21E103416 | 21G105082 | 21G100611 | 21G100846 | 21G101304 |  | 21H105437 | [9 empty] | 4102C02 | Lena Kury
CPAP | MP2B (southern MPAP) | 2 | 21E103418 | 21G105084 | 21G100610 | 21G100848 | 21G101301 |  | 21H105435 | [9 empty] | 21A20510133
TMAP | MP3C (middle MPAP) |  |  | 21G105935 |  |  | 21G101303
| CHWW1 (West Wall of CHAP) | [5 empty] | 23F106588
CPAH2
Note on AA1: Howe, Samantha: 0 to 100: ± 0.5% of reading or 0.001 mS/cm 100 to 200: ± 1% of reading Wiped Sensor: ±1% of the reading or 0.002 mS/cm (whichever is greater)

## Example
Aquatic Preserve Continuous Water Quality Monitoring Calibration Log
|  | File Name | BBSCH_110221 | [5 empty] | .csv | [5 empty] | File Name | BBSCH19E101881_110221_071500 |  |  |  |  | .bin
[7 empty] | Datasonde and Probe Identification Numbers | [12 empty] | Calibration Standards
|  | Sonde Code |  |  | Serial Number |  |  | Serial Number |  | Model Number | [5 empty] | Standard | Lot No. |  |  | Received | Opened | Exp. Date
| Datasonde | 156019 |  |  | 19E101880 |  | pH | 19K101465 |  | 577602 |  |  |  |  | SpCond (mS/cm) | 20 | 2008A95 |  |  | 2021-07-12 | 08/23/2021 | 07/2022
| Vented | Y/N |  |  | Model |  | ODO | 20A103173 |  | 599100-01 |  |  |  |  | SpCond (mS/cm) | 50 | 2105829 |  |  | 2021-07-01 | 09/16/2021 | 04/2023
| Nickname | CH |  |  | EXO2 |  | Turb | 20A103291 |  | 599101-01 |  |  |  |  | pH | 4 | 2004922 |  |  | 2021-07-12 | 08/27/2021 | 03/2022
| Preserve | BBSAP |  |  |  |  | Cond | 20A103494 |  | 599827 |  |  |  |  | pH | 7 | 2103C60 |  |  | 2021-05-03 | 08/04/2021 | 02/2023
| Station | BBSCH (Chassahowitzka) |  |  |  |  | Chl | 19J104091 |  | 599103-01 |  |  |  |  | pH | 10 | 4102C02 |  |  | 2021-04-30 | 10/06/2021 | 07/2022
[7 empty] | Wiper | 19J105081 |  | 599090-01 |  |  |  |  | Turb (FNU) | 124 | 21H21280150 |  |  | 2021-08-24 | 08/27/2021 | 08/2022
|  |  |  | Wiper Replaced | N |  |  |  | Batteries Replaced | N | [5 empty] | Technician(s) | Trisha Green
|  |  |  | Wiper Parked in Garage | Y |  |  |  | Date of Calibration | 2021-11-02 | [5 empty] | Date of Post-Calibration | 2021-11-23
| Pre-deployment Comments/ Failed-calibration values: | If a calibration QC score or ICV fails, place a note with the failed value in the pre-deployment comments section. Passing recalibrated value should go in the actual log.
| Pre-Deployment (CAL & ICV) | [12 empty] | Post-Deployment (CCV) | [5 empty] | Programming
|  | Standards |  |  | Before Cal |  |  | Calibrated |  | Specs |  | P/ F | Time |  | Post value |  | P/ F | Time |  | Interval |  | 15
| Temp | 21.75 | °C (Therm) |  | 21.693 | °C (DL) |  |  |  | ±0.5(w)+0.1°C |  | P | 7:10 |  | 20.5 | °C (Therm) | P | 12:56 |  | Duration |  | 365 days
| SpCond in air |  |  |  | 0 |  |  |  |  | ≤ 0.002 mS/cm |  | P |  |  | 20.46 | °C (DL) |  |  |  | Free memory % |  | 99.55
| SpCond | 50 | mS/cm |  | 49.767 | mS/cm |  | 50 | mS/cm | ±5% |  |  | 7:08 | ✔ | 50.02 | mS/cm | P | 13:00 |  | Memory (days) |  | 24432.81
[6 empty] | Cell Constant |  | 0.47 |  | 0.419-0.519 (w) |  | P | [7 empty] | Start time |  | 9:30
| SpCond 20 (ICV) | 20 | mS/cm |  |  | ✔ |  | 20.112 | mS/cm | ±5% |  | P | 7:14 | [6 empty] | Battery % |  | 65
| pH 7 | 7 |  |  | 7.05 |  |  | 7.02 |  | 7 ±0.2 |  |  | 7:11 | ✔ | 7.1 |  | P | 13:04 |  | Set clock |  | Y/N
| pH 10 | 10 |  |  | 10.08 |  |  | 10.05 |  | 10 ±0.2 |  |  | 7:14 | ✔ | 10.06 |  | P | 13:05 |  | End time |  | 10:00
[5 empty] | pH 7 |  |  | -30.7 | mV | 0 ±50 mV |  | P |  |  | -35.5 | mV | P |  |  |  | Other/notes
[5 empty] | pH 10 |  |  | -204.1 | mV | -180 ±50 mV |  | P |  |  | -203.9 | mV | P |  |  | Change locks |  | Y
[5 empty] | pH 4 |  |  | 148.8 | mV | +180 ±50 mV |  | P | [7 empty] | Check for DST |  | Y
[5 empty] | pH Slope |  |  | 173.4 |  | ≥ 165 mV |  | P |  |  | 168.4 |  | P
| pH 4 or 10 (ICV) | 4 |  |  |  | ✔ |  | 3.9 |  | 4 ±0.2 |  | P | 7:16
| BP at Depth Cal |  |  |  | 765.556 | mmHg | [8 empty] | 765.56 | mmHg
| Depth | 0.075534 | Offset |  | 0.062 | m |  | 0.076 | m |  |  |  | 7:24 |  | 0.072 | m | P | 13:10
| Turbidity | 0 | FNU |  | -0.1 |  |  | 0 | FNU | 0.00 ±0.3 |  |  | 7:27 | ✔ | 0.04 | FNU | P | 13:15 |  |  | General Maintenance
| Turbidity | 124 | FNU |  | 123.24 | FNU |  | 124 | FNU | 124 ±6.2 |  |  | 7:31 | ✔ | 117.72 | FNU | P | 13:17 | This section is for any maintenance (in-field maintenance, probe repair etc.). Please include dates, probe numbers, and any information that is needed to track the incident.
| Turbidity 124 (ICV) | 124 | FNU |  |  |  | ✔ | 123.89 | FNU | 124 ±6.2 |  | P | 7:36
| BP at ODO Cal |  |  |  | 765.556 | mmHg | [8 empty] | 765.556 | mmHg
| Optical DO % |  |  |  | 101.7 | % |  | 100.9 | % | NA |  | NA | 7:38 |  | 100.5 | % | P | 13:22
[6 empty] | DO gain |  | 1.02 | mg/L (table) | 0.87-1.25 |  | P |  |  | 100.5 | %
| Optical DO mg/L (ICV) | 8.85 | mg/L (DL) |  | 21.75 | °C |  | 8.743 |  | ± 0.3 mg/L |  | P | 7:43 |  | 9.05 | mg/L (DL)
| Battery Voltage | [6 empty] | 3.2 | V | ~3V (EXO 3) |  | P |  |  | 20.5 | °C | P | 13:29
[6 empty] | Optional Parameters | [8 empty] | 2.24 | V | P
| Chlorophyll | 0 | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| Chlorophyll |  | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [5 empty] | °C
| Chlorophyll | 0 | µg/L |  |  | µg/L |  |  | µg/L | [6 empty] | µg/L |  |  |  |  | Post-Deployment Notes
| Chlorophyll |  | µg/L |  |  | µg/L |  |  | µg/L | [6 empty] | µg/L |  |  | Turbidity post-cal kind of low. Watch this sensor. May need to be replaced soon or just needed to be cleaned well.
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [5 empty] | °C
| BGA | 0 | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| BGA |  | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [5 empty] | °C
Field Log
Fouling Presence | [9 empty] | Deployment |  |  | Retrieval |  |  |  | Field Notes:
A=algae, AM=amphipods, B=barnacles, BR=bryozoans, C=crabs, E=eggs, F=fish, FA=filamentous algae, G=gastropod, H=hydroids, I=isopod, MD=mud, S=sponges, SI=silt, SL=shell, SP=shrimp, T=tunicates, W=worm, O=other | [8 empty] | Date deploy/retrieval | 2021-11-02 |  |  | 2021-11-23 |  |  | Deployment:
[9 empty] | Time deploy/retrieval | 8:52 |  |  | 10:17 |  |  | Low tide, calm, light winds.
[9 empty] | First/last sample time | 9:00 |  |  | 10:00
| H= heavy, M= moderate, L= light | [7 empty] | Technician initials | TG, KS |  |  | TG, KS, TJ
Mount: 2 Fish; A/H, B/M, Si/M | [9 empty] | Discrete Samples (Required)
[9 empty] | Depth (m) | 0.0634 | m |  | 0.0787 |  | m
Sonde Body: Fish in guard; E/H, Si/M, T/L, W/L | [8 empty] | Temp (°C) | 19.3 | °C |  | 17.5 |  | °C
[9 empty] | DO % | 103.9 | % |  | 94.2 |  | %
Field guard: | [8 empty] | DO mg/L | 8.21 | mg/L |  | 7.94 |  | mg/L
All Sesnors: Si/M, EH, W/L | [8 empty] | SpCond (mS/cm) | 41.065 | mS/cm |  | 33.298 |  | mS/cm | Retrieval:
C/T: G/L | [8 empty] | Salinity (psu) | 26.35 | psu |  | 20.91 |  | psu | Low tide, cold, windy, choppy, brown water, clear skies.
pH: | [8 empty] | pH | 8.2 |  |  | 8.03
DO: | [11 empty] | Optional Samples
Turb: | [8 empty] | pHmV | -76.9 | mV |  | -74.1 |  | mV
Wiper: Bad; fell off | [8 empty] | Conductivity |  | mS/cm |  |  |  | mS/cm
TAL-PE:
Other/Notes:

## APCODE_MMDDYY
Aquatic Preserve Continuous Water Quality Monitoring Calibration Log
|  | File Name | [6 empty] | .csv | [5 empty] | File Name |  |  |  |  | .bin
[7 empty] | Datasonde and Probe Identification Numbers | [11 empty] | Calibration Standards
|  | Sonde Code |  |  | Serial Number |  |  | Serial Number |  | Model Number | [5 empty] | Standard | Lot No. |  | Received | Opened | Exp. Date
| Datasonde | Datasonde |  |  | Sonde No. |  | pH | pH Serial No. |  | pH Model |  |  |  |  | SpCond (mS/cm) | 20
| Vented | Y/N |  |  | Model |  | ODO | ODO Serial No. |  | ODO Model |  |  |  |  | SpCond (mS/cm) | 50
| Nickname |  |  |  | Sonde Model |  | Turb | Turb Serial No. |  | Turb Model |  |  |  |  | pH | 4
| Preserve | AP |  |  |  |  | Cond | Cond Serial No. |  | Cond Model |  |  |  |  | pH | 7
| Station | Station |  |  |  |  | Chl | Chl Serial No. |  | Chl Model |  |  |  |  | pH | 10
[7 empty] | Wiper | Wiper Serial No. |  | Wiper Model |  |  |  |  | Turb (FNU) | 124
|  |  |  | Wiper Replaced |  |  |  |  | Batteries Replaced | [6 empty] | Technician(s) | Technician
|  |  |  | Wiper Parked in Garage |  |  |  |  | Date of Calibration | [6 empty] | Date of Post-Calibration
Pre-deployment Comments/ Failed-calibration values:
| Pre-Deployment (CAL & ICV) | [12 empty] | Post-Deployment (CCV) |  |  |  |  | Programming
|  | Standards |  |  | Before Cal |  |  | Calibrated |  | Specs |  | P/ F | Time |  | Post value |  | P/ F | Time | Interval |  | 15
| Temp |  | °C (Therm) |  |  | °C (DL) |  |  |  | ±0.5(w)+0.1°C | [5 empty] | °C (Therm) |  |  | Duration |  | 365 days
| SpCond in air | [8 empty] | ≤ 0.002 mS/cm | [5 empty] | °C (DL) |  |  | Free memory %
| SpCond | 50 | mS/cm |  |  | mS/cm |  |  | mS/cm | ±5% |  |  |  | ✔ |  | mS/cm |  |  | Memory (days)
[6 empty] | Cell Constant |  |  |  | 0.419-0.519 (w) | [8 empty] | Start time
| SpCond 20 (ICV) | 20 | mS/cm |  |  | ✔ |  |  | mS/cm | ±5% | [8 empty] | Battery %
| pH 7 | 7 | [7 empty] | 7 ±0.2 |  |  |  | ✔ |  |  |  |  | Set clock |  | Y/N
| pH 10 | 10 | [7 empty] | 10 ±0.2 |  |  |  | ✔ |  |  |  |  | End time
[5 empty] | pH 7 |  |  |  | mV | 0 ±50 mV | [5 empty] | mV |  |  |  | Other/notes
[5 empty] | pH 10 |  |  |  | mV | -180 ±50 mV | [5 empty] | mV
[5 empty] | pH 4 |  |  |  | mV | +180 ±50 mV
[5 empty] | pH Slope |  |  | 0 |  | ≥ 165 mV |  |  |  |  | 0
| pH 4 or 10 (ICV) | 4 |  |  |  | ✔ |  |  |  | 4 ±0.2
| BP at Depth Cal |  |  |  |  | mmHg | [9 empty] | mmHg
| Depth |  | Offset |  |  | m |  |  | m | ±5% | [5 empty] | m
| Turbidity | 0 | FNU | [5 empty] | FNU | 0.00 ±0.3 |  |  |  | ✔ |  | FNU |  |  |  | General Maintenance
| Turbidity | 124 | FNU |  |  | FNU |  |  | FNU | 124 ±6.2 |  |  |  | ✔ |  | FNU
| Turbidity 124 (ICV) | 124 | FNU |  |  |  | ✔ |  | FNU | 124 ±6.2
| BP at ODO Cal |  |  |  |  | mmHg | [9 empty] | mmHg
| Optical DO % |  |  |  |  | % |  |  | % | NA |  | NA |  |  |  | %
[6 empty] | DO gain |  |  | mg/L (table) | 0.87-1.25 | [5 empty] | %
| Optical DO mg/L (ICV) |  | mg/L (DL) |  |  | °C |  |  |  | ±0.3 mg/L | [5 empty] | mg/L (DL)
| Battery Voltage | [7 empty] | V | ~6V (EXO 2) | [5 empty] | °C
[6 empty] | Optional Parameters | [9 empty] | V
| Chlorophyll | 0 | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| Chlorophyll |  | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [5 empty] | °C
| Chlorophyll | 0 | µg/L |  |  | µg/L |  |  | µg/L | [6 empty] | µg/L |  |  |  | Post-Deployment Notes
| Chlorophyll |  | µg/L |  |  | µg/L |  |  | µg/L | [6 empty] | µg/L
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [5 empty] | °C
| BGA | 0 | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| BGA |  | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [5 empty] | °C
Field Log
Fouling Presence | [9 empty] | Deployment |  |  | Retrieval |  |  |  | Field Notes:
A=algae, AM=amphipods, B=barnacles, BR=bryozoans, C=crabs, E=eggs, F=fish, FA=filamentous algae, G=gastropod, H=hydroids, I=isopod, MD=mud, S=sponges, SI=silt, SL=shell, SP=shrimp, T=tunicates, W=worm, O=other | [8 empty] | Date deploy/retrieval | [6 empty] | Deployment:
Time deploy/retrieval
First/last sample time
| H= heavy, M= moderate, L= light | [7 empty] | Technician initials
Mount: | [9 empty] | Discrete Samples (Required)
[9 empty] | Depth (m) |  | m |  |  |  | m
Sonde Body: | [8 empty] | Temp (°C) |  | °C |  |  |  | °C
[9 empty] | DO % |  | % |  |  |  | %
Field guard: | [8 empty] | DO mg/L |  | mg/L |  |  |  | mg/L
All Sensors: | [8 empty] | SpCond (mS/cm) |  | mS/cm |  |  |  | mS/cm | Retrieval:
C/T: | [8 empty] | Salinity (psu) |  | psu |  |  |  | psu
pH: | [8 empty] | pH
DO: | [11 empty] | Optional Samples
Turb: | [8 empty] | pHmV |  | mV |  |  |  | mV
Wiper: | [8 empty] | Conductivity |  | mS/cm |  |  |  | mS/cm
TAL-PE:
Other/notes:

## CPAH2_122223
Aquatic Preserve Continuous Water Quality Monitoring Calibration Log
|  | File Name | CPAH2_122223 | [5 empty] | .csv | [5 empty] | File Name | CPAH221E103416_122223_084500 |  |  |  |  | .bin
[7 empty] | Datasonde and Probe Identification Numbers | [12 empty] | Calibration Standards
|  | Sonde Code |  |  | Serial Number |  |  | Serial Number |  | Model Number | [5 empty] | Standard | Lot No. |  |  | Received | Opened | Exp. Date
| Datasonde | 1 |  |  | 21E103416 |  | pH | 21G105935 |  | 577602 |  |  |  |  | SpCond (mS/cm) | 20 | 4203H74 |  |  |  | 4/11/2022 | 3/2024
| Vented | N |  |  | Model |  | ODO | 21G100611 |  | 599100-01 |  |  |  |  | SpCond (mS/cm) | 50 | 2908A58 |  |  |  |  | 7/2021
| Nickname | AH1 |  |  | EXO2 |  | Turb | 21G100846 |  | 599101-01 |  |  |  |  | pH | 4
| Preserve | CPAP |  |  |  |  | Cond | 21G101303 |  | 599827 |  |  |  |  | pH | 7 | 21M2S |  |  |  |  | 12/8/2023
| Station | CPAH2 |  |  |  |  | Chl | [7 empty] | pH | 10 | 22A1T |  |  |  |  | 1/3/2024
[7 empty] | Wiper | 21H105437 |  | 599090-01 |  |  |  |  | Turb (FNU) | 124 | 22D22130237 |  |  |  |  | 4/2023
|  |  |  | Wiper Replaced | Y |  |  |  | Batteries Replaced | N | [5 empty] | Technician(s) | Lena Kury
|  |  |  | Wiper Parked in Garage | Y |  |  |  | Date of Calibration | 2023-12-15 | [5 empty] | Date of Post-Calibration | 2024-01-15
Pre-deployment Comments/ Failed-calibration values:
| Pre-Deployment (CAL & ICV) | [12 empty] | Post-Deployment (CCV) | [5 empty] | Programming
|  | Standards |  |  | Before Cal |  |  | Calibrated |  | Specs |  | P/ F | Time |  | Post value |  | P/ F | Time |  | Interval |  | 15
| Temp | 18.5 | °C (Therm) |  | 18.501 | °C (DL) |  |  |  | ±0.5(w)+0.1°C |  | P | 14:28 |  | 19.6 | °C (Therm) | P | 10:11 |  | Duration |  | 365 days
| SpCond in air |  |  |  | 0.001 |  |  |  |  | ≤ 0.002 mS/cm |  | P |  |  | 19.217 | °C (DL) |  |  |  | Free memory % |  | 100
| SpCond | 50 | mS/cm |  | 50.151 | mS/cm |  | 50 | mS/cm | ±5% |  |  | 14:03 | ✔ | 49.708 | mS/cm | P | 10:19 |  | Memory (days) |  | 24221
[6 empty] | Cell Constant |  | 0.47 |  | 0.419-0.519 (w) |  | P | [7 empty] | Start time |  | 8:45
| SpCond 20 (ICV) | 20 | mS/cm |  |  | ✔ |  | 20.556 | mS/cm | ±5% |  | P | 14:06 | [6 empty] | Battery % |  | 82.5
| pH 7 | 7 |  |  | 6.93 |  |  | 7.01 |  | 7 ±0.2 |  |  | 14:09 | ✔ | 6.97 |  | P | 10:20 |  | Set clock |  | Y
| pH 10 | 10 |  |  | 9.87 |  |  | 10.04 |  | 10 ±0.2 |  |  | 14:13 | ✔ | 10.1 |  | P | 10:21 |  | End time |  | 15:45
[5 empty] | pH 7 |  |  | -51.2 | mV | 0 ±50 mV |  | P |  |  | -48.3 | mV | P |  |  |  | Other/notes
[5 empty] | pH 10 |  |  | -223.6 | mV | -180 ±50 mV |  | P |  |  | -226.7 | mV | P |  |  | BP = 30.28 = 769.112
[5 empty] | pH 4 |  |  |  | mV | +180 ±50 mV |  | P
[5 empty] | pH Slope |  |  | 172.4 |  | ≥ 165 mV |  | P |  |  | 178.4 |  | P
| pH 4 or 10 (ICV) | 10 |  |  |  | ✔ |  | 10.02 |  | 10 ±0.2 |  | P | 14:14
| BP at Depth Cal |  |  |  | 769.112 | mmHg | [8 empty] | 765.81 | mmHg
| Depth | 0.122 | Offset |  | 0.202 | m |  | 0.122 | m | ±5% |  | P | 14:18 |  | 0.087 | m
| Turbidity | 0 | FNU |  | -0.21 |  |  | -0.01 | FNU | 0.00 ±0.3 |  |  | 14:20 | ✔ | 0.26 | FNU | P | 10:28 |  |  | General Maintenance
| Turbidity | 124 | FNU |  | 122.48 | FNU |  | 124.01 | FNU | 124 ±6.2 |  |  | 14:24 | ✔ | 125.1 | FNU | P | 10:30 | Used ProDSS for temp.
| Turbidity 124 (ICV) | 124 | FNU |  |  |  | ✔ | 123.98 | FNU | 124 ±6.2 |  | P | 14:24
| BP at ODO Cal |  |  |  | 769.112 | mmHg | [8 empty] | 765.81 | mmHg
| Optical DO % |  |  |  | 104.2 | % |  | 101 | % | NA |  | NA | 14:41 |  | 100.7 | % | P | 10:08
[6 empty] | DO gain |  | 0.99 | mg/L (table) | 0.87-1.25 |  | P |  |  | 100.8 | %
| Optical DO mg/L (ICV) | 9.45 | mg/L (DL) |  | 18.609 | °C |  | 9.276 |  | ±0.3 mg/L |  | P | 14:42 |  | 9.37 | mg/L (DL)
| Battery Voltage | [6 empty] | 5.6 | V | ~6V (EXO 2) |  | P |  |  | 19.216 | °C | P | 10:11
[6 empty] | Optional Parameters | [8 empty] | 5.42 | V | F
| Chlorophyll | 0 | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| Chlorophyll |  | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [5 empty] | °C
| Chlorophyll | 0 | µg/L |  |  | µg/L |  |  | µg/L | [6 empty] | µg/L |  |  |  |  | Post-Deployment Notes
| Chlorophyll |  | µg/L |  |  | µg/L |  |  | µg/L | [6 empty] | µg/L |  |  | BP = 30.15 = 765.81
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [5 empty] | °C
| BGA | 0 | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| BGA |  | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [5 empty] | °C
Field Log
Fouling Presence | [9 empty] | Deployment |  |  | Retrieval |  |  |  | Field Notes:
A=algae, AM=amphipods, B=barnacles, BR=bryozoans, C=crabs, E=eggs, F=fish, FA=filamentous algae, G=gastropod, H=hydroids, I=isopod, MD=mud, S=sponges, SI=silt, SL=shell, SP=shrimp, T=tunicates, W=worm, O=other | [8 empty] | Date deploy/retrieval | 2023-12-22 |  |  | 2024-01-11 |  |  | Deployment:
[9 empty] | Time deploy/retrieval | 10:35 |  |  | 13:49 |  |  | Turb: 0.78
[9 empty] | First/last sample time | 10:45 |  |  | 13:45
| H= heavy, M= moderate, L= light | [7 empty] | Technician initials | LK |  |  | LK
Mount: | [9 empty] | Discrete Samples (Required)
[9 empty] | Depth (m) | 1.6 | m |  | 2.1 |  | m
Sonde Body: A, MD - L | [8 empty] | Temp (°C) | 12.4 | °C |  | 13.2 |  | °C
[9 empty] | DO % | 97.6 | % |  | 104.9 |  | %
Field guard: A, MD - L | [8 empty] | DO mg/L | 8.54 | mg/L |  | 9.07 |  | mg/L
All Sensors: | [8 empty] | SpCond (mS/cm) | 48428 | mS/cm |  | 47598 |  | mS/cm | Retrieval:
C/T: A, MD - L | [8 empty] | Salinity (psu) | 31.54 | psu |  | 30.96 |  | psu | Turb: 2.58
pH: A, MD - L | [8 empty] | pH | 8.07 |  |  | 8.05
DO: A, MD - L | [11 empty] | Optional Samples
Turb: A, MD - L | [8 empty] | pHmV |  | mV |  |  |  | mV
Wiper: stayed on! | [8 empty] | Conductivity |  | mS/cm |  |  |  | mS/cm
TAL-PE:
Other/notes: nothing on faces

## CPAH2_011124
Aquatic Preserve Continuous Water Quality Monitoring Calibration Log
|  | File Name | CPAH2_011124 | [5 empty] | .csv | [5 empty] | File Name | CPAH221E103416_011124_111500 |  |  |  |  | .bin
[7 empty] | Datasonde and Probe Identification Numbers | [12 empty] | Calibration Standards
|  | Sonde Code |  |  | Serial Number |  |  | Serial Number |  | Model Number | [5 empty] | Standard | Lot No. |  |  | Received | Opened | Exp. Date
| Datasonde | 2 |  |  | 21E103418 |  | pH | 21G105084 |  | 577602 |  |  |  |  | SpCond (mS/cm) | 20 | 4203H74 |  |  |  | 4/11/2022 | 3/2024
| Vented | N |  |  | Model |  | ODO | 21G100610 |  | 599100-01 |  |  |  |  | SpCond (mS/cm) | 50 | 2908A58 |  |  |  |  | 7/2021
| Nickname | AH2 |  |  | EXO2 |  | Turb | 21G100848 |  | 599101-01 |  |  |  |  | pH | 4
| Preserve | CPAP |  |  |  |  | Cond | 23F106588 |  | 599827 |  |  |  |  | pH | 7 | 21M2S |  |  |  |  | 12/8/2023
| Station | CPAH2 |  |  |  |  | Chl | [7 empty] | pH | 10 | 22A1T |  |  |  |  | 1/3/2024
[7 empty] | Wiper | 21H105435 |  | 599090-01 |  |  |  |  | Turb (FNU) | 124 | 22D22130237 |  |  |  |  | 4/2023
|  |  |  | Wiper Replaced | Y |  |  |  | Batteries Replaced | Y | [5 empty] | Technician(s) | Lena Kury
|  |  |  | Wiper Parked in Garage | Y |  |  |  | Date of Calibration | 2024-01-08 | [5 empty] | Date of Post-Calibration | 2024-02-05
Pre-deployment Comments/ Failed-calibration values:
| Pre-Deployment (CAL & ICV) | [12 empty] | Post-Deployment (CCV) | [5 empty] | Programming
|  | Standards |  |  | Before Cal |  |  | Calibrated |  | Specs |  | P/ F | Time |  | Post value |  | P/ F | Time |  | Interval |  | 15
| Temp | 17.8 | °C (Therm) |  | 17.712 | °C (DL) |  |  |  | ±0.5(w)+0.1°C |  | P | 10:48 |  | 19.773 | °C (Therm) | P | 11:20 |  | Duration |  | 365 days
| SpCond in air |  |  |  | 0 |  |  |  |  | ≤ 0.002 mS/cm |  | P |  |  | 19.861 | °C (DL) |  |  |  | Free memory % |  | 100
| SpCond | 50 | mS/cm |  | 50.004 | mS/cm |  | 50 | mS/cm | ±5% |  |  | 11:12 | ✔ | 49.348 | mS/cm | P | 11:44 |  | Memory (days) |  | 24221
[6 empty] | Cell Constant |  | 0.47 |  | 0.419-0.519 (w) |  | P | [7 empty] | Start time |  | 11:15
| SpCond 20 (ICV) | 20 | mS/cm |  |  | ✔ |  | 20.563 | mS/cm | ±5% |  | P | 11:14 | [6 empty] | Battery % |  | 97.6
| pH 7 | 7 |  |  | 6.9 |  |  | 7.02 |  | 7 ±0.2 |  |  | 11:16 | ✔ | 7.09 |  | P | 11:46 |  | Set clock |  | Y
| pH 10 | 10 |  |  | 9.89 |  |  | 10.05 |  | 10 ±0.2 |  |  | 11:18 | ✔ | 10.18 |  | P | 11:49 |  | End time |  | 14:00
[5 empty] | pH 7 |  |  | -49.2 | mV | 0 ±50 mV |  | P |  |  | -53.5 | mV | P |  |  |  | Other/notes
[5 empty] | pH 10 |  |  | -227.5 | mV | -180 ±50 mV |  | P |  |  | -235.4 | mV | P |  |  | BP = 30.17 = 766.318
[5 empty] | pH 4 |  |  |  | mV | +180 ±50 mV |  | P
[5 empty] | pH Slope |  |  | 178.3 |  | ≥ 165 mV |  | P |  |  | 181.9 |  | P
| pH 4 or 10 (ICV) | 10 |  |  |  | ✔ |  | 10.01 |  | 10 ±0.2 |  | P | 11:19
| BP at Depth Cal |  |  |  | 766.318 | mmHg | [8 empty] | 752.094 | mmHg
| Depth | 0.082 | Offset |  | 0.12 | m |  | 0.082 | m | ±5% |  | P | 11:06 |  | -0.14 | m
| Turbidity | 0 | FNU |  | -0.43 |  |  | 0 | FNU | 0.00 ±0.3 |  |  | 11:23 | ✔ | 0.01 | FNU | P | 11:54 |  |  | General Maintenance
| Turbidity | 124 | FNU |  | 121.31 | FNU |  | 123.98 | FNU | 124 ±6.2 |  |  | 11:25 | ✔ | 121.14 | FNU | P | 11:55 | Used ProDSS for temp.
| Turbidity 124 (ICV) | 124 | FNU |  |  |  | ✔ | 123.96 | FNU | 124 ±6.2 |  | P | 11:25
| BP at ODO Cal |  |  |  | 766.318 | mmHg | [8 empty] | 752.094 | mmHg
| Optical DO % |  |  |  | 101.9 | % |  | 100.8 | % | NA |  | NA | 11:03 |  | 99.6 | % | P | 11:16
[6 empty] | DO gain |  | 1.02 | mg/L (table) | 0.87-1.25 |  | P |  |  | 99.1 | %
| Optical DO mg/L (ICV) | 9.56 | mg/L (DL) |  | 17.825 | °C |  | 9.467 |  | ±0.3 mg/L |  | P | 11:04 |  | 8.86 | mg/L (DL)
| Battery Voltage | [6 empty] | 5.91 | V | ~6V (EXO 2) |  | P |  |  | 19.849 | °C | P | 11:16
[6 empty] | Optional Parameters | [8 empty] | 5.72 | V | P
| Chlorophyll | 0 | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| Chlorophyll |  | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [5 empty] | °C
| Chlorophyll | 0 | µg/L |  |  | µg/L |  |  | µg/L | [6 empty] | µg/L |  |  |  |  | Post-Deployment Notes
| Chlorophyll |  | µg/L |  |  | µg/L |  |  | µg/L | [6 empty] | µg/L |  |  | BP = 29.61 = 752.094
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [5 empty] | °C
| BGA | 0 | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| BGA |  | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [5 empty] | °C
Field Log
Fouling Presence | [9 empty] | Deployment |  |  | Retrieval |  |  |  | Field Notes:
A=algae, AM=amphipods, B=barnacles, BR=bryozoans, C=crabs, E=eggs, F=fish, FA=filamentous algae, G=gastropod, H=hydroids, I=isopod, MD=mud, S=sponges, SI=silt, SL=shell, SP=shrimp, T=tunicates, W=worm, O=other | [8 empty] | Date deploy/retrieval | 2024-01-11 |  |  | 2024-02-02 |  |  | Deployment:
[9 empty] | Time deploy/retrieval | 13:53 |  |  | 11:56 |  |  | Turb: 2.58
[9 empty] | First/last sample time | 14:00 |  |  | 11:45
| H= heavy, M= moderate, L= light | [7 empty] | Technician initials | LK |  |  | LK
Mount: | [9 empty] | Discrete Samples (Required)
[9 empty] | Depth (m) | 2.1 | m |  | 2 |  | m
Sonde Body: | [8 empty] | Temp (°C) | 13.2 | °C |  | 15 |  | °C
[9 empty] | DO % | 104.9 | % |  | 97.2 |  | %
Field guard: MD, A - L | [8 empty] | DO mg/L | 9.07 | mg/L |  | 8.11 |  | mg/L
All Sensors: | [8 empty] | SpCond (mS/cm) | 47598 | mS/cm |  | 47025 |  | mS/cm | Retrieval:
C/T: MD, A - L | [8 empty] | Salinity (psu) | 30.96 | psu |  | 30.6 |  | psu | Turb: 1.38
pH: MD, A - L | [8 empty] | pH | 8.05 |  |  | 7.98
DO: MD, A - L | [11 empty] | Optional Samples
Turb: MD, A - L | [8 empty] | pHmV |  | mV |  |  |  | mV
Wiper: stayed on! | [8 empty] | Conductivity |  | mS/cm |  |  |  | mS/cm
TAL-PE:
Other/notes: nothing on faces

## CPAH2_020224
Aquatic Preserve Continuous Water Quality Monitoring Calibration Log
|  | File Name | CPAH2_020224 | [5 empty] | .csv | [5 empty] | File Name | CPAH221E103416_020224_093000 |  |  |  |  | .bin
[7 empty] | Datasonde and Probe Identification Numbers | [12 empty] | Calibration Standards
|  | Sonde Code |  |  | Serial Number |  |  | Serial Number |  | Model Number | [5 empty] | Standard | Lot No. |  |  | Received | Opened | Exp. Date
| Datasonde | 1 |  |  | 21E103416 |  | pH | 21G105935 |  | 577602 |  |  |  |  | SpCond (mS/cm) | 20 | 4203H74 |  |  |  | 4/11/2022 | 3/2024
| Vented | N |  |  | Model |  | ODO | 21G100611 |  | 599100-01 |  |  |  |  | SpCond (mS/cm) | 50 | 2908A58 |  |  |  |  | 7/2021
| Nickname | AH1 |  |  | EXO2 |  | Turb | 21G100846 |  | 599101-01 |  |  |  |  | pH | 4
| Preserve | CPAP |  |  |  |  | Cond | 21G101303 |  | 599827 |  |  |  |  | pH | 7 | 21M2S |  |  |  |  | 12/8/2023
| Station | CPAH2 |  |  |  |  | Chl | [7 empty] | pH | 10 | 22A1T |  |  |  |  | 1/3/2024
[7 empty] | Wiper | 21H105437 |  | 599090-01 |  |  |  |  | Turb (FNU) | 124 | 22D22130237 |  |  |  |  | 4/2023
|  |  |  | Wiper Replaced | Y |  |  |  | Batteries Replaced | Y | [5 empty] | Technician(s) | Lena Kury
|  |  |  | Wiper Parked in Garage | Y |  |  |  | Date of Calibration | 2024-02-01 | [5 empty] | Date of Post-Calibration | 2024-03-15
Pre-deployment Comments/ Failed-calibration values:
| Pre-Deployment (CAL & ICV) | [12 empty] | Post-Deployment (CCV) | [5 empty] | Programming
|  | Standards |  |  | Before Cal |  |  | Calibrated |  | Specs |  | P/ F | Time |  | Post value |  | P/ F | Time |  | Interval |  | 15
| Temp | 16.561 | °C (Therm) |  | 16.555 | °C (DL) |  |  |  | ±0.5(w)+0.1°C |  | P | 15:21 |  | 19.795 | °C (Therm) | P | 15:45 |  | Duration |  | 365 days
| SpCond in air |  |  |  | 0.001 |  |  |  |  | ≤ 0.002 mS/cm |  | P |  |  | 19.907 | °C (DL) |  |  |  | Free memory % |  | 100
| SpCond | 50 | mS/cm |  | 49.712 | mS/cm |  | 50 | mS/cm | ±5% |  |  | 15:31 | ✔ | 49.985 | mS/cm | P | 15:54 |  | Memory (days) |  | 24221
[6 empty] | Cell Constant |  | 0.47 |  | 0.419-0.519 (w) |  | P | [7 empty] | Start time |  | 9:30
| SpCond 20 (ICV) | 20 | mS/cm |  |  | ✔ |  | 20.733 | mS/cm | ±5% |  | P | 15:34 | [6 empty] | Battery % |  | 100%
| pH 7 | 7 |  |  | 6.99 |  |  | 7.02 |  | 7 ±0.2 |  |  | 15:37 | ✔ | 7.25 |  | F | 15:55 |  | Set clock |  | Y
| pH 10 | 10 |  |  | 10.07 |  |  | 10.05 |  | 10 ±0.2 |  |  | 15:39 | ✔ | 10.2 |  | P | 15:56 |  | End time |  | 14:30
[5 empty] | pH 7 |  |  | -49.9 | mV | 0 ±50 mV |  | P |  |  | -63.1 | mV | P |  |  |  | Other/notes
[5 empty] | pH 10 |  |  | -225 | mV | -180 ±50 mV |  | P |  |  | -233.9 | mV | P |  |  | BP = 30.13 = 765.302
[5 empty] | pH 4 |  |  |  | mV | +180 ±50 mV |  | P
[5 empty] | pH Slope |  |  | 175.1 |  | ≥ 165 mV |  | P |  |  | 170.8 |  | P
| pH 4 or 10 (ICV) | 10 |  |  |  | ✔ |  | 10.03 |  | 10 ±0.2 |  | P | 15:40
| BP at Depth Cal |  |  |  | 765.302 | mmHg | [8 empty] | 763.016 | mmHg
| Depth | 0.068 | Offset |  | 0.031 | m |  | 0.068 | m | ±5% |  | P | 15:26 |  | 0.081 | m
| Turbidity | 0 | FNU |  | 0.1 |  |  | 0 | FNU | 0.00 ±0.3 |  |  | 15:42 | ✔ | 0.07 | FNU | P | 15:57 |  |  | General Maintenance
| Turbidity | 124 | FNU |  | 124.02 | FNU |  | 124 | FNU | 124 ±6.2 |  |  | 15:44 | ✔ | 122.66 | FNU | P | 15:59
| Turbidity 124 (ICV) | 124 | FNU |  |  |  | ✔ | 124.01 | FNU | 124 ±6.2 |  | P | 15:45
| BP at ODO Cal |  |  |  | 765.302 | mmHg | [8 empty] | 763.96 | mmHg
| Optical DO % |  |  |  | 102.1 | % |  | 100.7 | % | NA |  | NA | 15:23 |  | 99.3 | % | P | 15:45
[6 empty] | DO gain |  | 0.98 | mg/L (table) | 0.87-1.25 |  | P |  |  | 99.3 | %
| Optical DO mg/L (ICV) | 9.8 | mg/L (DL) |  | 16.596 | °C |  | 9.665 |  | ±0.3 mg/L |  | P | 15:24 |  | 8.47 | mg/L (DL)
| Battery Voltage | [6 empty] | 6.37 | V | ~6V (EXO 2) |  | P |  |  | 19.909 | °C | P | 15:45
[6 empty] | Optional Parameters | [8 empty] | 5.79 | V | P
| Chlorophyll | 0 | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| Chlorophyll |  | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [5 empty] | °C
| Chlorophyll | 0 | µg/L |  |  | µg/L |  |  | µg/L | [6 empty] | µg/L |  |  |  |  | Post-Deployment Notes
| Chlorophyll |  | µg/L |  |  | µg/L |  |  | µg/L | [6 empty] | µg/L |  |  | BP = 30.04 = 763.016
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [5 empty] | °C
| BGA | 0 | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| BGA |  | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [5 empty] | °C
Field Log
Fouling Presence | [9 empty] | Deployment |  |  | Retrieval |  |  |  | Field Notes:
A=algae, AM=amphipods, B=barnacles, BR=bryozoans, C=crabs, E=eggs, F=fish, FA=filamentous algae, G=gastropod, H=hydroids, I=isopod, MD=mud, S=sponges, SI=silt, SL=shell, SP=shrimp, T=tunicates, W=worm, O=other | [8 empty] | Date deploy/retrieval | 2024-02-02 |  |  | 2024-03-12 |  |  | Deployment:
[9 empty] | Time deploy/retrieval | 11:58 |  |  | 11:53 |  |  | Turb: 1.38
[9 empty] | First/last sample time | 12:00 |  |  | 11:45
| H= heavy, M= moderate, L= light | [7 empty] | Technician initials | LK |  |  | LK
Mount: | [9 empty] | Discrete Samples (Required)
[9 empty] | Depth (m) | 2 | m |  | 1.9 |  | m
Sonde Body: MD - L | [8 empty] | Temp (°C) | 15 | °C |  | 19.1 |  | °C
[9 empty] | DO % | 97.2 | % |  | 103 |  | %
Field guard: MD, A, C, anemones - M | [8 empty] | DO mg/L | 8.11 | mg/L |  | 8.02 |  | mg/L
All Sensors: | [8 empty] | SpCond (mS/cm) | 47025 | mS/cm |  | 45527 |  | mS/cm | Retrieval:
C/T: B, A, MD - L | [8 empty] | Salinity (psu) | 30.6 | psu |  | 29.55 |  | psu | Turb: 3.27
pH: B, A, MD - L | [8 empty] | pH | 7.98 |  |  | 7.98
DO: B, A, MD - L | [11 empty] | Optional Samples
Turb: B, A, MD - L | [8 empty] | pHmV |  | mV |  |  |  | mV
Wiper: stayed on! | [8 empty] | Conductivity |  | mS/cm |  |  |  | mS/cm
TAL-PE:
Other/notes: small barnacle on edge of C/T probe face

## CPAH2_031224
Aquatic Preserve Continuous Water Quality Monitoring Calibration Log
|  | File Name | CPAH2_031224 | [5 empty] | .csv | [5 empty] | File Name | CPAH221E103416_031224_091500 |  |  |  |  | .bin
[7 empty] | Datasonde and Probe Identification Numbers | [12 empty] | Calibration Standards
|  | Sonde Code |  |  | Serial Number |  |  | Serial Number |  | Model Number | [5 empty] | Standard | Lot No. |  |  | Received | Opened | Exp. Date
| Datasonde | 2 |  |  | 21E103417 |  | pH | 23G101728 |  | 577602 |  |  |  |  | SpCond (mS/cm) | 20 | 4203H74 |  |  |  | 4/11/2022 | 3/2024
| Vented | N |  |  | Model |  | ODO | 21G100608 |  | 599100-01 |  |  |  |  | SpCond (mS/cm) | 50 | 2908A58 |  |  |  |  | 7/2021
| Nickname | AH2 |  |  | EXO2 |  | Turb | 21G100845 |  | 599101-01 |  |  |  |  | pH | 4
| Preserve | CPAP |  |  |  |  | Cond | 23E103845 |  | 599827 |  |  |  |  | pH | 7 | 21M2S |  |  |  |  | 12/8/2023
| Station | CPAH2 |  |  |  |  | Chl | [7 empty] | pH | 10 | 22A1T |  |  |  |  | 1/3/2024
[7 empty] | Wiper | 21H105434 |  | 599090-01 |  |  |  |  | Turb (FNU) | 124 | 22D22130237 |  |  |  |  | 4/2023
|  |  |  | Wiper Replaced | Y |  |  |  | Batteries Replaced | Y | [5 empty] | Technician(s) | Lena Kury
|  |  |  | Wiper Parked in Garage | Y |  |  |  | Date of Calibration | 2024-02-22 | [5 empty] | Date of Post-Calibration | 2024-04-23
| Pre-deployment Comments/ Failed-calibration values: | Back-up sonde; usual one in for maintenance
| Pre-Deployment (CAL & ICV) | [12 empty] | Post-Deployment (CCV) | [5 empty] | Programming
|  | Standards |  |  | Before Cal |  |  | Calibrated |  | Specs |  | P/ F | Time |  | Post value |  | P/ F | Time |  | Interval |  | 15
| Temp | 19.824 | °C (Therm) |  | 19.967 | °C (DL) |  |  |  | ±0.5(w)+0.1°C |  | P | 15:31 |  | 19.375 | °C (Therm) | P | 9:34 |  | Duration |  | 365 days
| SpCond in air |  |  |  | 0 |  |  |  |  | ≤ 0.002 mS/cm |  | P |  |  | 19.408 | °C (DL) |  |  |  | Free memory % |  | 100
| SpCond | 50 | mS/cm |  | 49.717 | mS/cm |  | 50 | mS/cm | ±5% |  |  | 15:43 | ✔ | 50.056 | mS/cm | P | 10:02 |  | Memory (days) |  | 24221
[6 empty] | Cell Constant |  | 0.47 |  | 0.419-0.519 (w) |  | P | [7 empty] | Start time |  | 9:15
| SpCond 20 (ICV) | 20 | mS/cm |  |  | ✔ |  | 20.627 | mS/cm | ±5% |  | P | 15:45 | [6 empty] | Battery % |  | 100%
| pH 7 | 7 |  |  | 7.3 |  |  | 7.02 |  | 7 ±0.2 |  |  | 15:47 | ✔ | 7.2 |  | P | 10:04 |  | Set clock |  | Y
| pH 10 | 10 |  |  | 10.31 |  |  | 10.05 |  | 10 ±0.2 |  |  | 15:50 | ✔ | 10.22 |  | F | 10:07 |  | End time |  | 15:30
[5 empty] | pH 7 |  |  | -17.6 | mV | 0 ±50 mV |  | P |  |  | -28.2 | mV | P |  |  |  | Other/notes
[5 empty] | pH 10 |  |  | -192.7 | mV | -180 ±50 mV |  | P |  |  | -203.4 | mV | P |  |  | BP = 30.12 = 765.048
[5 empty] | pH 4 |  |  |  | mV | +180 ±50 mV |  | P | [7 empty] | DO calibrated on 3/11/2024; BP = 30.23 = 767.842
[5 empty] | pH Slope |  |  | 175.1 |  | ≥ 165 mV |  | P |  |  | 175.2 |  | P
| pH 4 or 10 (ICV) | 10 |  |  |  | ✔ |  | 10.03 |  | 10 ±0.2 |  | P | 15:51
| BP at Depth Cal |  |  |  | 765.048 | mmHg | [8 empty] | 767.588 | mmHg
| Depth | 0.068 | Offset |  | 10.392 | m |  | 0.068 | m | ±5% |  | P | 15:53 |  | 0.15 | m
| Turbidity | 0 | FNU |  | -0.17 |  |  | 0 | FNU | 0.00 ±0.3 |  |  | 16:00 | ✔ | 0.08 | FNU | P | 10:09 |  |  | General Maintenance
| Turbidity | 124 | FNU |  | 123.67 | FNU |  | 124 | FNU | 124 ±6.2 |  |  | 16:03 | ✔ | 124.23 | FNU | P | 10:13 | Batteries changed 3/12/2024; ~6.41 V
| Turbidity 124 (ICV) | 124 | FNU |  |  |  | ✔ | 123.88 | FNU | 124 ±6.2 |  | P | 16:04
| BP at ODO Cal |  |  |  | 767.842 | mmHg | [8 empty] | 767.588 | mmHg
| Optical DO % |  |  |  | 97.8 | % |  | 100.9 | % | NA |  | NA | 14:27 |  | 99.5 | % | P | 9:45
[6 empty] | DO gain |  | 1.03 | mg/L (table) | 0.87-1.25 |  | P |  |  | 99.3 | %
| Optical DO mg/L (ICV) | 9.32 | mg/L (DL) |  | 19.186 | °C |  | 9.276 |  | ±0.3 mg/L |  | P | 14:30 |  | 9.15 | mg/L (DL)
| Battery Voltage | [6 empty] | 6.41 | V | ~6V (EXO 2) |  | P |  |  | 19.413 | °C | P | 9:45
[6 empty] | Optional Parameters | [8 empty] | 5.87 | V | P
| Chlorophyll | 0 | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| Chlorophyll |  | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [5 empty] | °C
| Chlorophyll | 0 | µg/L |  |  | µg/L |  |  | µg/L | [6 empty] | µg/L |  |  |  |  | Post-Deployment Notes
| Chlorophyll |  | µg/L |  |  | µg/L |  |  | µg/L | [6 empty] | µg/L |  |  | BP = 30.22 = 767.588
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [5 empty] | °C
| BGA | 0 | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| BGA |  | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [5 empty] | °C
Field Log
Fouling Presence | [9 empty] | Deployment |  |  | Retrieval |  |  |  | Field Notes:
A=algae, AM=amphipods, B=barnacles, BR=bryozoans, C=crabs, E=eggs, F=fish, FA=filamentous algae, G=gastropod, H=hydroids, I=isopod, MD=mud, S=sponges, SI=silt, SL=shell, SP=shrimp, T=tunicates, W=worm, O=other | [8 empty] | Date deploy/retrieval | 2024-03-12 |  |  | 2024-04-17 |  |  | Deployment:
[9 empty] | Time deploy/retrieval | 11:56 |  |  | 10:41 |  |  | Turb: 3.27
[9 empty] | First/last sample time | 12:00 |  |  | 10:30
| H= heavy, M= moderate, L= light | [7 empty] | Technician initials | LK |  |  | LK
Mount: | [9 empty] | Discrete Samples (Required)
[9 empty] | Depth (m) | 1.9 | m |  | 2.3 |  | m
Sonde Body: MD - L | [8 empty] | Temp (°C) | 19.1 | °C |  | 23.7 |  | °C
[9 empty] | DO % | 103 | % |  | 93.1 |  | %
Field guard: MD, C, A - M | [8 empty] | DO mg/L | 8.02 | mg/L |  | 6.65 |  | mg/L
All Sensors: | [8 empty] | SpCond (mS/cm) | 45527 | mS/cm |  | 45646 |  | mS/cm | Retrieval:
C/T: MD, T, A, B, C - M | [8 empty] | Salinity (psu) | 29.55 | psu |  | 29.59 |  | psu | Turb: 6.94
pH: MD, T, A, B, C - M | [8 empty] | pH | 7.98 |  |  | 7.94
DO: MD, T, A, B, C - M | [11 empty] | Optional Samples
Turb: MD, T, A, B, C - M | [8 empty] | pHmV |  | mV |  |  |  | mV
Wiper: stayed on | [8 empty] | Conductivity |  | mS/cm |  |  |  | mS/cm
TAL-PE:
Other/notes: Nothing on probe faces

## CPAH2_041724
Aquatic Preserve Continuous Water Quality Monitoring Calibration Log
|  | File Name | CPAH2_041724 | [5 empty] | .csv | [5 empty] | File Name | CPAH221E103416_041724_074500 |  |  |  |  | .bin
[7 empty] | Datasonde and Probe Identification Numbers | [12 empty] | Calibration Standards
|  | Sonde Code |  |  | Serial Number |  |  | Serial Number |  | Model Number | [5 empty] | Standard | Lot No. |  |  | Received | Opened | Exp. Date
| Datasonde | 2 |  |  | 21E103418 |  | pH | 21G105084 |  | 577602 |  |  |  |  | SpCond (mS/cm) | 20 | 4203H74 |  |  |  | 4/11/2022 | 3/2024
| Vented | N |  |  | Model |  | ODO | 21G100610 |  | 599100-01 |  |  |  |  | SpCond (mS/cm) | 50 | 2908A58 |  |  |  |  | 7/2021
| Nickname | AH2 |  |  | EXO2 |  | Turb | 21G100848 |  | 599101-01 |  |  |  |  | pH | 4
| Preserve | CPAP |  |  |  |  | Cond | 23F106588 |  | 599827 |  |  |  |  | pH | 7 | 21M2S |  |  |  |  | 12/8/2023
| Station | CPAH2 |  |  |  |  | Chl | [7 empty] | pH | 10 | 22A1T |  |  |  |  | 1/3/2024
[7 empty] | Wiper | 21H105435 |  | 599090-01 |  |  |  |  | Turb (FNU) | 124 | 22D22130237 |  |  |  |  | 4/2023
|  |  |  | Wiper Replaced | Y |  |  |  | Batteries Replaced | N | [5 empty] | Technician(s) | Lena Kury
|  |  |  | Wiper Parked in Garage | Y |  |  |  | Date of Calibration | 2024-04-11 | [5 empty] | Date of Post-Calibration | 2024-05-23
Pre-deployment Comments/ Failed-calibration values:
| Pre-Deployment (CAL & ICV) | [12 empty] | Post-Deployment (CCV) | [5 empty] | Programming
|  | Standards |  |  | Before Cal |  |  | Calibrated |  | Specs |  | P/ F | Time |  | Post value |  | P/ F | Time |  | Interval |  | 15
| Temp | 20.372 | °C (Therm) |  | 20.39 | °C (DL) |  |  |  | ±0.5(w)+0.1°C |  | P | 13:35 |  | 19.598 | °C (Therm) | P | 14:17 |  | Duration |  | 365 days
| SpCond in air |  |  |  | 0 |  |  |  |  | ≤ 0.002 mS/cm |  | P |  |  | 19.598 | °C (DL) |  |  |  | Free memory % |  | 100
| SpCond | 50 | mS/cm |  | 48.335 | mS/cm |  | 50 | mS/cm | ±5% |  |  | 13:41 | ✔ | 49.784 | mS/cm | P | 14:34 |  | Memory (days) |  | 24221
[6 empty] | Cell Constant |  | 0.48 |  | 0.419-0.519 (w) |  | P | [7 empty] | Start time
| SpCond 20 (ICV) | 20 | mS/cm |  |  | ✔ |  | 20.56 | mS/cm | ±5% |  | P | 13:43 | [6 empty] | Battery % |  | 93.9
| pH 7 | 7 |  |  | 7.2 |  |  | 7.01 |  | 7 ±0.2 |  |  | 13:50 | ✔ | 7 |  | P | 14:35 |  | Set clock |  | Y
| pH 10 | 10 |  |  | 10.12 |  |  | 10.03 |  | 10 ±0.2 |  |  | 13:54 | ✔ | 10.01 |  | P | 14:37 |  | End time
[5 empty] | pH 7 |  |  | -21.8 | mV | 0 ±50 mV |  | P |  |  | -20.6 | mV | P |  |  |  | Other/notes
[5 empty] | pH 10 |  |  | -194.5 | mV | -180 ±50 mV |  | P |  |  | -193 | mV | P |  |  | BP = 29.69 = 754.126
[5 empty] | pH 4 |  |  |  | mV | +180 ±50 mV |  | P
[5 empty] | pH Slope |  |  | 172.7 |  | ≥ 165 mV |  | P |  |  | 172.4 |  | P
| pH 4 or 10 (ICV) | 10 |  |  |  | ✔ |  | 10 |  | 10 ±0.2 |  | P | 13:54
| BP at Depth Cal |  |  |  | 754.126 | mmHg | [8 empty] | 762 | mmHg
| Depth | -0.082 | Offset |  | 0.27 | m |  | -0.082 | m | ±5% |  | P | 13:38 |  | 0.005 | m
| Turbidity | 0 | FNU |  | 0.02 |  |  | 0 | FNU | 0.00 ±0.3 |  |  | 13:57 | ✔ | 0.22 | FNU | P | 14:39 |  |  | General Maintenance
| Turbidity | 124 | FNU |  | 127.1 | FNU |  | 124 | FNU | 124 ±6.2 |  |  | 14:01 | ✔ | 122.92 | FNU | P | 14:41
| Turbidity 124 (ICV) | 124 | FNU |  |  |  | ✔ | 123.99 | FNU | 124 ±6.2 |  | P | 14:01
| BP at ODO Cal |  |  |  | 754.126 | mmHg | [8 empty] | 762 | mmHg
| Optical DO % |  |  |  | 104.5 | % |  | 99.2 | % | NA |  | NA | 13:36 |  | 101.3 | % | P | 14:16
[6 empty] | DO gain |  | 1.03 | mg/L (table) | 0.87-1.25 |  | P |  |  | 101.4 | %
| Optical DO mg/L (ICV) | 8.94 | mg/L (DL) |  | 20.387 | °C |  | 9.092 |  | ±0.3 mg/L |  | P | 13:37 |  | 9.23 | mg/L (DL)
| Battery Voltage | [6 empty] | 5.78 | V | ~6V (EXO 2) |  | P |  |  | 19.599 | °C | P | 14:16
[6 empty] | Optional Parameters | [8 empty] | 5.45 | V | P
| Chlorophyll | 0 | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| Chlorophyll |  | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [5 empty] | °C
| Chlorophyll | 0 | µg/L |  |  | µg/L |  |  | µg/L | [6 empty] | µg/L |  |  |  |  | Post-Deployment Notes
| Chlorophyll |  | µg/L |  |  | µg/L |  |  | µg/L | [6 empty] | µg/L |  |  | BP = 30.00 = 762.00
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [5 empty] | °C
| BGA | 0 | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| BGA |  | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [5 empty] | °C
Field Log
Fouling Presence | [9 empty] | Deployment |  |  | Retrieval |  |  |  | Field Notes:
A=algae, AM=amphipods, B=barnacles, BR=bryozoans, C=crabs, E=eggs, F=fish, FA=filamentous algae, G=gastropod, H=hydroids, I=isopod, MD=mud, S=sponges, SI=silt, SL=shell, SP=shrimp, T=tunicates, W=worm, O=other | [8 empty] | Date deploy/retrieval | 2024-04-17 |  |  | 2024-05-20 |  |  | Deployment:
[9 empty] | Time deploy/retrieval | 10:43 |  |  | 9:41 |  |  | Turb: 6.94
[9 empty] | First/last sample time | 10:45 |  |  | 9:30
| H= heavy, M= moderate, L= light | [7 empty] | Technician initials | LK |  |  | LK
Mount: | [9 empty] | Discrete Samples (Required)
[9 empty] | Depth (m) | 2.3 | m |  | 2.4 |  | m
Sonde Body: MD - L | [8 empty] | Temp (°C) | 23.7 | °C |  | 26.5 |  | °C
[9 empty] | DO % | 93.1 | % |  | 96.3 |  | %
Field guard: C, T, A, MD - L | [8 empty] | DO mg/L | 6.65 | mg/L |  | 6.54 |  | mg/L
All Sensors: | [8 empty] | SpCond (mS/cm) | 45646 | mS/cm |  | 46274 |  | mS/cm | Retrieval:
C/T: T, MD, A - L | [8 empty] | Salinity (psu) | 29.59 | psu |  | 29.98 |  | psu | Turb: 6.49
pH: T, MD, A - L | [8 empty] | pH | 7.94 |  |  | 8
DO: T, MD, A - L | [11 empty] | Optional Samples
Turb: MD, A - L | [8 empty] | pHmV |  | mV |  |  |  | mV
Wiper: stayed on! | [8 empty] | Conductivity |  | mS/cm |  |  |  | mS/cm
TAL-PE:
Other/notes: some T on depth ports

## CPAH2_052024
Aquatic Preserve Continuous Water Quality Monitoring Calibration Log
|  | File Name | CPAH2_052024 | [5 empty] | .csv | [5 empty] | File Name | CPAH221E103416_052024_ |  |  |  |  | .bin
[7 empty] | Datasonde and Probe Identification Numbers | [12 empty] | Calibration Standards
|  | Sonde Code |  |  | Serial Number |  |  | Serial Number |  | Model Number | [5 empty] | Standard | Lot No. |  |  | Received | Opened | Exp. Date
| Datasonde | 1 |  |  | 21E103417 |  | pH | 23G101728 |  | 577602 |  |  |  |  | SpCond (mS/cm) | 20 | 4203H74 |  |  |  | 4/11/2022 | 3/2024
| Vented | N |  |  | Model |  | ODO | 21G100608 |  | 599100-01 |  |  |  |  | SpCond (mS/cm) | 50 | 2908A58 |  |  |  |  | 7/2021
| Nickname | AH1 |  |  | EXO2 |  | Turb | 21G100845 |  | 599101-01 |  |  |  |  | pH | 4
| Preserve | CPAP |  |  |  |  | Cond | 23E103845 |  | 599827 |  |  |  |  | pH | 7 | 21M2S |  |  |  |  | 12/8/2023
| Station | CPAH2 |  |  |  |  | Chl | [7 empty] | pH | 10 | 22A1T |  |  |  |  | 1/3/2024
[7 empty] | Wiper | 21H105434 |  | 599090-01 |  |  |  |  | Turb (FNU) | 124 | 22D22130237 |  |  |  |  | 4/2023
|  |  |  | Wiper Replaced | Y |  |  |  | Batteries Replaced | N | [5 empty] | Technician(s) | Lena Kury
|  |  |  | Wiper Parked in Garage | Y |  |  |  | Date of Calibration | 2024-05-14 | [5 empty] | Date of Post-Calibration | 2024-06-28
| Pre-deployment Comments/ Failed-calibration values: | Back-up sonde; usual one in for maintenance
| Pre-Deployment (CAL & ICV) | [12 empty] | Post-Deployment (CCV) | [5 empty] | Programming
|  | Standards |  |  | Before Cal |  |  | Calibrated |  | Specs |  | P/ F | Time |  | Post value |  | P/ F | Time |  | Interval |  | 15
| Temp | 20.669 | °C (Therm) |  | 20.806 | °C (DL) |  |  |  | ±0.5(w)+0.1°C |  | P | 14:10 |  | 20.927 | °C (Therm) | P | 14:37 |  | Duration |  | 365 days
| SpCond in air |  |  |  | 0 |  |  |  |  | ≤ 0.002 mS/cm |  | P |  |  | 20.677 | °C (DL) |  |  |  | Free memory % |  | 100
| SpCond | 50 | mS/cm |  | 49.936 | mS/cm |  | 50 | mS/cm | ±5% |  |  | 13:51 | ✔ | 50.055 | mS/cm | P | 14:32 |  | Memory (days) |  | 24221
[6 empty] | Cell Constant |  | 0.47 |  | 0.419-0.519 (w) |  | P | [7 empty] | Start time
| SpCond 20 (ICV) | 10 | mS/cm |  |  | ✔ |  | 10.255 | mS/cm | ±5% |  | P | 13:53 | [6 empty] | Battery % |  | 97.5%
| pH 7 | 7 |  |  | 7.12 |  |  | 7.02 |  | 7 ±0.2 |  |  | 13:55 | ✔ | 7.2 |  | P | 14:54 |  | Set clock |  | Y
| pH 10 | 10 |  |  | 10.22 |  |  | 10.03 |  | 10 ±0.2 |  |  | 13:57 | ✔ | 9.99 |  | F | 14:56 |  | End time
[5 empty] | pH 7 |  |  | -23.7 | mV | 0 ±50 mV |  | P |  |  | -34.6 | mV | P |  |  |  | Other/notes
[5 empty] | pH 10 |  |  | -203.2 | mV | -180 ±50 mV |  | P |  |  | -200.8 | mV | P |  |  | BP = 29.82 = 757.43
[5 empty] | pH 4 |  |  |  | mV | +180 ±50 mV |  | P
[5 empty] | pH Slope |  |  | 179.5 |  | ≥ 165 mV |  | P |  |  | 166.2 |  | P
| pH 4 or 10 (ICV) | 10 |  |  |  | ✔ |  | 10.03 |  | 10 ±0.2 |  | P | 13:58
| BP at Depth Cal |  |  |  | 757.43 | mmHg | [8 empty] | 760.984 | mmHg
| Depth | -0.041 | Offset |  | -0.108 | m |  | -0.041 | m | ±5% |  | P | 14:02 |  | 0.01 | m
| Turbidity | 0 | FNU |  | 0 |  |  | 0 | FNU | 0.00 ±0.3 |  |  | 14:05 | ✔ | 0.18 | FNU | P | 15:01 |  |  | General Maintenance
| Turbidity | 124 | FNU |  | 123.62 | FNU |  | 124 | FNU | 124 ±6.2 |  |  | 14:08 | ✔ | 123.91 | FNU | P | 15:02
| Turbidity 124 (ICV) | 124 | FNU |  |  |  | ✔ | 124.1 | FNU | 124 ±6.2 |  | P | 14:08
| BP at ODO Cal |  |  |  | 757.43 | mmHg | [8 empty] | 760.984 | mmHg
| Optical DO % |  |  |  | 98.7 | % |  | 99.6 | % | NA |  | NA | 14:25 |  | 99.7 | % | P | 14:40
[6 empty] | DO gain |  | 1.04 | mg/L (table) | 0.87-1.25 |  | P |  |  | 99.7 | %
| Optical DO mg/L (ICV) | 8.9 | mg/L (DL) |  | 20.79 | °C |  | 8.915 |  | ±0.3 mg/L |  | P | 14:26 |  | 8.9 | mg/L (DL)
| Battery Voltage | [6 empty] | 5.91 | V | ~6V (EXO 2) |  | P |  |  | 20.683 | °C | P | 14:40
[6 empty] | Optional Parameters | [8 empty] | 5.59 | V | P
| Chlorophyll | 0 | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| Chlorophyll |  | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [5 empty] | °C
| Chlorophyll | 0 | µg/L |  |  | µg/L |  |  | µg/L | [6 empty] | µg/L |  |  |  |  | Post-Deployment Notes
| Chlorophyll |  | µg/L |  |  | µg/L |  |  | µg/L | [6 empty] | µg/L |  |  | BP = 29.96 = 760.984
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [5 empty] | °C
| BGA | 0 | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| BGA |  | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [5 empty] | °C
Field Log
Fouling Presence | [9 empty] | Deployment |  |  | Retrieval |  |  |  | Field Notes:
A=algae, AM=amphipods, B=barnacles, BR=bryozoans, C=crabs, E=eggs, F=fish, FA=filamentous algae, G=gastropod, H=hydroids, I=isopod, MD=mud, S=sponges, SI=silt, SL=shell, SP=shrimp, T=tunicates, W=worm, O=other | [8 empty] | Date deploy/retrieval | 2024-05-20 |  |  | 2024-06-28 |  |  | Deployment:
[9 empty] | Time deploy/retrieval | 9:44 |  |  | 10:46 |  |  | Turb: 6.49
[9 empty] | First/last sample time | 9:45 |  |  | 10:45
| H= heavy, M= moderate, L= light | [7 empty] | Technician initials | LK |  |  | LK
Mount: | [9 empty] | Discrete Samples (Required)
[9 empty] | Depth (m) | 2.4 | m |  | 2.2 |  | m
Sonde Body: A, T - L, oyster = 23mm | [8 empty] | Temp (°C) | 26.5 | °C |  | 29.9 |  | °C
[9 empty] | DO % | 96.3 | % |  | 101.9 |  | %
Field guard: T, B - H, some crabs | [8 empty] | DO mg/L | 6.54 | mg/L |  | 6.52 |  | mg/L
All Sensors: | [8 empty] | SpCond (mS/cm) | 46274 | mS/cm |  | 47815 |  | mS/cm | Retrieval:
C/T: T - H | [8 empty] | Salinity (psu) | 29.98 | psu |  | 31.01 |  | psu | Turb: 8.35
pH: T - H | [8 empty] | pH | 8 |  |  | 8
DO: T - H | [11 empty] | Optional Samples
Turb: T - H | [8 empty] | pHmV |  | mV |  |  |  | mV
Wiper: stayed on! | [8 empty] | Conductivity |  | mS/cm |  |  |  | mS/cm
TAL-PE:
Other/notes: nothing on probe faces, some tunicates on depth holes

## CPAH2_062824
Aquatic Preserve Continuous Water Quality Monitoring Calibration Log
|  | File Name | CPAH2_062824 | [5 empty] | .csv | [5 empty] | File Name | CPAH221E103416_062824_091500 |  |  |  |  | .bin
[7 empty] | Datasonde and Probe Identification Numbers | [12 empty] | Calibration Standards
|  | Sonde Code |  |  | Serial Number |  |  | Serial Number |  | Model Number | [5 empty] | Standard | Lot No. |  |  | Received | Opened | Exp. Date
| Datasonde | 1 |  |  | 21E103416 |  | pH | 21G105935 |  | 577602 |  |  |  |  | SpCond (mS/cm) | 20 | 4203H74 |  |  |  | 4/11/2022 | 3/2024
| Vented | N |  |  | Model |  | ODO | 21G100611 |  | 599100-01 |  |  |  |  | SpCond (mS/cm) | 50 | 2908A58 |  |  |  |  | 7/2021
| Nickname | AH1 |  |  | EXO2 |  | Turb | 21G100846 |  | 599101-01 |  |  |  |  | pH | 4
| Preserve | CPAP |  |  |  |  | Cond | 21G101303 |  | 599827 |  |  |  |  | pH | 7 | 21M2S |  |  |  |  | 12/8/2023
| Station | CPAH2 |  |  |  |  | Chl | [7 empty] | pH | 10 | 22A1T |  |  |  |  | 1/3/2024
[7 empty] | Wiper | 21H105437 |  | 599090-01 |  |  |  |  | Turb (FNU) | 124 | 22D22130237 |  |  |  |  | 4/2023
|  |  |  | Wiper Replaced | Y |  |  |  | Batteries Replaced | N | [5 empty] | Technician(s) | Lena Kury
|  |  |  | Wiper Parked in Garage | Y |  |  |  | Date of Calibration | 2024-06-27 | [5 empty] | Date of Post-Calibration
Pre-deployment Comments/ Failed-calibration values:
| Pre-Deployment (CAL & ICV) | [12 empty] | Post-Deployment (CCV) | [5 empty] | Programming
|  | Standards |  |  | Before Cal |  |  | Calibrated |  | Specs |  | P/ F | Time |  | Post value |  | P/ F | Time |  | Interval |  | 15
| Temp | 20.468 | °C (Therm) |  | 20.395 | °C (DL) |  |  |  | ±0.5(w)+0.1°C |  | P | 14:50 |  | 20.682 | °C (Therm) | P | 13:46 |  | Duration |  | 365 days
| SpCond in air |  |  |  | 0.001 |  |  |  |  | ≤ 0.002 mS/cm |  | P |  |  | 20.799 | °C (DL) |  |  |  | Free memory % |  | 100
| SpCond | 50 | mS/cm |  | 50.047 | mS/cm |  | 50 | mS/cm | ±5% |  |  | 13:59 | ✔ | 49.483 | mS/cm | P | 14:02 |  | Memory (days) |  | 24221
[6 empty] | Cell Constant |  | 0.47 |  | 0.419-0.519 (w) |  | P | [7 empty] | Start time |  | 9:15
| SpCond 20 (ICV) | 10 | mS/cm |  |  | ✔ |  | 10.078 | mS/cm | ±5% |  | P | 14:03 | [6 empty] | Battery % |  | 96.4%
| pH 7 | 7 |  |  | 7.22 |  |  | 7.01 |  | 7 ±0.2 |  |  | 14:06 | ✔ | 7.15 |  | P | 14:09 |  | Set clock |  | Y
| pH 10 | 10 |  |  | 10.17 |  |  | 10.02 |  | 10 ±0.2 |  |  | 14:08 | ✔ | 10.19 |  | P | 14:13 |  | End time
[5 empty] | pH 7 |  |  | -13 | mV | 0 ±50 mV |  | P |  |  | -20.4 | mV | P |  |  |  | Other/notes
[5 empty] | pH 10 |  |  | -185.9 | mV | -180 ±50 mV |  | P |  |  | -196 | mV | P |  |  | BP = 29.85 = 758.19
[5 empty] | pH 4 |  |  |  | mV | +180 ±50 mV |  | P
[5 empty] | pH Slope |  |  | 172.9 |  | ≥ 165 mV |  | P |  |  | 175.6 |  | P
| pH 4 or 10 (ICV) | 10 |  |  |  | ✔ |  | 10.01 |  | 10 ±0.2 |  | P | 14:09
| BP at Depth Cal |  |  |  | 758.19 | mmHg | [8 empty] | 765.81 | mmHg
| Depth | -0.027 | Offset |  | 0.301 | m |  | -0.027 | m | ±5% |  | P | 14:13 |  | 0.167 | m
| Turbidity | 0 | FNU |  | -0.39 |  |  | 0 | FNU | 0.00 ±0.3 |  |  | 14:17 | ✔ | 4.34 | FNU | F | 14:18 |  |  | General Maintenance
| Turbidity | 124 | FNU |  | 133.78 | FNU |  | 124 | FNU | 124 ±6.2 |  |  | 14:20 | ✔ | 103.48 | FNU | F | 14:21 | DO % readings taken from 6/30/24 01:00 right before it shut off
| Turbidity 124 (ICV) | 124 | FNU |  |  |  | ✔ | 124.04 | FNU | 124 ±6.2 |  | P | 14:21
| BP at ODO Cal |  |  |  | 758.19 | mmHg | [8 empty] | 765.81 | mmHg
| Optical DO % |  |  |  | 102.5 | % |  | 99.5 | % | NA |  | NA | 13:50 |  | 102.6 | % | P | 13:48
[6 empty] | DO gain |  | 1.06 | mg/L (table) | 0.87-1.25 |  | P |  |  | 102 | %
| Optical DO mg/L (ICV) | 8.98 | mg/L (DL) |  | 20.396 | °C |  | 9.09 |  | ±0.3 mg/L |  | P | 13:51 |  | 1.7 | mg/L (DL)
| Battery Voltage | [6 empty] | 5.88 | V | ~6V (EXO 2) |  | P |  |  | 20.798 | °C | P | 13:49
[6 empty] | Optional Parameters | [8 empty] | 5.86 | V | P
| Chlorophyll | 0 | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| Chlorophyll |  | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [5 empty] | °C
| Chlorophyll | 0 | µg/L |  |  | µg/L |  |  | µg/L | [6 empty] | µg/L |  |  |  |  | Post-Deployment Notes
| Chlorophyll |  | µg/L |  |  | µg/L |  |  | µg/L | [6 empty] | µg/L |  |  |  | BP = 30.15 = 765.81
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [5 empty] | °C |  |  |  | Sonde died ~2 days into deployment but voltage reads normal in the lab, maybe a sensor was not going to sleep when it was supposed to be
| BGA | 0 | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| BGA |  | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [5 empty] | °C
Field Log
Fouling Presence | [9 empty] | Deployment |  |  | Retrieval |  |  |  | Field Notes:
A=algae, AM=amphipods, B=barnacles, BR=bryozoans, C=crabs, E=eggs, F=fish, FA=filamentous algae, G=gastropod, H=hydroids, I=isopod, MD=mud, S=sponges, SI=silt, SL=shell, SP=shrimp, T=tunicates, W=worm, O=other | [8 empty] | Date deploy/retrieval | 2024-06-28 |  |  | 2024-07-30 |  |  | Deployment:
[9 empty] | Time deploy/retrieval | 10:48 |  |  | 12:17 |  |  | Turb: 8.35
[9 empty] | First/last sample time | 11:00 |  |  | 12:15
| H= heavy, M= moderate, L= light | [7 empty] | Technician initials | LK |  |  | LK
Mount: | [9 empty] | Discrete Samples (Required)
[9 empty] | Depth (m) | 2.2 | m |  | 2.3 |  | m
Sonde Body: MD - L | [8 empty] | Temp (°C) | 29.9 | °C |  | 31.1 |  | °C
[9 empty] | DO % | 101.9 | % |  | 104.8 |  | %
Field guard: T, A, MD, B - L | [8 empty] | DO mg/L | 6.52 | mg/L |  | 6.47 |  | mg/L
All Sensors: | [8 empty] | SpCond (mS/cm) | 47815 | mS/cm |  | 51154 |  | mS/cm | Retrieval:
C/T: T, MD - L | [8 empty] | Salinity (psu) | 31.01 | psu |  | 33.41 |  | psu | Turb: 3.45
pH: T, MD - L | [8 empty] | pH | 8 |  |  | 8.15
DO: T, MD - L | [11 empty] | Optional Samples
Turb: T, MD - L | [8 empty] | pHmV |  | mV |  |  |  | mV
Wiper: stayed on | [8 empty] | Conductivity |  | mS/cm |  |  |  | mS/cm
TAL-PE:
Other/notes: T on a depth probe hole

## CPAH2_073024
Aquatic Preserve Continuous Water Quality Monitoring Calibration Log
|  | File Name | CPAH2_073024 | [5 empty] | .csv | [5 empty] | File Name | CPAH221E103416_073024 |  |  |  | .bin
[7 empty] | Datasonde and Probe Identification Numbers | [11 empty] | Calibration Standards
|  | Sonde Code |  |  | Serial Number |  |  | Serial Number |  | Model Number | [5 empty] | Standard | Lot No. |  | Received | Opened | Exp. Date
| Datasonde | 2 |  |  | 21E103418 |  | pH | 21G105084 |  | 577602 |  |  |  |  | SpCond (mS/cm) | 20 | 4203H74 |  |  | 4/11/2022 | 3/2024
| Vented | N |  |  | Model |  | ODO | 21G100610 |  | 599100-01 |  |  |  |  | SpCond (mS/cm) | 50 | 2908A58 |  |  |  | 7/2021
| Nickname | AH2 |  |  | EXO2 |  | Turb | 21G100848 |  | 599101-01 |  |  |  |  | pH | 4
| Preserve | CPAP |  |  |  |  | Cond | 23F106588 |  | 599827 |  |  |  |  | pH | 7 | 21M2S |  |  |  | 12/8/2023
| Station | CPAH2 |  |  |  |  | Chl | [7 empty] | pH | 10 | 22A1T |  |  |  | 1/3/2024
[7 empty] | Wiper | 21H105435 |  | 599090-01 |  |  |  |  | Turb (FNU) | 124 | 22D22130237 |  |  |  | 4/2023
|  |  |  | Wiper Replaced | Y |  |  |  | Batteries Replaced | N | [5 empty] | Technician(s) | Lena Kury
|  |  |  | Wiper Parked in Garage | Y |  |  |  | Date of Calibration | 2024-07-30 | [5 empty] | Date of Post-Calibration
Pre-deployment Comments/ Failed-calibration values:
| Pre-Deployment (CAL & ICV) | [12 empty] | Post-Deployment (CCV) |  |  |  |  | Programming
|  | Standards |  |  | Before Cal |  |  | Calibrated |  | Specs |  | P/ F | Time |  | Post value |  | P/ F | Time | Interval |  | 15
| Temp |  | °C (Therm) |  |  | °C (DL) |  |  |  | ±0.5(w)+0.1°C |  | P |  |  |  | °C (Therm) | P |  | Duration |  | 365 days
| SpCond in air | [8 empty] | ≤ 0.002 mS/cm |  | P |  |  |  | °C (DL) |  |  | Free memory % |  | 100
| SpCond | 50 | mS/cm |  |  | mS/cm |  |  | mS/cm | ±5% |  |  |  | ✔ |  | mS/cm | P |  | Memory (days) |  | 24221
[6 empty] | Cell Constant |  |  |  | 0.419-0.519 (w) |  | P | [6 empty] | Start time
| SpCond 20 (ICV) | 20 | mS/cm |  |  | ✔ |  |  | mS/cm | ±5% |  | P | [6 empty] | Battery %
| pH 7 | 7 | [7 empty] | 7 ±0.2 |  |  |  | ✔ |  |  | P |  | Set clock |  | Y
| pH 10 | 10 | [7 empty] | 10 ±0.2 |  |  |  | ✔ |  |  | P |  | End time
[5 empty] | pH 7 |  |  |  | mV | 0 ±50 mV |  | P |  |  |  | mV | P |  |  | Other/notes
[5 empty] | pH 10 |  |  |  | mV | -180 ±50 mV |  | P |  |  |  | mV | P
[5 empty] | pH 4 |  |  |  | mV | +180 ±50 mV |  | P
[5 empty] | pH Slope |  |  | 0 |  | ≥ 165 mV |  | P |  |  | 0 |  | P
| pH 4 or 10 (ICV) | 10 |  |  |  | ✔ |  |  |  | 10 ±0.2 |  | P
| BP at Depth Cal |  |  |  |  | mmHg | [9 empty] | mmHg
| Depth |  | Offset |  |  | m |  |  | m | ±5% |  | P |  |  |  | m
| Turbidity | 0 | FNU | [5 empty] | FNU | 0.00 ±0.3 |  |  |  | ✔ |  | FNU | P |  |  | General Maintenance
| Turbidity | 124 | FNU |  |  | FNU |  |  | FNU | 124 ±6.2 |  |  |  | ✔ |  | FNU | P
| Turbidity 124 (ICV) | 124 | FNU |  |  |  | ✔ |  | FNU | 124 ±6.2 |  | P
| BP at ODO Cal |  |  |  |  | mmHg | [9 empty] | mmHg
| Optical DO % |  |  |  |  | % |  |  | % | NA |  | NA |  |  |  | % | P
[6 empty] | DO gain |  |  | mg/L (table) | 0.87-1.25 |  | P |  |  |  | %
| Optical DO mg/L (ICV) |  | mg/L (DL) |  |  | °C |  |  |  | ±0.3 mg/L |  | P |  |  |  | mg/L (DL)
| Battery Voltage | [7 empty] | V | ~6V (EXO 2) |  | P |  |  |  | °C | P
[6 empty] | Optional Parameters | [9 empty] | V | P
| Chlorophyll | 0 | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| Chlorophyll |  | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [5 empty] | °C
| Chlorophyll | 0 | µg/L |  |  | µg/L |  |  | µg/L | [6 empty] | µg/L |  |  |  | Post-Deployment Notes
| Chlorophyll |  | µg/L |  |  | µg/L |  |  | µg/L | [6 empty] | µg/L
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [5 empty] | °C
| BGA | 0 | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| BGA |  | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [5 empty] | °C
Field Log
Fouling Presence | [9 empty] | Deployment |  |  | Retrieval |  |  |  | Field Notes:
A=algae, AM=amphipods, B=barnacles, BR=bryozoans, C=crabs, E=eggs, F=fish, FA=filamentous algae, G=gastropod, H=hydroids, I=isopod, MD=mud, S=sponges, SI=silt, SL=shell, SP=shrimp, T=tunicates, W=worm, O=other | [8 empty] | Date deploy/retrieval | 2024-07-30 | [5 empty] | Deployment:
[9 empty] | Time deploy/retrieval | 12:20 | [5 empty] | Turb: 3.45
[9 empty] | First/last sample time | 12:30
| H= heavy, M= moderate, L= light | [7 empty] | Technician initials | LK |  |  | LK
Mount: | [9 empty] | Discrete Samples (Required)
[9 empty] | Depth (m) | 2.3 | m |  |  |  | m
Sonde Body: | [8 empty] | Temp (°C) | 31.1 | °C |  |  |  | °C
[9 empty] | DO % | 104.8 | % |  |  |  | %
Field guard: | [8 empty] | DO mg/L | 6.47 | mg/L |  |  |  | mg/L
All Sensors: | [8 empty] | SpCond (mS/cm) | 51154 | mS/cm |  |  |  | mS/cm | Retrieval:
C/T: | [8 empty] | Salinity (psu) | 33.41 | psu |  |  |  | psu | Turb:
pH: | [8 empty] | pH | 8.15
DO: | [11 empty] | Optional Samples
Turb: | [8 empty] | pHmV |  | mV |  |  |  | mV
Wiper: | [8 empty] | Conductivity |  | mS/cm |  |  |  | mS/cm
TAL-PE:
Other/notes: