## Instructions
Instructions
● | Do not move, edit, swap, or shift any columns, rows or cells in the calibration log.
● | In the Drop-downs tab, edit the white columns to match your station, probe, and sonde information. Gray columns should not be edited, only selected in the log drop downs.
● | Within the log, only in white boxes labeled notes/other, you may type in extra items you wish the calibration log to prompt (see example).
● | If a calibration or ICV fails, place a note with the failed value in the pre-deployment comments section. Passing recalibrated value should go in the actual log.
● | General maintenance section is for any maintenance (in-field maintenance, probe repair date/ reason, etc.).
● | Parameters should be calibrated in the order of the calibration log from top to bottom; this matters especially for temperature and SpCond.
● | Your NIST device should be at least as accurate as your sensor. The accuracy of the NIST device must be included in your temperature accuracy determination, meaning that the temperature probe specification should be added to the accuracy of your thermometer.
● | Submit calibration log in .xls format named by AP station code and year, for example, BBSCH_2021_CalLog
● | Tabs at the bottom should be named by the station code and deployment date, for example, BBSCH_010222. There should be approximately 12 tabs in each yearly calibration log file.
Other information
● | Optional parameters will not be uploaded to the AP Data Portal at this time but will remain in the dataset for possible future upload.
● | Check marks represent DEP SOP requirement for ICV (initial calibration verification) and CCV (continuing calibration verification) steps.
● | The DL next to a parameter's units indicates that the reading comes from the datalogger's probe.

## FT1500
Solubility of Oxygen in Water at Atmospheric Pressure
Temperature | Oxygen Solubility | Use this table to DEP verify that DO mg/L are reading within range based on the temperature reading.
°C | mg/L | DEP SOP FT 1500
0 | 14.621
1 | 14.216
2 | 13.829
3 | 13.46
4 | 13.107
5 | 12.77
6 | 12.447
7 | 12.139
8 | 11.843
9 | 11.559
10 | 11.288
11 | 11.027
12 | 10.777
13 | 10.537
14 | 10.306
15 | 10.084
16 | 9.87
17 | 9.665
18 | 9.467
19 | 9.276
20 | 9.092
21 | 8.915
22 | 8.743
23 | 8.578
24 | 8.418
25 | 8.263
26 | 8.113
27 | 7.968
28 | 7.827
29 | 7.691
30 | 7.559
31 | 7.43
32 | 7.305
33 | 7.183
34 | 7.065
35 | 9.65
36 | 6.837
37 | 6.727
38 | 6.62
39 | 6.515
40 | 6.412
41 | 6.312
42 | 6.213
43 | 6.116
44 | 6.021
45 | 5.927
46 | 8.835
47 | 5.744
48 | 5.654
49 | 5.565
50 | 5.477

## Drop-downs
AP | Station | Datasonde | Sonde No. | pH Serial No. | ODO Serial No. | Turb Serial No. | Cond Serial No. | Chl Serial No. | Wiper Serial No. | Y/N | Pass/Fail | pH Model | ODO Model | Turb Model | Cond Model | Chl Model | Wiper Model | Sonde Model | Lot No. | Technician | Cell Constant |  | Battery Voltage | DO Gain |  | Sp Cond (mS/cm) |  | Temperature |  | DO |  | Depth |  | Turbidity
CPAP | CPAH2 | 1 | 21E103418 | 21G105084 | 21G100610 | 21G100848 | 21G101301 |  | 21H105435 | Y | P | 577602 | 599100-01 | 599101-01 | 599827 | 599103-01 | 599090-01 | EXO2 | 2712749 | Jon Brucker | EXO Wiped CT | 0.419-0.519 (w) | ~6V (EXO 2) | EXO Optical | 0.87-1.25 | Non-wiped (10 std.) | 10 ± 0.05 (n) | Non-wiped (-5 to 35°C) | ±0.01(n) +0.1°C | Optical (0-200% air sat.) | ± 1% (op) | EXO (10m) | ±0.004 m | 0.00 ±0.3
| CPFS | 2 | 21E103415 | 21G105934 | 21G100609 | 21G100849 | 21G101300 |  | 21H105436 | N | F | 6151 | 6150 | 6136 | 6560 |  |  | EXO3 | 2906F57 | Michael Palandri | EXO Non-Wiped CT | 5.05-5.95 (n) | ~3V (EXO 3) | 6600 | 0.7 - 1.40 | Non-wiped (20 std.) | 20 ± 0.01 (n) | Wiped sensor | ±0.2(w) +0.1°C |  |  | 6600 (10m) | ±0.06 ft (0.018 m) | 124 ±2.48
| CPWD |  | 21E103417 | 21G105935 | 21G100608 | 21G100845 | 21G101303 |  | 21H105434 | Auto | [7 empty] | 6600 | 106779 | [6 empty] | Non-wiped (50 std.) | 50 ± 0.25 (n)
|  |  | 21E103419 | 21G105083 | 21G100607 | 21G100847 | 21G101304 |  | 21H105438 | [9 empty] | 21H2S | [6 empty] | Wiped Sensor (10 std.) | 10 ± 0.1 (w)
|  |  | 21E103416 | 21G105082 | 21G100611 | 21G100846 | 16G100119 |  | 21H105437 | [9 empty] | 21G3T | [6 empty] | Wiped Sensor (20 std.) | 20 ± 0.2 (w)
|  | 140387 | 10B100491 | 19F | 13C100355 | 15B104385 | 13C100687 | [11 empty] | 21G21280150 | [6 empty] | Wiped Sensor (50 std.) | 50 ± 0.5 (w)
|  | 145459 | 13C100687 | 19F | 15L103284 | 12C100448 | 15G100312
|  | 120803 | 02A0876AI | 20G | 15L103286 | 12C100447 | 16G100618
|  | 145154 | 12H102535 | 20C | 12H101596 | 14C103217
Note on AA1: Howe, Samantha: 0 to 100: ± 0.5% of reading or 0.001 mS/cm 100 to 200: ± 1% of reading Wiped Sensor: ±1% of the reading or 0.002 mS/cm (whichever is greater)
Note on AD2: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: +0.1 represents the accuracy of your thermometer. The total accuracy should include the addition of your NIST device. Your NIST should be at least this accurate.

## Example
Aquatic Preserve Continuous Water Quality Monitoring Calibration Log
|  | File Name | BBSCH_110221 | [5 empty] | .csv | [5 empty] | File Name | BBSCH19E101881_110221_071500 |  |  |  |  | .bin
[7 empty] | Datasonde and Probe Identification Numbers | [12 empty] | Calibration Standards
|  | Sonde Code |  |  | Serial Number |  |  | Serial Number |  | Model Number | [5 empty] | Standard | Lot No. |  |  | Received | Opened | Exp. Date
| Datasonde | 156019 |  |  | 19E101880 |  | pH | 19K101465 |  | 577602 |  |  |  |  | SpCond (mS/cm) | 20 | 2008A95 |  |  | 2021-07-12 | 08/23/2021 | 07/2022
| Vented | Y/N |  |  | Model |  | ODO | 20A103173 |  | 599100-01 |  |  |  |  | SpCond (mS/cm) | 50 | 2105829 |  |  | 2021-07-01 | 09/16/2021 | 04/2023
| Nickname | CH |  |  | EXO2 |  | Turb | 20A103291 |  | 599101-01 |  |  |  |  | pH | 4 | 2004922 |  |  | 2021-07-12 | 08/27/2021 | 03/2022
| Preserve | BBSAP |  |  |  |  | Cond | 20A103494 |  | 599827 |  |  |  |  | pH | 7 | 2103C60 |  |  | 2021-05-03 | 08/04/2021 | 02/2023
| Station | BBSCH (Chassahowitzka) |  |  |  |  | Chl | 19J104091 |  | 599103-01 |  |  |  |  | pH | 10 | 4102C02 |  |  | 2021-04-30 | 10/06/2021 | 07/2022
[7 empty] | Wiper | 19J105081 |  | 599090-01 |  |  |  |  | Turb (FNU) | 124 | 21H21280150 |  |  | 2021-08-24 | 08/27/2021 | 08/2022
|  |  |  | Wiper Replaced | N |  |  |  | Batteries Replaced | N | [5 empty] | Technician(s) | Trisha Green
|  |  |  | Wiper Parked in Garage | Y |  |  |  | Date of Calibration | 2021-11-02 | [5 empty] | Date of Post-Calibration | 2021-11-23
| Pre-deployment Comments/ Failed-calibration values: | If a calibration QC score or ICV fails, place a note with the failed value in the pre-deployment comments section. Passing recalibrated value should go in the actual log.
| Pre-Deployment (CAL & ICV) | [12 empty] | Post-Deployment (CCV) | [5 empty] | Programming
|  | Standards |  |  | Before Cal |  |  | Calibrated |  | Specs |  | P/ F | Time |  | Post value |  | P/ F | Time |  | Interval |  | 15
| Temp | 21.75 | °C (Therm) |  | 21.693 | °C (DL) |  |  |  | ±0.2(w) +0.1°C |  | P | 7:10 |  | 20.5 | °C (Therm) | P | 12:56 |  | Duration |  | 365 days
| SpCond in air |  |  |  | 0 |  |  |  |  | ≤ 0.002 mS/cm |  | P |  |  | 20.46 | °C (DL) |  |  |  | Free memory % |  | 99.55
| SpCond | 50 | mS/cm |  | 49.767 | mS/cm |  | 50 | mS/cm | 50 ± 0.25 (n) |  |  | 7:08 | ✔ | 50.02 | mS/cm | P | 13:00 |  | Memory (days) |  | 24432.81
[6 empty] | Cell Constant |  | 0.47 |  | 0.419-0.519 (w) |  | P | [7 empty] | Start time |  | 9:30
| SpCond 20 (ICV) | 20 | mS/cm |  |  | ✔ |  | 20.112 | mS/cm | 20 ± 0.2 (w) |  | P | 7:14 | [6 empty] | Battery % |  | 65
| pH 7 | 7 |  |  | 7.05 |  |  | 7.02 |  | ± 0.2 |  |  | 7:11 | ✔ | 7.1 |  | P | 13:04 |  | Set clock |  | Y/N
| pH 10 | 10 |  |  | 10.08 |  |  | 10.05 |  | ± 0.2 |  |  | 7:11 | ✔ | 10.06 |  | P | 13:05 |  | End time |  | 10:00
[5 empty] | pH 7 |  |  | -30.7 | mV | 0 ±50 mV |  | P |  |  | -35.5 | mV | P |  |  |  | Other/notes
[5 empty] | pH 10 |  |  | -204.1 | mV | -180 ±50 mV |  | P |  |  | -203.9 | mV | P |  |  | Change locks |  | Y
[5 empty] | pH 4 |  |  | 148.8 | mV | +180 ±50 mV |  | P | [7 empty] | Check for DST |  | Y
[5 empty] | pH Slope |  |  | 173.4 |  | > 155 mV |  | P |  |  | 168.4 |  | P
| pH 4 or 10 (ICV) | 4 |  |  |  | ✔ |  | 3.9 |  | ± 0.1 |  | P | 7:16
| BP at Depth Cal |  |  |  | 765.556 | mmHg | [8 empty] | 765.56 | mmHg
| Depth | 0.075534 | Offset |  | 0.062 | m |  | 0.076 | m | ±0.004 m |  | P | 7:24 |  | 0.072 | m | P | 13:10
| Turbidity | 0 | FNU |  | -0.1 |  |  | 0 | FNU | 0.00 ±0.3 |  |  | 7:27 | ✔ | 0.04 | FNU | P | 13:15 |  |  | General Maintenance
| Turbidity | 124 | FNU |  | 123.24 | FNU |  | 124 | FNU | 124 ±2.48 |  |  | 7:31 | ✔ | 117.72 | FNU | P | 13:17 | This section is for any maintenance (in-field maintenance, probe repair etc.). Please include dates, probe numbers, and any information that is needed to track the incident.
| Turbidity 124 (ICV) | 124 | FNU |  |  |  | ✔ | 123.89 | FNU | 124 ±2.48 |  | P | 7:36
| BP at ODO Cal |  |  |  | 765.556 | mmHg | [8 empty] | 765.556 | mmHg
| Optical DO % |  |  |  | 101.7 | % |  | 100.9 | % | ± 1% |  | P | 7:38 |  | 100.5 | % | P | 13:22
[6 empty] | DO gain |  | 1.02 | mg/L (table) | 0.87-1.25 |  | P |  |  | 100.5 | %
| Optical DO mg/L (ICV) | 8.85 | mg/L (DL) |  | 21.75 | °C |  | 8.743 |  | ± 0.1 mg/L |  | P | 7:43 |  | 9.05 | mg/L (DL)
| Battery Voltage | [6 empty] | 3.2 | V | ~3V (EXO 3) |  | P |  |  | 20.5 | °C | P | 13:29
[6 empty] | Optional Parameters | [8 empty] | 2.24 | V | P
| Chlorophyll | 0 | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| Chlorophyll |  | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [5 empty] | °C
| Chlorophyll | 0 | µg/L |  |  | µg/L |  |  | µg/L | [6 empty] | µg/L |  |  |  |  | Post-Deployment Notes
| Chlorophyll |  | µg/L |  |  | µg/L |  |  | µg/L | [6 empty] | µg/L |  |  | Turbidity post-cal kind of low. Watch this sensor. May need to be replaced soon or just needed to be cleaned well.
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [5 empty] | °C
| BGA | 0 | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| BGA |  | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [5 empty] | °C
Field Log
Fouling Presence | [9 empty] | Deployment |  |  | Retrieval |  |  |  | Field Notes:
A=algae, AM=amphipods, B=barnacles, BR=bryozoans, C=crabs, E=eggs, F=fish, FA=filamentous algae, G=gastropod, H=hydroids, I=isopod, MD=mud, S=sponges, SI=silt, SL=shell, SP=shrimp, T=tunicates, W=worm, O=other | [8 empty] | Date deploy/retrieval | 2021-11-02 |  |  | 2021-11-23 |  |  | Deployment:
[9 empty] | Time deploy/retrieval | 8:52 |  |  | 10:17 |  |  | Low tide, calm, light winds.
[9 empty] | First/last sample time | 9:00 |  |  | 10:00
| H= heavy, M= moderate, L= light | [7 empty] | Technician initials | TG, KS |  |  | TG, KS, TJ
Mount: 2 Fish; A/H, B/M, Si/M | [9 empty] | Discrete Samples (Required)
[9 empty] | Depth (m) | 0.0634 | m |  | 0.0787 |  | m
Sonde Body: Fish in guard; E/H, Si/M, T/L, W/L | [8 empty] | Temp (°C) | 19.3 | °C |  | 17.5 |  | °C
[9 empty] | DO % | 103.9 | % |  | 94.2 |  | %
Field guard: | [8 empty] | DO mg/L | 8.21 | mg/L |  | 7.94 |  | mg/L
All Sesnors: Si/M, EH, W/L | [8 empty] | SpCond (mS/cm) | 41.065 | mS/cm |  | 33.298 |  | mS/cm | Retrieval:
C/T: G/L | [8 empty] | Salinity (psu) | 26.35 | psu |  | 20.91 |  | psu | Low tide, cold, windy, choppy, brown water, clear skies.
pH: | [8 empty] | pH | 8.2 |  |  | 8.03
DO: | [11 empty] | Optional Samples
Turb: | [8 empty] | pHmV | -76.9 | mV |  | -74.1 |  | mV
Wiper: Bad; fell off | [8 empty] | Conductivity |  | mS/cm |  |  |  | mS/cm
TAL-PE:
Other/Notes:

## APCODE_MMDDYY
Aquatic Preserve Continuous Water Quality Monitoring Calibration Log
|  | File Name | [6 empty] | .csv | [5 empty] | File Name |  |  |  |  | .bin
[7 empty] | Datasonde and Probe Identification Numbers | [11 empty] | Calibration Standards
|  | Sonde Code |  |  | Serial Number |  |  | Serial Number |  | Model Number | [5 empty] | Standard | Lot No. |  | Received | Opened | Exp. Date
| Datasonde | Datasonde |  |  | Sonde No. |  | pH | pH Serial No. |  | pH Model |  |  |  |  | SpCond (mS/cm) | 20
| Vented | Y/N |  |  | Model |  | ODO | ODO Serial No. |  | ODO Model |  |  |  |  | SpCond (mS/cm) | 50
| Nickname |  |  |  | Sonde Model |  | Turb | Turb Serial No. |  | Turb Model |  |  |  |  | pH | 4
| Preserve | AP |  |  |  |  | Cond | Cond Serial No. |  | Cond Model |  |  |  |  | pH | 7
| Station | Station |  |  |  |  | Chl | Chl Serial No. |  | Chl Model |  |  |  |  | pH | 10
[7 empty] | Wiper | Wiper Serial No. |  | Wiper Model |  |  |  |  | Turb (FNU) | 124
|  |  |  | Wiper Replaced |  |  |  |  | Batteries Replaced | [6 empty] | Technician(s) | Technician
|  |  |  | Wiper Parked in Garage |  |  |  |  | Date of Calibration | [6 empty] | Date of Post-Calibration
Pre-deployment Comments/ Failed-calibration values:
| Pre-Deployment (CAL & ICV) | [12 empty] | Post-Deployment (CCV) |  |  |  |  | Programming
|  | Standards |  |  | Before Cal |  |  | Calibrated |  | Specs |  | P/ F | Time |  | Post value |  | P/ F | Time | Interval |  | 15
| Temp |  | °C (Therm) |  |  | °C (DL) |  |  |  | ±0.2(w) +0.1°C | [5 empty] | °C (Therm) |  |  | Duration |  | 365 days
| SpCond in air | [8 empty] | ≤ 0.002 mS/cm | [5 empty] | °C (DL) |  |  | Free memory %
| SpCond | 50 | mS/cm |  |  | mS/cm |  |  | mS/cm | 50 ± 0.25 (n) |  |  |  | ✔ |  | mS/cm |  |  | Memory (days)
[6 empty] | Cell Constant |  |  |  | 0.419-0.519 (w) | [8 empty] | Start time
| SpCond 20 (ICV) | 20 | mS/cm |  |  | ✔ |  |  | mS/cm | 20 ± 0.2 (w) | [8 empty] | Battery %
| pH 7 | 7 | [7 empty] | ± 0.1 |  |  |  | ✔ |  |  |  |  | Set clock |  | Y/N
| pH 10 | 10 | [7 empty] | ± 0.1 |  |  |  | ✔ |  |  |  |  | End time
[5 empty] | pH 7 |  |  |  | mV | 0 ±50 mV | [5 empty] | mV |  |  |  | Other/notes
[5 empty] | pH 10 |  |  |  | mV | -180 ±50 mV | [5 empty] | mV
[5 empty] | pH 4 |  |  |  | mV | +180 ±50 mV
[5 empty] | pH Slope |  |  | 0 |  | > 155 mV |  |  |  |  | 0
| pH 4 or 10 (ICV) | 10 |  |  |  | ✔ |  |  |  | ± 0.1
| BP at Depth Cal |  |  |  |  | mmHg | [9 empty] | mmHg
| Depth |  | Offset |  |  | m |  |  | m | ±0.004 m | [5 empty] | m
| Turbidity | 0 | FNU | [5 empty] | FNU | 0.00 ±0.3 |  |  |  | ✔ |  | FNU |  |  |  | General Maintenance
| Turbidity | 124 | FNU |  |  | FNU |  |  | FNU | 124 ±2.48 |  |  |  | ✔ |  | FNU
| Turbidity 124 (ICV) | 124 | FNU |  |  |  | ✔ |  | FNU | 124 ±2.48
| BP at ODO Cal |  |  |  |  | mmHg | [9 empty] | mmHg
| Optical DO % |  |  |  |  | % |  |  | % | ± 1% | [5 empty] | %
[6 empty] | DO gain |  |  | mg/L (table) | 0.87-1.25 | [5 empty] | %
| Optical DO mg/L (ICV) |  | mg/L (DL) |  |  | °C |  |  |  | ± 0.1 mg/L | [5 empty] | mg/L (DL)
| Battery Voltage | [7 empty] | V | ~6V (EXO 2) | [5 empty] | °C
[6 empty] | Optional Parameters | [9 empty] | V
| Chlorophyll | 0 | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| Chlorophyll |  | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [5 empty] | °C
| Chlorophyll | 0 | µg/L |  |  | µg/L |  |  | µg/L | [6 empty] | µg/L |  |  |  | Post-Deployment Notes
| Chlorophyll |  | µg/L |  |  | µg/L |  |  | µg/L | [6 empty] | µg/L
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [5 empty] | °C
| BGA | 0 | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| BGA |  | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [5 empty] | °C
Field Log
Fouling Presence | [9 empty] | Deployment |  |  | Retrieval |  |  |  | Field Notes:
A=algae, AM=amphipods, B=barnacles, BR=bryozoans, C=crabs, E=eggs, F=fish, FA=filamentous algae, G=gastropod, H=hydroids, I=isopod, MD=mud, S=sponges, SI=silt, SL=shell, SP=shrimp, T=tunicates, W=worm, O=other | [8 empty] | Date deploy/retrieval | [6 empty] | Deployment:
Time deploy/retrieval
First/last sample time
| H= heavy, M= moderate, L= light | [7 empty] | Technician initials
Mount: | [9 empty] | Discrete Samples (Required)
[9 empty] | Depth (m) |  | m |  |  |  | m
Sonde Body: | [8 empty] | Temp (°C) |  | °C |  |  |  | °C
[9 empty] | DO % |  | % |  |  |  | %
Field guard: | [8 empty] | DO mg/L |  | mg/L |  |  |  | mg/L
All Sesnors: | [8 empty] | SpCond (mS/cm) |  | mS/cm |  |  |  | mS/cm | Retrieval:
C/T: | [8 empty] | Salinity (psu) |  | psu |  |  |  | psu
pH: | [8 empty] | pH
DO: | [11 empty] | Optional Samples
Turb: | [8 empty] | pHmV |  | mV |  |  |  | mV
Wiper: | [8 empty] | Conductivity |  | mS/cm |  |  |  | mS/cm
TAL-PE:
Other/notes:

## CPAH2_011422
Aquatic Preserve Continuous Water Quality Monitoring Calibration Log
|  | File Name | CPAH2_011422 | [5 empty] | .csv | [5 empty] | File Name | CPAH2_011422.DAT |  |  |  | .bin
[7 empty] | Datasonde and Probe Identification Numbers | [11 empty] | Calibration Standards
|  | Sonde Code |  |  | Serial Number |  |  | Serial Number |  | Model Number | [5 empty] | Standard | Lot No. |  | Received | Opened | Exp. Date
| Datasonde | 140387 |  |  | 10B100491 |  | pH | 19F |  | 6151 |  |  |  |  | SpCond (mS/cm) | 20 | - |  | - | - | -
| Vented | N |  |  | Model |  | ODO | 13C100355 |  | 6150 |  |  |  |  | SpCond (mS/cm) | 50 | 2906F57 |  | - | - | 7/2021
| Nickname | AH |  |  | 6600 |  | Turb | 15B104385 |  | 6136 |  |  |  |  | pH | 4 | 160779 |  | - | - | 7/13/23
| Preserve | CPAP |  |  |  |  | Cond | 16G100119 |  | 6560 |  |  |  |  | pH | 7 | 21H2S |  | - | - | 8/11/23
| Station | CPAH2 |  |  |  |  | Chl | Chl Serial No. |  | Chl Model |  |  |  |  | pH | 10 | 21G3T |  | - | - | 7/13/23
[7 empty] | Wiper | Wiper Serial No. |  | Wiper Model |  |  |  |  | Turb (FNU) | 124 | 21G21280150 |  |  | 11/1/21 | 7/2022
|  |  |  | Wiper Replaced | N |  |  |  | Batteries Replaced | Y | [5 empty] | Technician(s) | Jon Brucker
|  |  |  | Wiper Parked in Garage | Y |  |  |  | Date of Calibration | 2022-01-13 | [5 empty] | Date of Post-Calibration | 2022-02-18
| Pre-deployment Comments/ Failed-calibration values: | New batteries installed 12/28 and at retrieval 2/18. Cannot complete IVC for sp cond until get new standard. Very old and now on backorder with Fisher Sci.
| Pre-Deployment (CAL & ICV) | [12 empty] | Post-Deployment (CCV) |  |  |  |  | Programming
|  | Standards |  |  | Before Cal |  |  | Calibrated |  | Specs |  | P/ F | Time |  | Post value |  | P/ F | Time | Interval |  | 15
| Temp | 20.4 | °C (Therm) |  | 20.2 | °C (DL) |  |  |  | ±0.2(w) +0.1°C |  | P |  |  | 21.2 | °C (Therm) | P |  | Duration |  | 365 days
| SpCond in air |  |  |  | 0 |  |  |  |  | ≤ 0.002 mS/cm |  | P |  |  | 21 | °C (DL) |  |  | Free memory %
| SpCond | 50 | mS/cm |  | 49.32 | mS/cm |  | 50 | mS/cm | 50 ± 0.25 (n) |  |  |  | ✔ | 47.39 | mS/cm | P |  | Memory (days) |  | 58.8
[6 empty] | Cell Constant |  | 5.1573 |  | 0.419-0.519 (w) |  | P | [6 empty] | Start time |  | 9:59
| SpCond 20 (ICV) | 20 | mS/cm |  |  | ✔ |  | N/A | mS/cm | 20 ± 0.2 (w) | [8 empty] | Battery % |  | 62.8 days
| pH 7 | 7 |  |  | 7.1 |  |  | 7 |  | ± 0.2 |  |  |  | ✔ | 7.07 |  | P |  | Set clock |  | Y
| pH 10 | 10 |  |  | 9.75 |  |  | 10 |  | ± 0.2 |  |  |  | ✔ |  |  | P |  | End time |  | 13:45
[5 empty] | pH 7 |  |  | -45.4 | mV | 0 ±50 mV |  | P |  |  | -57.5 | mV | P |  |  | Other/notes
[5 empty] | pH 10 |  |  | -209.1 | mV | -180 ±50 mV |  | P |  |  |  | mV | P
[5 empty] | pH 4 |  |  |  | mV | +180 ±50 mV |  | P
[5 empty] | pH Slope |  |  | 163.7 |  | > 155 mV |  | P |  |  | -57.5 |  | P
| pH 4 or 10 (ICV) | 10 |  |  |  | ✔ |  | 10 |  | ± 0.1 |  | P
| BP at Depth Cal |  |  |  | 764.54 | mmHg | [8 empty] | 767.842 | mmHg
| Depth | 0.062 | Offset |  | -0.004 | m |  | 0.062 | m | ±0.004 m |  | P |  |  | 0.134 | m | P
| Turbidity | 0 | FNU |  | 0.6 |  |  | 0 | FNU | 0.00 ±0.3 |  |  |  | ✔ | -0.5 | FNU | P |  |  | General Maintenance
| Turbidity | 126 | FNU |  | 129.3 | FNU |  | 126 | FNU | 124 ±2.48 |  |  |  | ✔ |  | FNU
| Turbidity 124 (ICV) | 126 | FNU |  |  |  | ✔ | 126 | FNU | 124 ±2.48 |  | P
| BP at ODO Cal |  |  |  | 764.54 | mmHg | [8 empty] | 767.842 | mmHg
| Optical DO % |  |  |  | 97.6 | % |  | 100.6 | % | ± 1% |  | P |  |  | 103.1 | % | P
[6 empty] | DO gain |  | 1.0418 | mg/L (table) | 0.87-1.25 |  | P |  |  | 103.5 | %
| Optical DO mg/L (ICV) | N/A | mg/L (DL) |  | N/A | °C |  | N/A |  | ± 0.1 mg/L |  |  |  |  | N/A | mg/L (DL)
| Battery Voltage | [6 empty] | 12.8 | V | ~6V (EXO 2) |  | P |  |  |  | °C
[6 empty] | Optional Parameters | [9 empty] | V
| Chlorophyll | 0 | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| Chlorophyll |  | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [5 empty] | °C
| Chlorophyll | 0 | µg/L |  |  | µg/L |  |  | µg/L | [6 empty] | µg/L |  |  |  | Post-Deployment Notes
| Chlorophyll |  | µg/L |  |  | µg/L |  |  | µg/L | [6 empty] | µg/L
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [5 empty] | °C
| BGA | 0 | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| BGA |  | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [5 empty] | °C
Field Log
Fouling Presence | [9 empty] | Deployment |  |  | Retrieval |  |  |  | Field Notes:
A=algae, AM=amphipods, B=barnacles, BR=bryozoans, C=crabs, E=eggs, F=fish, FA=filamentous algae, G=gastropod, H=hydroids, I=isopod, MD=mud, S=sponges, SI=silt, SL=shell, SP=shrimp, T=tunicates, W=worm, O=other | [8 empty] | Date deploy/retrieval | 2022-01-14 |  |  | 2022-02-18 |  |  | Deployment: Calm, sunny, low tide (flood). Staff did not record in field discrete data
[9 empty] | Time deploy/retrieval | 11:44 |  |  | 13:35
[9 empty] | First/last sample time | 11:45 |  |  | 13:30
| H= heavy, M= moderate, L= light | [7 empty] | Technician initials | JB |  |  | JB
None | [9 empty] | Discrete Samples (Required)
[9 empty] | Depth (m) |  | m |  |  |  | m
Sonde Body: None | [8 empty] | Temp (°C) |  | °C |  |  |  | °C
[9 empty] | DO % |  | % |  |  |  | %
Field guard: | [8 empty] | DO mg/L |  | mg/L |  |  |  | mg/L
All Sesnors: | [8 empty] | SpCond (mS/cm) |  | mS/cm |  |  |  | mS/cm | Retrieval:
C/T: N | [8 empty] | Salinity (psu) |  | psu |  |  |  | psu | Staff did not record in field discrete samples
pH:N | [8 empty] | pH
DO: N | [11 empty] | Optional Samples
Turb: N | [8 empty] | pHmV |  | mV |  |  |  | mV
Wiper: N | [8 empty] | Conductivity |  | mS/cm |  |  |  | mS/cm
TAL-PE:
Other/notes:

## CPAH2_021822
Aquatic Preserve Continuous Water Quality Monitoring Calibration Log
|  | File Name | CPAH2_021822 | [5 empty] | .csv | [6 empty] | File Name | CPAH2_21E103416_021822_104500 |  |  |  |  | .bin
[7 empty] | Datasonde and Probe Identification Numbers | [13 empty] | Calibration Standards
|  | Sonde Code |  |  | Serial Number |  |  | Serial Number |  | Model Number | [6 empty] | Standard | Lot No. |  |  | Received | Opened | Exp. Date
| Datasonde | 1 |  |  | 21E103416 |  | pH | 21G105082 |  | 577602 | [5 empty] | SpCond (mS/cm) | 20
| Vented | N |  |  | Model |  | ODO | 21G100611 |  | 599100-01 | [5 empty] | SpCond (mS/cm) | 50 | 2906F57 |  |  |  |  | 7/2021
| Nickname | AH1 |  |  | EXO2 |  | Turb | 21G100846 |  | 599101-01 | [5 empty] | pH | 4 | 160779 |  |  |  |  | 7/13/23
| Preserve | CPAP |  |  |  |  | Cond | 21G101304 |  | 599827 | [5 empty] | pH | 7 | 21H2S |  |  |  |  | 8/11/23
| Station | CPAH2 |  |  |  |  | Chl | Chl Serial No. |  | Chl Model | [5 empty] | pH | 10 | 21G3T |  |  |  |  | 7/13/23
[7 empty] | Wiper | 21H105497 |  | 599090-01 | [5 empty] | Turb (FNU) | 124 | 21G21280150 |  |  |  | 11/1/21 | 7/2022
|  |  |  | Wiper Replaced | Y |  |  |  | Batteries Replaced | Y | [6 empty] | Technician(s) | Jon Brucker
|  |  |  | Wiper Parked in Garage | Y |  |  |  | Date of Calibration | 2022-02-08 | [6 empty] | Date of Post-Calibration | 2022-03-18
| Pre-deployment Comments/ Failed-calibration values: | First deployment with sonde and probes. New batteries, probes, and wiper. Do not have any 10 or 20 conductivty standard. Product on backorder with Fisher Sci.
| Pre-Deployment (CAL & ICV) | [13 empty] | Post-Deployment (CCV) | [5 empty] | Programming
|  | Standards |  |  | Before Cal |  |  | Calibrated |  | Specs |  | P/ F |  | Time |  | Post value |  | P/ F | Time |  | Interval |  | 15
| Temp | 16.6 | °C (Therm) |  | 16.67 | °C (DL) |  |  |  | ±0.2(w) +0.1°C |  | P |  | 9:52 |  | 19.6 | °C (Therm) | P | 8:34 |  | Duration |  | 365 days
| SpCond in air |  |  |  | 0 |  |  |  |  | ≤ 0.002 mS/cm |  | P |  |  |  | 19.442 | °C (DL) |  |  |  | Free memory % |  | 100
| SpCond | 50 | mS/cm |  | 49.679 | mS/cm |  | 50 | mS/cm | 50 ± 0.25 (n) |  |  |  | 9:53 | ✔ | 50.237 | mS/cm | P | 8:51 |  | Memory (days) |  | 24542
[6 empty] | Cell Constant |  | 0.47202 |  | 0.419-0.519 (w) |  | P | [8 empty] | Start time |  | 10:45
| SpCond 20 (ICV) | 20 | mS/cm |  |  | ✔ |  | N/A | mS/cm | 20 ± 0.2 (w) |  |  |  | N/A | [6 empty] | Battery % |  | 100
| pH 7 | 7 |  |  | 7.29 |  |  | 7 |  | ± 0.2 |  |  |  | 10:04 | ✔ | 7.24 |  | F | 8:54 |  | Set clock |  | Y
| pH 10 | 10 |  |  | 10.46 |  |  | 10 |  | ± 0.2 |  |  |  | 10:06 | ✔ | 10 |  | P | 9:00 |  | End time |  | 12:00
[5 empty] | pH 7 |  |  | -16.8 | mV | 0 ±50 mV |  | P |  |  |  | -31.5 | mV | P |  |  |  | Other/notes
[5 empty] | pH 10 |  |  | -198.4 | mV | -180 ±50 mV |  | P |  |  |  | -195.3 | mV | P
[5 empty] | pH 4 |  |  |  | mV | +180 ±50 mV |  | P
[5 empty] | pH Slope |  |  | 181.6 |  | > 155 mV |  | P |  |  |  | 163.8 |  | P
| pH 4 or 10 (ICV) | 10 |  |  |  | ✔ |  | 10.08 |  | ± 0.1 |  | P |  | 10:31
| BP at Depth Cal |  |  |  | 766.572 | mmHg | [9 empty] | 761.46 | mmHg
| Depth | 0.095 | Offset |  | 10.552 | m |  | 0.095 | m | ±0.004 m |  | P |  | 10:36 |  | 0.028 | m | P | 8:46
| Turbidity | 0 | FNU |  | 0.35 |  |  | 0 | FNU | 0.00 ±0.3 |  |  |  | 10:36 | ✔ | -0.14 | FNU | P | 9:04 |  |  | General Maintenance
| Turbidity | 124 | FNU |  | 124.38 | FNU |  | 124 | FNU | 124 ±2.48 |  |  |  | 10:48 | ✔ | 120.7 | FNU | F | 9:08 | Staff did not record discrete in field samples.Changed shackle, cleaned sonde tube.
| Turbidity 124 (ICV) | 124 | FNU |  |  |  | ✔ | 123.83 | FNU | 124 ±2.48 |  | P |  | 10:48
| BP at ODO Cal |  |  |  | 762 | mmHg | [9 empty] | 761.46 | mmHg
| Optical DO % |  |  |  | 96 | % |  | 100.3 | % | ± 1% |  | P | 7:56 | 7:56 |  | 101.4 | % | P | 9:10
[6 empty] | DO gain |  | 1.04 | mg/L (table) | 0.87-1.25 |  | P |  |  |  | 101.3 | %
| Optical DO mg/L (ICV) | 9.14 | mg/L (DL) |  | 19.755 | °C |  | 9.092 |  | ± 0.1 mg/L |  | P |  | 7:58 |  | 9.04 | mg/L (DL)
| Battery Voltage | [7 empty] | V | ~6V (EXO 2) |  | P |  |  |  | 20.988 | °C | F | 9:10
[6 empty] | Optional Parameters | [9 empty] | 5.83 | V | P
| Chlorophyll | 0 | RFU |  |  | RFU |  |  | RFU | [7 empty] | RFU
| Chlorophyll |  | RFU |  |  | RFU |  |  | RFU | [7 empty] | RFU
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [6 empty] | °C
| Chlorophyll | 0 | µg/L |  |  | µg/L |  |  | µg/L | [7 empty] | µg/L |  |  |  |  | Post-Deployment Notes
| Chlorophyll |  | µg/L |  |  | µg/L |  |  | µg/L | [7 empty] | µg/L
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [6 empty] | °C
| BGA | 0 | RFU |  |  | RFU |  |  | RFU | [7 empty] | RFU
| BGA |  | RFU |  |  | RFU |  |  | RFU | [7 empty] | RFU
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [6 empty] | °C
Field Log
Fouling Presence | [9 empty] | Deployment |  |  |  | Retrieval |  |  |  | Field Notes:
A=algae, AM=amphipods, B=barnacles, BR=bryozoans, C=crabs, E=eggs, F=fish, FA=filamentous algae, G=gastropod, H=hydroids, I=isopod, MD=mud, S=sponges, SI=silt, SL=shell, SP=shrimp, T=tunicates, W=worm, O=other | [8 empty] | Date deploy/retrieval | 2022-02-18 |  |  |  | 2022-03-16 |  |  | Staff did not record discrete samples upon deployment.
[9 empty] | Time deploy/retrieval | 13:37 |  |  |  | 12:07
[9 empty] | First/last sample time | 13:45 |  |  |  | 12:00
| H= heavy, M= moderate, L= light | [7 empty] | Technician initials | JB |  |  |  | MP
Mount: N | [9 empty] | Discrete Samples (Required)
[9 empty] | Depth (m) | N/A | m |  |  | 2 |  | m
Sonde Body: N | [8 empty] | Temp (°C) | N/A | °C |  |  | 16.1 |  | °C
[9 empty] | DO % | N/A | % |  |  | 105.6 |  | %
Field guard: A/M, O/M (anemone), B/M, SL/M | [8 empty] | DO mg/L | N/A | mg/L |  |  | 8.55 |  | mg/L
All Sesnors: | [8 empty] | SpCond (mS/cm) | N/A | mS/cm |  |  | 49737 |  | mS/cm | Retrieval:
C/T: N | [8 empty] | Salinity (psu) | N/A | psu |  |  | 32.61 |  | psu | Calm, sunny, low tide rising.
pH: A/L | [8 empty] | pH | N/A |  |  |  | 8.17
DO: O/L (anemone), B/L | [11 empty] | Optional Samples
Turb: A/L | [8 empty] | pHmV | N/A | mV |  |  | N/A |  | mV
Wiper: B/L, A/L | [8 empty] | Conductivity | N/A | mS/cm |  |  | N/A |  | mS/cm
TAL-PE: N/A
Other/notes:

## CPAH2_031622
Aquatic Preserve Continuous Water Quality Monitoring Calibration Log
|  | File Name | CPAH2_031622 | [5 empty] | .csv | [5 empty] | File Name | CPAH2_21E103418_031622_074500 |  |  |  |  | .bin
[7 empty] | Datasonde and Probe Identification Numbers | [12 empty] | Calibration Standards
|  | Sonde Code |  |  | Serial Number |  |  | Serial Number |  | Model Number | [5 empty] | Standard | Lot No. |  |  | Received | Opened | Exp. Date
| Datasonde | 2 |  |  | 21E103418 |  | pH | 21G105084 |  | 577602 |  |  |  |  | SpCond (mS/cm) | 20
| Vented | N |  |  | Model |  | ODO | 21G100610 |  | 599100-01 |  |  |  |  | SpCond (mS/cm) | 50 | 2906F57 |  |  |  |  | 7/2021
| Nickname | AH2 |  |  | EXO2 |  | Turb | 21G100848 |  | 599101-01 |  |  |  |  | pH | 4
| Preserve | CPAP |  |  |  |  | Cond | 21G101301 |  | 599827 |  |  |  |  | pH | 7 | 21H2S |  |  |  |  | 8/11/2023
| Station | CPAH2 |  |  |  |  | Chl | [7 empty] | pH | 10 | 21G3T |  |  |  |  | 7/13/2023
[7 empty] | Wiper | 21H105435 |  | 599090-01 |  |  |  |  | Turb (FNU) | 124 | 21G21280150 |  |  |  | 11/1/2021 | 7/2022
|  |  |  | Wiper Replaced | Y |  |  |  | Batteries Replaced | Y | [5 empty] | Technician(s) | Michael Palandri
|  |  |  | Wiper Parked in Garage | Y |  |  |  | Date of Calibration | 2022-03-08 | [5 empty] | Date of Post-Calibration | 2022-04-18
| Pre-deployment Comments/ Failed-calibration values: | Waiting on back order 20mS/cm conductivity status. Depth and DO probe off initially before calibration, likely due to new sonde and probe never calibrated/deployed.
| Pre-Deployment (CAL & ICV) | [12 empty] | Post-Deployment (CCV) | [5 empty] | Programming
|  | Standards |  |  | Before Cal |  |  | Calibrated |  | Specs |  | P/ F | Time |  | Post value |  | P/ F | Time |  | Interval |  | 15
| Temp | 20.2 | °C (Therm) |  | 20.4 | °C (DL) |  |  |  | ±0.2(w) +0.1°C |  | P | 13:26 |  | 22.1 | °C (Therm) | P | 11:57 |  | Duration |  | 365 days
| SpCond in air |  |  |  | 0 |  |  |  |  | ≤ 0.002 mS/cm |  | P |  |  | 21.866 | °C (DL) |  |  |  | Free memory % |  | 100
| SpCond | 50 | mS/cm |  | 50.025 | mS/cm |  | 50 | mS/cm | 50 ± 0.25 (n) |  |  | 13:32 | ✔ | 50.38 | mS/cm | F | 12:05 |  | Memory (days) |  | 24542
[6 empty] | Cell Constant |  | 0.46876 |  | 0.419-0.519 (w) |  | P | [7 empty] | Start time |  | 7:45
| SpCond 20 (ICV) | 20 | mS/cm |  |  | ✔ |  | N/A | mS/cm | 20 ± 0.2 (w) | [9 empty] | Battery % |  | 100
| pH 7 | 7 |  |  | 7.41 |  |  | 7 |  | ± 0.1 |  |  | 13:36 | ✔ | 7.05 |  | P | 12:08 |  | Set clock |  | Y
| pH 10 | 10 |  |  | 10.37 |  |  | 10 |  | ± 0.1 |  |  | 13:37 | ✔ | 10.19 |  | F | 12:15 |  | End time |  | 13:00
[5 empty] | pH 7 |  |  | -24 | mV | 0 ±50 mV |  | P |  |  | -26.9 | mV | P |  |  |  | Other/notes
[5 empty] | pH 10 |  |  | -196.4 | mV | -180 ±50 mV |  | P |  |  | -207.6 | mV | P
[5 empty] | pH 4 |  |  |  | mV | +180 ±50 mV
[5 empty] | pH Slope |  |  | 172.4 |  | > 155 mV |  | P |  |  | 180.7 |  | P
| pH 4 or 10 (ICV) | 10 |  |  |  | ✔ |  | 9.97 |  | ± 0.1 |  | P | 13:39
| BP at Depth Cal |  |  |  | 766.572 | mmHg | [8 empty] | 760.5 | mmHg
| Depth | 0.095 | Offset |  | 10.52 | m |  | 0.095 | m | ±0.004 m |  | P | 13:45 |  | 0.036 | m | F | 12:01
| Turbidity | 0 | FNU |  | 1.19 |  |  | 0 | FNU | 0.00 ±0.3 |  |  | 13:47 | ✔ | -1.27 | FNU | F | 12:22 |  |  | General Maintenance
| Turbidity | 124 | FNU |  | 122.57 | FNU |  | 124 | FNU | 124 ±2.48 |  |  | 13:50 | ✔ | 122.36 | FNU | P | 12:27 | Changed shackle, sonde tube cleaned
| Turbidity 124 (ICV) | 124 | FNU |  |  |  | ✔ | 124.1 | FNU | 124 ±2.48 |  | P | 13:51
| BP at ODO Cal |  |  |  | 766.572 | mmHg | [8 empty] | 760.5 | mmHg
| Optical DO % |  |  |  | 87 | % |  | 100.7 | % | ± 1% |  | P | 13:59 |  | 106.1 | % | F | 11:58
[6 empty] | DO gain |  | 1.1594 | mg/L (table) | 0.87-1.25 |  | P |  |  | 106.2 | %
| Optical DO mg/L (ICV) | 9.09 | mg/L (DL) |  | 20.367 | °C |  | 9.092 |  | ± 0.1 mg/L |  | P | 14:00 |  | 9.09 | mg/L (DL)
| Battery Voltage | [6 empty] | 6.18 | V | ~6V (EXO 2) |  | P |  |  | 21.861 | °C | F | 12:00
[6 empty] | Optional Parameters | [8 empty] | 5.82 | V | P
| Chlorophyll | 0 | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| Chlorophyll |  | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [5 empty] | °C
| Chlorophyll | 0 | µg/L |  |  | µg/L |  |  | µg/L | [6 empty] | µg/L |  |  |  |  | Post-Deployment Notes
| Chlorophyll |  | µg/L |  |  | µg/L |  |  | µg/L | [6 empty] | µg/L
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [5 empty] | °C
| BGA | 0 | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| BGA |  | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [5 empty] | °C
Field Log
Fouling Presence | [9 empty] | Deployment |  |  | Retrieval |  |  |  | Field Notes:
A=algae, AM=amphipods, B=barnacles, BR=bryozoans, C=crabs, E=eggs, F=fish, FA=filamentous algae, G=gastropod, H=hydroids, I=isopod, MD=mud, S=sponges, SI=silt, SL=shell, SP=shrimp, T=tunicates, W=worm, O=other | [8 empty] | Date deploy/retrieval | 2022-03-16 |  |  | 2022-04-15 |  |  | Deployment:
[9 empty] | Time deploy/retrieval | 12:12 |  |  | 10:22 |  |  | SW Wind 10mph; Low tide rising
[9 empty] | First/last sample time | 12:15 |  |  | 10:15
| H= heavy, M= moderate, L= light | [7 empty] | Technician initials | MP |  |  | MP
Mount: N | [9 empty] | Discrete Samples (Required)
[9 empty] | Depth (m) | 2 | m |  | 1.7 |  | m
Sonde Body: B/L, MD/L, FA/L | [8 empty] | Temp (°C) | 16.1 | °C |  | 20.5 |  | °C
[9 empty] | DO % | 105.6 | % |  | 94.3 |  | %
Field guard: B/L, MD/M, FA/M, T/M | [8 empty] | DO mg/L | 8.55 | mg/L |  | 7.27 |  | mg/L
All Sesnors: | [8 empty] | SpCond (mS/cm) | 49.737 | mS/cm |  | 42.178 |  | mS/cm | Retrieval:
C/T: T/L, FA/L, MD/L | [8 empty] | Salinity (psu) | 32.61 | psu |  | 27.14 |  | psu | Turbidity: 4.1 NTU; sunny, wind NE 10-15mph, low tide rising.
pH: T/L, FA/L, MD/L | [8 empty] | pH | 8.17 |  |  | 7.92
DO: T/L, FA/L, MD/L | [11 empty] | Optional Samples
Turb: T/L, FA/L, MD/L | [8 empty] | pHmV |  | mV |  |  |  | mV
Wiper: T/L, FA/L, B/L, MD/L | [8 empty] | Conductivity |  | mS/cm |  |  |  | mS/cm
TAL-PE:
Other/notes:

## CPAH2_041522
Aquatic Preserve Continuous Water Quality Monitoring Calibration Log
|  | File Name | CPAH2_041522 | [5 empty] | .csv | [5 empty] | File Name | CPAH2_21E103416_041522_080000 |  |  |  |  | .bin
[7 empty] | Datasonde and Probe Identification Numbers | [12 empty] | Calibration Standards
|  | Sonde Code |  |  | Serial Number |  |  | Serial Number |  | Model Number | [5 empty] | Standard | Lot No. |  |  | Received | Opened | Exp. Date
| Datasonde | 1 |  |  | 21E103416 |  | pH | 21G105082 |  | pH Model |  |  |  |  | SpCond (mS/cm) | 20 | 4203H74 |  |  |  | 4/11/2022 | 3/2024
| Vented | N |  |  | Model |  | ODO | 21G100611 |  | ODO Model |  |  |  |  | SpCond (mS/cm) | 50 | 2906F57 |  |  |  |  | 6/2021
| Nickname | AH1 |  |  | EXO2 |  | Turb | 21G100846 |  | Turb Model |  |  |  |  | pH | 4
| Preserve | CPAP |  |  |  |  | Cond | 21G101304 |  | Cond Model |  |  |  |  | pH | 7 | 21H2S |  |  |  | 3/23/2022 | 8/11/2023
| Station | CPAH2 |  |  |  |  | Chl | Chl Serial No. |  | Chl Model |  |  |  |  | pH | 10 | 21G3T |  |  |  |  | 7/13/2023
[7 empty] | Wiper | 21H105437 |  | Wiper Model |  |  |  |  | Turb (FNU) | 124 | 607300 |  |  |  | 3/23/2022 | 11/2022
|  |  |  | Wiper Replaced | Y |  |  |  | Batteries Replaced | N | [5 empty] | Technician(s) | Michael Palandri
|  |  |  | Wiper Parked in Garage | Y |  |  |  | Date of Calibration | 2022-04-12 | [5 empty] | Date of Post-Calibration | 2022-05-19
| Pre-deployment Comments/ Failed-calibration values: | SpCond for 20 standard was reading high during ICV - most likely due to old 50 standard which is ordered. Was deployed to begin logging on 4/14/2022 at 8:00a.m., but due to weather could not deploy that day - was left on and logging but deployed the next day 4/15/2022.
| Pre-Deployment (CAL & ICV) | [12 empty] | Post-Deployment (CCV) | [5 empty] | Programming
|  | Standards |  |  | Before Cal |  |  | Calibrated |  | Specs |  | P/ F | Time |  | Post value |  | P/ F | Time |  | Interval |  | 15
| Temp | 20.8 | °C (Therm) |  | 20.778 | °C (DL) |  |  |  | ±0.2(w) +0.1°C |  | P | 11:00 |  | 23.8 | °C (Therm) | P | 9:46 |  | Duration |  | 365 days
| SpCond in air |  |  |  | 0 |  |  |  |  | ≤ 0.002 mS/cm |  | P |  |  | 23.642 | °C (DL) |  |  |  | Free memory %
| SpCond | 50 | mS/cm |  | 49.95 | mS/cm |  | 50 | mS/cm | 50 ± 0.25 (n) |  |  | 13:34 | ✔ | 49.94 | mS/cm | P | 9:54 |  | Memory (days) |  | 24461.11
[6 empty] | Cell Constant |  | 0.46982 |  | 0.419-0.519 (w) |  | P | [7 empty] | Start time |  | 8:00
| SpCond 20 (ICV) | 20 | mS/cm |  |  | ✔ |  | 20.361 | mS/cm | 20 ± 0.2 (w) |  | F | 13:37 | [6 empty] | Battery % |  | 79.5
| pH 7 | 7 |  |  | 7.09 |  |  | 7.02 |  | ± 0.1 |  |  | 13:44 | ✔ | 7.1 |  | P | 9:59 |  | Set clock |  | Y
| pH 10 | 10 |  |  | 10.07 |  |  | 10.04 |  | ± 0.1 |  |  | 13:48 | ✔ | 10.07 |  | P | 10:05 |  | End time |  | 16:04
[5 empty] | pH 7 |  |  | -20.3 | mV | 0 ±50 mV |  | P |  |  | -25.2 | mV | P |  |  |  | Other/notes
[5 empty] | pH 10 |  |  | -200.4 | mV | -180 ±50 mV |  | P |  |  | -201.7 | mV | P
[5 empty] | pH 4 |  |  |  | mV | +180 ±50 mV
[5 empty] | pH Slope |  |  | 180.1 |  | > 155 mV |  | P |  |  | 176.5 |  | P
| pH 4 or 10 (ICV) | 10 |  |  |  | ✔ |  | 10.03 |  | ± 0.1 |  | P | 13:48
| BP at Depth Cal |  |  |  | 765.5 | mmHg | [8 empty] | 761.24 | mmHg
| Depth | 0.082 | Offset |  | 0.085 | m |  | 0.082 | m | ±0.004 m |  | P | 13:52 |  | -0.005 | m | F | 9:50
| Turbidity | 0 | FNU |  | -0.21 |  |  | 0 | FNU | 0.00 ±0.3 |  |  | 14:01 | ✔ | 0.06 | FNU | P | 10:10 |  |  | General Maintenance
| Turbidity | 124 | FNU |  | 124.12 | FNU |  | 124 | FNU | 124 ±2.48 |  |  | 14:07 | ✔ | 122.46 | FNU | P | 10:15 | Cleaned sonde tube with ram rod, greased locks.
| Turbidity 124 (ICV) | 124 | FNU |  |  |  | ✔ | 124.2 | FNU | 124 ±2.48 |  | P | 14:08
| BP at ODO Cal |  |  |  | 765.5 | mmHg | [8 empty] | 761.24 | mmHg
| Optical DO % |  |  |  | 105.3 | % |  | 100.7 | % | ± 1% |  | P | 14:39 |  | 100.2 | % | P | 9:46
[6 empty] | DO gain |  | 1.044 | mg/L (table) | 0.87-1.25 |  | P |  |  | 100.4 | %
| Optical DO mg/L (ICV) | 9.01 | mg/L (DL) |  | 20.714 | °C |  | 8.915 |  | ± 0.1 mg/L |  | P | 14:41 |  | 8.35 | mg/L (DL)
| Battery Voltage | [6 empty] | 5.42 | V | ~6V (EXO 2) |  | P |  |  | 23.635 | °C | P | 9:47
[6 empty] | Optional Parameters | [8 empty] | 5.34 | V | P
| Chlorophyll | 0 | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| Chlorophyll |  | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [5 empty] | °C
| Chlorophyll | 0 | µg/L |  |  | µg/L |  |  | µg/L | [6 empty] | µg/L |  |  |  |  | Post-Deployment Notes
| Chlorophyll |  | µg/L |  |  | µg/L |  |  | µg/L | [6 empty] | µg/L |  |  | Cleaned sonde tube with ram rod and greased locks, heavy fouling. Thermometer for post-cal range: 21.8-23.9. Depth sensor had heavy fouling, may have led to post-cal reading.
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [5 empty] | °C
| BGA | 0 | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| BGA |  | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [5 empty] | °C
Field Log
Fouling Presence | [9 empty] | Deployment |  |  | Retrieval |  |  |  | Field Notes:
A=algae, AM=amphipods, B=barnacles, BR=bryozoans, C=crabs, E=eggs, F=fish, FA=filamentous algae, G=gastropod, H=hydroids, I=isopod, MD=mud, S=sponges, SI=silt, SL=shell, SP=shrimp, T=tunicates, W=worm, O=other | [8 empty] | Date deploy/retrieval | 2022-04-15 |  |  | 2022-05-10 |  |  | Deployment:
[9 empty] | Time deploy/retrieval | 10:25 |  |  | 14:25 |  |  | Turbidity: 4.1 NTU; Sunny, wind NE 10-15 mph, low tide rising
[9 empty] | First/last sample time | 10:30 |  |  | 14:15
| H= heavy, M= moderate, L= light | [7 empty] | Technician initials | MP |  |  | MP
Mount: N | [9 empty] | Discrete Samples (Required)
[9 empty] | Depth (m) | 1.7 | m |  | 1.7 |  | m
Sonde Body: MD/L; T/L | [8 empty] | Temp (°C) | 20.5 | °C |  | 23.1 |  | °C
[9 empty] | DO % | 94.3 | % |  | 105.8 |  | %
Field guard: A/H; E/H; T/H, B/L, C/L | [8 empty] | DO mg/L | 7.27 | mg/L |  | 7.7 |  | mg/L
All Sesnors: | [8 empty] | SpCond (mS/cm) | 42.178 | mS/cm |  | 44.614 |  | mS/cm | Retrieval:
C/T: T/H, A/L, B/L | [8 empty] | Salinity (psu) | 27.14 | psu |  | 28.85 |  | psu | Turbidity: 5.3 FNU. Sunny, warm, mid-tide rising, wind 5-10E.
pH: T/H, A/L, B/L | [8 empty] | pH | 7.92 |  |  | 7.62
DO: T/H, A/L, B/L | [11 empty] | Optional Samples
Turb: T/H, A/L, B/L | [8 empty] | pHmV |  | mV |  |  |  | mV
Wiper: T/H, A/L, B/L | [8 empty] | Conductivity |  | mS/cm |  |  |  | mS/cm
TAL-PE:
Other/notes:

## CPAH2_051022
Aquatic Preserve Continuous Water Quality Monitoring Calibration Log
|  | File Name | CPAH2_051022 | [5 empty] | .csv | [5 empty] | File Name | CPAH2_21E103418_051022_130000 |  |  |  |  | .bin
[7 empty] | Datasonde and Probe Identification Numbers | [12 empty] | Calibration Standards
|  | Sonde Code |  |  | Serial Number |  |  | Serial Number |  | Model Number | [5 empty] | Standard | Lot No. |  |  | Received | Opened | Exp. Date
| Datasonde | 2 |  |  | 21E103418 |  | pH | 21G105084 |  | 577602 |  |  |  |  | SpCond (mS/cm) | 20 | 4203H74 |  |  |  | 4/11/2022 | 3/2024
| Vented | N |  |  | Model |  | ODO | 21G100610 |  | 599100-01 |  |  |  |  | SpCond (mS/cm) | 50 | 2906F57 |  |  |  |  | 6/2021
| Nickname | AH2 |  |  | EXO2 |  | Turb | 21G100848 |  | 599101-01 |  |  |  |  | pH | 4
| Preserve | CPAP |  |  |  |  | Cond | 21G101301 |  | 599827 |  |  |  |  | pH | 7 | 21M2S |  |  |  |  | 12/8/2023
| Station | CPAH2 |  |  |  |  | Chl | [7 empty] | pH | 10 | 22A1T |  |  |  |  | 1/3/2024
[7 empty] | Wiper | 21H105435 |  | 599090-01 |  |  |  |  | Turb (FNU) | 124 | 21L21460003 |  |  |  | 3/23/2022 | 11/2022
|  |  |  | Wiper Replaced | Y |  |  |  | Batteries Replaced | N | [5 empty] | Technician(s) | Michael Palandri
|  |  |  | Wiper Parked in Garage | Y |  |  |  | Date of Calibration | 2022-05-09 | [5 empty] | Date of Post-Calibration | 2022-06-09
| Pre-deployment Comments/ Failed-calibration values: | Thermometer range: 23.2-24.1. High SpCond for ICV 20mS/cm likely due to old 50mS/cm standard, ordering.
| Pre-Deployment (CAL & ICV) | [12 empty] | Post-Deployment (CCV) | [5 empty] | Programming
|  | Standards |  |  | Before Cal |  |  | Calibrated |  | Specs |  | P/ F | Time |  | Post value |  | P/ F | Time |  | Interval |  | 15
| Temp | 23.8 | °C (Therm) |  | 23.506 | °C (DL) |  |  |  | ±0.2(w) +0.1°C |  | P | 14:36 |  | 21.5 | °C (Therm) | P | 8:14 |  | Duration |  | 365 days
| SpCond in air |  |  |  | 0 |  |  |  |  | ≤ 0.002 mS/cm |  | P |  |  | 21.35 | °C (DL) |  |  |  | Free memory %
| SpCond | 50 | mS/cm |  | 49.631 | mS/cm |  | 50 | mS/cm | 50 ± 0.25 (n) |  |  | 14:43 | ✔ | 50.766 | mS/cm | F | 8:38 |  | Memory (days) |  | 24513.98
[6 empty] | Cell Constant |  | 0.47225 |  | 0.419-0.519 (w) |  | P | [7 empty] | Start time |  | 13:00
| SpCond 20 (ICV) | 20 | mS/cm |  |  | ✔ |  | 20.59 | mS/cm | 20 ± 0.2 (w) |  | F | 14:50 | [6 empty] | Battery % |  | 85.4
| pH 7 | 7 |  |  | 7.1 |  |  | 7.01 |  | ± 0.1 |  |  | 14:56 | ✔ | 6.96 |  | P | 8:47 |  | Set clock |  | Y
| pH 10 | 10 |  |  | 10.23 |  |  | 10.04 |  | ± 0.1 |  |  | 15:00 | ✔ | 9.94 |  | P | 8:50 |  | End time |  | 12:43
[5 empty] | pH 7 |  |  | -29.6 | mV | 0 ±50 mV |  | P |  |  | -26 | mV | P |  |  |  | Other/notes
[5 empty] | pH 10 |  |  | -209.8 | mV | -180 ±50 mV |  | P |  |  | -204.2 | mV | P |  | Keep eye on this turbidity probe, read low in 124 FNU during post-cal.
[5 empty] | pH 4 |  |  |  | mV | +180 ±50 mV
[5 empty] | pH Slope |  |  | 180.2 |  | > 155 mV |  | P |  |  | 178.2 |  | P
| pH 4 or 10 (ICV) | 10 |  |  |  | ✔ |  | 10.02 |  | ± 0.1 |  | P | 15:00
| BP at Depth Cal |  |  |  | 762.4 | mmHg | [8 empty] | 761.54 | mmHg
| Depth | 0.027 | Offset |  | 0.019 | m |  | 0.027 | m | ±0.004 m |  | P | 15:05 |  | -0.016 | m | F | 8:33
| Turbidity | 0 | FNU |  | -1.23 |  |  | 0 | FNU | 0.00 ±0.3 |  |  | 15:13 | ✔ | -0.29 | FNU | P | 8:55 |  |  | General Maintenance
| Turbidity | 124 | FNU |  | 121.24 | FNU |  | 124 | FNU | 124 ±2.48 |  |  | 15:16 | ✔ | 117.55 | FNU | F | 8:59 | Cleaned tube with ram rod, greased locks.
| Turbidity 124 (ICV) | 124 | FNU |  |  |  | ✔ | 123.9 | FNU | 124 ±2.48 |  | P | 15:19
| BP at ODO Cal |  |  |  | 762.4 | mmHg | [8 empty] | 761.54 | mmHg
| Optical DO % |  |  |  | 105.9 | % |  | 100.3 | % | ± 1% |  | P | 15:30 |  | 103.8 | % | F | 8:13
[6 empty] | DO gain |  | 1.0984 | mg/L (table) | 0.87-1.25 |  | P |  |  | 103.7 | %
| Optical DO mg/L (ICV) | 8.6 | mg/L (DL) |  | 22.927 | °C |  | 8.578 |  | ± 0.1 mg/L |  | P | 15:31 |  | 9.01 | mg/L (DL)
| Battery Voltage | [6 empty] | 5.87 | V | ~6V (EXO 2) |  | P |  |  | 21.3 | °C | P | 9:01
[6 empty] | Optional Parameters | [8 empty] | 5.65 | V | P
| Chlorophyll | 0 | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| Chlorophyll |  | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [5 empty] | °C
| Chlorophyll | 0 | µg/L |  |  | µg/L |  |  | µg/L | [6 empty] | µg/L |  |  |  |  | Post-Deployment Notes
| Chlorophyll |  | µg/L |  |  | µg/L |  |  | µg/L | [6 empty] | µg/L |  |  | Sonde turned off on 6/6/2022 at 12:43.
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [5 empty] | °C
| BGA | 0 | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| BGA |  | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [5 empty] | °C
Field Log
Fouling Presence | [9 empty] | Deployment |  |  | Retrieval |  |  |  | Field Notes:
A=algae, AM=amphipods, B=barnacles, BR=bryozoans, C=crabs, E=eggs, F=fish, FA=filamentous algae, G=gastropod, H=hydroids, I=isopod, MD=mud, S=sponges, SI=silt, SL=shell, SP=shrimp, T=tunicates, W=worm, O=other | [8 empty] | Date deploy/retrieval | 2022-05-10 |  |  | 2022-06-02 |  |  | Deployment:
[9 empty] | Time deploy/retrieval | 14:28 |  |  | 12:12 |  |  | Turbidity: 5.3 FNU, Sunny, midd-tide rising, wind east 5-10mph.
[9 empty] | First/last sample time | 14:30 |  |  | 12:00
| H= heavy, M= moderate, L= light | [7 empty] | Technician initials | MP |  |  | MC/LC
Mount: None | [9 empty] | Discrete Samples (Required)
[9 empty] | Depth (m) | 1.7 | m |  | N/A |  | m
Sonde Body: None | [8 empty] | Temp (°C) | 23.1 | °C |  | 27 |  | °C
[9 empty] | DO % | 105.8 | % |  | 103.7 |  | %
Field guard: None | [8 empty] | DO mg/L | 7.7 | mg/L |  | 6.94 |  | mg/L
All Sesnors: | [8 empty] | SpCond (mS/cm) | 44.614 | mS/cm |  | 47.022 |  | mS/cm | Retrieval:
C/T: T/M | [8 empty] | Salinity (psu) | 28.85 | psu |  | 30.51 |  | psu | Turbidity: 9.8 FNU, sunny, mid-tide rising.
pH: T/M | [8 empty] | pH | 7.62 |  |  | 7.97
DO: T/M | [11 empty] | Optional Samples
Turb: T/M | [8 empty] | pHmV |  | mV |  |  |  | mV
Wiper: T/H | [8 empty] | Conductivity |  | mS/cm |  |  |  | mS/cm
TAL-PE: N/A
Other/notes:

## CPAH2_060222
Aquatic Preserve Continuous Water Quality Monitoring Calibration Log
|  | File Name | CPAH2_060222 | [5 empty] | .csv | [5 empty] | File Name | CPAH2_21E103416_060222_093000 |  |  |  |  | .bin
[7 empty] | Datasonde and Probe Identification Numbers | [12 empty] | Calibration Standards
|  | Sonde Code |  |  | Serial Number |  |  | Serial Number |  | Model Number | [5 empty] | Standard | Lot No. |  |  | Received | Opened | Exp. Date
| Datasonde | 1 |  |  | 21E103416 |  | pH | 21G105082 |  | 577602 |  |  |  |  | SpCond (mS/cm) | 20 | 4203H74 |  |  |  | 4/11/2022 | 3/2024
| Vented | N |  |  | Model |  | ODO | 21G100611 |  | 599100-01 |  |  |  |  | SpCond (mS/cm) | 50 | 2906F57 |  |  |  |  | 6/2021
| Nickname | AH1 |  |  | EXO2 |  | Turb | 21G100846 |  | 599101-01 |  |  |  |  | pH | 4
| Preserve | CPAP |  |  |  |  | Cond | 21G101304 |  | 599827 |  |  |  |  | pH | 7 | 21M2S |  |  |  |  | 12/8/2023
| Station | CPAH2 |  |  |  |  | Chl | Chl Serial No. |  | Chl Model |  |  |  |  | pH | 10 | 22A1T |  |  |  |  | 1/3/2024
[7 empty] | Wiper | 21H105437 |  | 599090-01 |  |  |  |  | Turb (FNU) | 124 | 21L21460003 |  |  |  | 3/23/2022 | 11/2022
|  |  |  | Wiper Replaced | N |  |  |  | Batteries Replaced | N | [5 empty] | Technician(s) | Michael Palandri
|  |  |  | Wiper Parked in Garage | N |  |  |  | Date of Calibration | 2022-06-01 | [5 empty] | Date of Post-Calibration | 2022-06-24
| Pre-deployment Comments/ Failed-calibration values: | ICV 20mS/cm reading high, likely due to old 50mS/cm standard.
| Pre-Deployment (CAL & ICV) | [12 empty] | Post-Deployment (CCV) | [5 empty] | Programming
|  | Standards |  |  | Before Cal |  |  | Calibrated |  | Specs |  | P/ F | Time |  | Post value |  | P/ F | Time |  | Interval |  | 15
| Temp | 23.4 | °C (Therm) |  | 23.255 | °C (DL) |  |  |  | ±0.2(w) +0.1°C |  | P | 9:32 |  | 20.5 | °C (Therm) | P | 8:59 |  | Duration |  | 365 days
| SpCond in air |  |  |  | 0 |  |  |  |  | ≤ 0.002 mS/cm |  | P |  |  | 20.391 | °C (DL) |  |  |  | Free memory %
| SpCond | 50 | mS/cm |  | 50.097 | mS/cm |  | 50 | mS/cm | 50 ± 0.25 (n) |  |  | 9:35 | ✔ | 49.758 | mS/cm | P | 9:11 |  | Memory (days) |  | 24435.72
[6 empty] | Cell Constant |  | 0.4689 |  | 0.419-0.519 (w) |  | P | [7 empty] | Start time |  | 9:30
| SpCond 20 (ICV) | 20 | mS/cm |  |  | ✔ |  | 20.659 | mS/cm | 20 ± 0.2 (w) |  | F | 9:38 | [6 empty] | Battery % |  | 81.8
| pH 7 | 7 |  |  | 7.07 |  |  | 7.02 |  | ± 0.1 |  |  | 9:43 | ✔ | 7.1 |  | P | 9:15 |  | Set clock |  | Y
| pH 10 | 10 |  |  | 10.08 |  |  | 10.04 |  | ± 0.1 |  |  | 9:47 | ✔ | 10.08 |  | P | 9:18 |  | End time |  | 12:26
[5 empty] | pH 7 |  |  | -23.3 | mV | 0 ±50 mV |  | P |  |  | -28.3 | mV | P |  |  |  | Other/notes
[5 empty] | pH 10 |  |  | -202.6 | mV | -180 ±50 mV |  | P |  |  | -205 | mV | P |  | Unsure if probes were zip tied or if wiper brush was added. Upon retrieval, found it was not zip tied or had wiper brush attached. Heavy sensor face fouling. Last two DO readings on were 5.76 mg/L at 27.8C. Listed mg/L and Temp for DO were in completely aerated water bath.
[5 empty] | pH 4 |  |  |  | mV | +180 ±50 mV
[5 empty] | pH Slope |  |  | 179.3 |  | > 155 mV |  | P |  |  | 176.7 |  | P
| pH 4 or 10 (ICV) | 10 |  |  |  | ✔ |  | 10.02 |  | ± 0.1 |  | P | 9:48
| BP at Depth Cal |  |  |  | 762.1 | mmHg | [8 empty] | 758.69 | mmHg
| Depth | 0.027 | Offset |  | -0.038 | m |  | 0.027 | m | ±0.004 m |  | P | 10:16 |  | -0.033 | m | F | 9:01
| Turbidity | 0 | FNU |  | -0.01 |  |  | 0 | FNU | 0.00 ±0.3 |  |  | 9:57 | ✔ | 1 | FNU | F | 9:23 |  |  | General Maintenance
| Turbidity | 124 | FNU |  | 121.66 | FNU |  | 124 | FNU | 124 ±2.48 |  |  | 10:15 | ✔ | 116.57 | FNU | F | 9:25 | Ram rod and grease on retrieval.
| Turbidity 124 (ICV) | 124 | FNU |  |  |  | ✔ | 123.88 | FNU | 124 ±2.48 |  | P | 10:16
| BP at ODO Cal |  |  |  | 762.1 | mmHg | [8 empty] | 758.69 | mmHg
| Optical DO % |  |  |  | 101.1 | % |  | 100.1 | % | ± 1% |  | P | 10:22 |  | 73.3 | % | F | 8:57
[6 empty] | DO gain |  | 0.99068 | mg/L (table) | 0.87-1.25 |  | P |  |  | 73.2 | %
| Optical DO mg/L (ICV) | 8.62 | mg/L (DL) |  | 22.821 | °C |  | 8.578 |  | ± 0.1 mg/L |  | P | 10:23 |  | 6.45 | mg/L (DL)
| Battery Voltage | [7 empty] | V | ~6V (EXO 2) |  |  |  |  | 20.397 | °C | F | 8:59
[6 empty] | Optional Parameters | [8 empty] | 5.23 | V | P
| Chlorophyll | 0 | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| Chlorophyll |  | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [5 empty] | °C
| Chlorophyll | 0 | µg/L |  |  | µg/L |  |  | µg/L | [6 empty] | µg/L |  |  |  |  | Post-Deployment Notes
| Chlorophyll |  | µg/L |  |  | µg/L |  |  | µg/L | [6 empty] | µg/L |  |  | Damage to field guard mesh. Heavy fouling on sensor faces due to no wiper brush deployed.
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [5 empty] | °C
| BGA | 0 | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| BGA |  | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [5 empty] | °C
Field Log
Fouling Presence | [9 empty] | Deployment |  |  | Retrieval |  |  |  | Field Notes:
A=algae, AM=amphipods, B=barnacles, BR=bryozoans, C=crabs, E=eggs, F=fish, FA=filamentous algae, G=gastropod, H=hydroids, I=isopod, MD=mud, S=sponges, SI=silt, SL=shell, SP=shrimp, T=tunicates, W=worm, O=other | [8 empty] | Date deploy/retrieval | 2022-06-02 |  |  | 2022-06-23 |  |  | Deployment:
[9 empty] | Time deploy/retrieval | 12:14 |  |  | 10:32 |  |  | Turbidity: 9.8FNU, sunny, mid-tide rising.
[9 empty] | First/last sample time | 12:15 |  |  | 10:30
| H= heavy, M= moderate, L= light | [7 empty] | Technician initials | MC/LC |  |  | MP
Mount: N | [9 empty] | Discrete Samples (Required)
[9 empty] | Depth (m) |  | m |  | 2.1 |  | m
Sonde Body: MD/L | [8 empty] | Temp (°C) | 27 | °C |  | 28.3 |  | °C
[9 empty] | DO % | 103.7 | % |  | 80.6 |  | %
Field guard: C/M, T/H, MD/M, E/M | [8 empty] | DO mg/L | 6.94 | mg/L |  | 5.27 |  | mg/L
All Sesnors: | [8 empty] | SpCond (mS/cm) | 47.022 | mS/cm |  | 47095 |  | mS/cm | Retrieval:
C/T: A/M, T/M | [8 empty] | Salinity (psu) | 30.51 | psu |  | 30.53 |  | psu | Turbidity: 3.6 FNU, sunny, N wind @ 5mph, high-tide falling
pH: A/M, T/M | [8 empty] | pH | 7.97 |  |  | 9.09
DO: A/L, T/L | [11 empty] | Optional Samples
Turb: A/M, T/M | [8 empty] | pHmV |  | mV |  |  |  | mV
Wiper: T/M, A/M, E/M | [8 empty] | Conductivity |  | mS/cm |  |  |  | mS/cm
TAL-PE:
Other/notes:

## CPAH2_062322
Aquatic Preserve Continuous Water Quality Monitoring Calibration Log
|  | File Name | CPAH2_062322 | [5 empty] | .csv | [5 empty] | File Name | CPAH2_21E103418_062322_061500 |  |  |  |  | .bin
[7 empty] | Datasonde and Probe Identification Numbers | [12 empty] | Calibration Standards
|  | Sonde Code |  |  | Serial Number |  |  | Serial Number |  | Model Number | [5 empty] | Standard | Lot No. |  |  | Received | Opened | Exp. Date
| Datasonde | 2 |  |  | 21E103418 |  | pH | 21G105084 |  | 577602 |  |  |  |  | SpCond (mS/cm) | 20 | 4203H74 |  |  |  | 4/11/2022 | 3/2024
| Vented | N |  |  | Model |  | ODO | 21G100610 |  | 599100-01 |  |  |  |  | SpCond (mS/cm) | 50 | 2966F57 |  |  |  |  | 6/2021
| Nickname | AH2 |  |  | EXO2 |  | Turb | 21G100848 |  | 599101-01 |  |  |  |  | pH | 4
| Preserve | CPAP |  |  |  |  | Cond | 21G101301 |  | 599827 |  |  |  |  | pH | 7 | 21M2S |  |  |  |  | 12/8/2023
| Station | CPAH2 |  |  |  |  | Chl | Chl Serial No. |  | Chl Model |  |  |  |  | pH | 10 | 22A1T |  |  |  |  | 1/3/2024
[7 empty] | Wiper | 21H105435 |  | 599090-01 |  |  |  |  | Turb (FNU) | 124 | 22D22130237 |  |  |  |  | 4/2023
|  |  |  | Wiper Replaced | Y |  |  |  | Batteries Replaced | N | [5 empty] | Technician(s) | Michael Palandri
|  |  |  | Wiper Parked in Garage | Y |  |  |  | Date of Calibration | 2022-06-22 | [5 empty] | Date of Post-Calibration | 2022-07-15
| Pre-deployment Comments/ Failed-calibration values: | ICV 20mS/cm reading hight (need new standard). Pre-cal 124 FNU reading low, but post-cal on last deployment showed low reading, keep eye on.
| Pre-Deployment (CAL & ICV) | [12 empty] | Post-Deployment (CCV) | [5 empty] | Programming
|  | Standards |  |  | Before Cal |  |  | Calibrated |  | Specs |  | P/ F | Time |  | Post value |  | P/ F | Time |  | Interval |  | 15
| Temp | 24.7 | °C (Therm) |  | 24.627 | °C (DL) |  |  |  | ±0.2(w) +0.1°C |  | P | 9:00 |  | 21.6 | °C (Therm) | P | 10:46 |  | Duration |  | 365 days
| SpCond in air |  |  |  | 0 |  |  |  |  | ≤ 0.002 mS/cm |  | P |  |  | 21.333 | °C (DL) |  |  |  | Free memory %
| SpCond | 50 | mS/cm |  | 50.504 | mS/cm |  | 50 | mS/cm | 50 ± 0.25 (n) |  |  | 9:04 | ✔ | 50.122 | mS/cm | P | 10:49 |  | Memory (days) |  | 24487.75
[6 empty] | Cell Constant |  | 0.46753 |  | 0.419-0.519 (w) |  | P | [7 empty] | Start time |  | 6:15
| SpCond 20 (ICV) | 20 | mS/cm |  |  | ✔ |  | 20.412 | mS/cm | 20 ± 0.2 (w) |  | F | 9:10 | [6 empty] | Battery % |  | 92.2
| pH 7 | 7 |  |  | 6.94 |  |  | 7.02 |  | ± 0.1 |  |  | 9:12 | ✔ | 7 |  | P | 10:51 |  | Set clock |  | Y
| pH 10 | 10 |  |  | 9.92 |  |  | 10.04 |  | ± 0.1 |  |  | 9:17 | ✔ | 10 |  | P | 10:54 |  | End time |  | 13:11
[5 empty] | pH 7 |  |  | -25.3 | mV | 0 ±50 mV |  | P |  |  | -24 | mV | P |  |  |  | Other/notes
[5 empty] | pH 10 |  |  | -202.6 | mV | -180 ±50 mV |  | P |  |  | -200.5 | mV | P
[5 empty] | pH 4 |  |  |  | mV | +180 ±50 mV
[5 empty] | pH Slope |  |  | 177.3 |  | > 155 mV |  | P |  |  | 176.5 |  | P
| pH 4 or 10 (ICV) | 10 |  |  |  | ✔ |  | 10.03 |  | ± 0.1 |  | P | 9:18
| BP at Depth Cal |  |  |  | 762.44 | mmHg | [8 empty] | 764.31 | mmHg
| Depth | 0.027 | Offset |  | -0.005 | m |  | 0.027 | m | ±0.004 m |  | P | 9:30 |  | 0.052 | m | P | 11:04
| Turbidity | 0 | FNU |  | -0.3 |  |  | 0 | FNU | 0.00 ±0.3 |  |  | 9:34 | ✔ | 0.05 | FNU | P | 10:57 |  |  | General Maintenance
| Turbidity | 124 | FNU |  | 111.63 | FNU |  | 124 | FNU | 124 ±2.48 |  |  | 9:46 | ✔ | 125.83 | FNU | P | 11:01 | Ram rod and lock grease on deployment and retrieval.
| Turbidity 124 (ICV) | 124 | FNU |  |  |  | ✔ | 124 | FNU | 124 ±2.48 |  | P | 9:47
| BP at ODO Cal |  |  |  | 762.29 | mmHg | [8 empty] | 764.39 | mmHg
| Optical DO % |  |  |  | 104.7 | % |  | 100.4 | % | ± 1% |  | P | 10:12 |  | 99.9 | % | P | 10:43
[6 empty] | DO gain |  | 1.0525 | mg/L (table) | 0.87-1.25 |  | P |  |  | 99.9 | %
| Optical DO mg/L (ICV) | 8.51 | mg/L (DL) |  | 23.486 | °C |  | 8.578 |  | ± 0.1 mg/L |  | P | 10:14 |  | 8.57 | mg/L (DL)
| Battery Voltage | [6 empty] | 5.71 | V | ~6V (EXO 2) |  | P |  |  | 22.899 | °C | P | 10:44
[6 empty] | Optional Parameters | [8 empty] | 5.53 | V | P
| Chlorophyll | 0 | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| Chlorophyll |  | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [5 empty] | °C
| Chlorophyll | 0 | µg/L |  |  | µg/L |  |  | µg/L | [6 empty] | µg/L |  |  |  |  | Post-Deployment Notes
| Chlorophyll |  | µg/L |  |  | µg/L |  |  | µg/L | [6 empty] | µg/L |  |  | Tunicate growing near depth senor, did not seem to affect readings.
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [5 empty] | °C
| BGA | 0 | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| BGA |  | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [5 empty] | °C
Field Log
Fouling Presence | [9 empty] | Deployment |  |  | Retrieval |  |  |  | Field Notes:
A=algae, AM=amphipods, B=barnacles, BR=bryozoans, C=crabs, E=eggs, F=fish, FA=filamentous algae, G=gastropod, H=hydroids, I=isopod, MD=mud, S=sponges, SI=silt, SL=shell, SP=shrimp, T=tunicates, W=worm, O=other | [8 empty] | Date deploy/retrieval | 2022-06-23 |  |  | 2022-07-13 |  |  | Deployment:
[9 empty] | Time deploy/retrieval | 10:34 |  |  | 7:36 |  |  | Turbidity: 3.6 FNU. Discrete pH sample high, probe was malfunctioning. Sunny, wind N @ 5mph, high-tide falling.
[9 empty] | First/last sample time | 10:45 |  |  | 7:30
| H= heavy, M= moderate, L= light | [7 empty] | Technician initials | MP |  |  | MP
Mount: N | [9 empty] | Discrete Samples (Required)
[9 empty] | Depth (m) | 2.1 | m |  | 1.9 |  | m
Sonde Body: MD/L, T/L | [8 empty] | Temp (°C) | 28.3 | °C |  | 25.6 |  | °C
[9 empty] | DO % | 80.6 | % |  | 76.5 |  | %
Field guard: T/H, SL/L, MD/L, A/L, C/M | [8 empty] | DO mg/L | 5.27 | mg/L |  | 5.26 |  | mg/L
All Sesnors: | [8 empty] | SpCond (mS/cm) | 47.095 | mS/cm |  | 47.4 |  | mS/cm | Retrieval:
C/T: A/L, B/L, T/M | [8 empty] | Salinity (psu) | 30.53 | psu |  | 30.83 |  | psu | Turbidity: 9.9 FNU, sunny, wind SW at 10-15mph, low tide ~1.5ft.
pH: A/L, B/L, T/M | [8 empty] | pH | 9.09* |  |  | 7.75
DO: A/L, B/M, T/L | [11 empty] | Optional Samples
Turb: A/L, B/L, T/M | [8 empty] | pHmV |  | mV |  |  |  | mV
Wiper: A/L, B/L | [8 empty] | Conductivity |  | mS/cm |  |  |  | mS/cm
TAL-PE:
Other/notes:

## CPAH2_071322
Aquatic Preserve Continuous Water Quality Monitoring Calibration Log
|  | File Name | CPAH2_071322 | [5 empty] | .csv | [5 empty] | File Name | CPAH2_21E103416_071322_060000 |  |  |  |  | .bin
[7 empty] | Datasonde and Probe Identification Numbers | [12 empty] | Calibration Standards
|  | Sonde Code |  |  | Serial Number |  |  | Serial Number |  | Model Number | [5 empty] | Standard | Lot No. |  |  | Received | Opened | Exp. Date
| Datasonde | 1 |  |  | 21E103416 |  | pH | 21G105082 |  | 577602 |  |  |  |  | SpCond (mS/cm) | 20 | 4203H74 |  |  |  | 4/11/2022 | 3/2024
| Vented | N |  |  | Model |  | ODO | 21G100611 |  | 599100-01 |  |  |  |  | SpCond (mS/cm) | 50 | 2908A58
| Nickname | AH1 |  |  | EXO2 |  | Turb | 21G100846 |  | 599101-01 |  |  |  |  | pH | 4
| Preserve | CPAP |  |  |  |  | Cond | 21G101304 |  | 599827 |  |  |  |  | pH | 7 | 21M2S |  |  |  | 7/11/2022 | 12/8/2023
| Station | CPAH2 |  |  |  |  | Chl | Chl Serial No. |  | Chl Model |  |  |  |  | pH | 10 | 22A1T |  |  |  | 7/11/2022 | 1/3/2024
[7 empty] | Wiper | 21H105437 |  | 599090-01 |  |  |  |  | Turb (FNU) | 124 | 22D22130237 |  |  |  |  | 4/2023
|  |  |  | Wiper Replaced | Y |  |  |  | Batteries Replaced | Y | [5 empty] | Technician(s) | Michael Palandri
|  |  |  | Wiper Parked in Garage | Y |  |  |  | Date of Calibration | 2022-07-11 | [5 empty] | Date of Post-Calibration | 2022-08-04
| Pre-deployment Comments/ Failed-calibration values: | SpCond ICV 20mS/cm standard reading high, need new 50 standard.
| Pre-Deployment (CAL & ICV) | [12 empty] | Post-Deployment (CCV) | [5 empty] | Programming
|  | Standards |  |  | Before Cal |  |  | Calibrated |  | Specs |  | P/ F | Time |  | Post value |  | P/ F | Time |  | Interval |  | 15
| Temp | 27.1 | °C (Therm) |  | 26.908 | °C (DL) |  |  |  | ±0.2(w) +0.1°C |  | P | 9:39 |  | 21.1 | °C (Therm) | P | 11:41 |  | Duration |  | 365 days
| SpCond in air |  |  |  | 0 |  |  |  |  | ≤ 0.002 mS/cm |  | P |  |  | 20.04 | °C (DL) |  |  |  | Free memory %
| SpCond | 50 | mS/cm |  | 49.811 | mS/cm |  | 50 | mS/cm | 50 ± 0.25 (n) |  |  | 9:45 | ✔ | 50.22 | mS/cm | P | 11:48 |  | Memory (days) |  | 24415.32
[6 empty] | Cell Constant |  | 0.47068 |  | 0.419-0.519 (w) |  | P | [7 empty] | Start time |  | 6:00
| SpCond 20 (ICV) | 20 | mS/cm |  |  | ✔ |  | 20.381 | mS/cm | 20 ± 0.2 (w) |  | F | 9:48 | [6 empty] | Battery % |  | 100
| pH 7 | 7 |  |  | 7.08 |  |  | 7.01 |  | ± 0.1 |  |  | 9:53 | ✔ | 6.9 |  | P | 11:55 |  | Set clock |  | Y
| pH 10 | 10 |  |  | 10.06 |  |  | 10.03 |  | ± 0.1 |  |  | 9:57 | ✔ | 9.9 |  | P | 12:10 |  | End time |  | 10:07
[5 empty] | pH 7 |  |  | -27 | mV | 0 ±50 mV |  | P |  |  | -20.3 | mV | P |  |  |  | Other/notes
[5 empty] | pH 10 |  |  | -204 | mV | -180 ±50 mV |  | P |  |  | -196.2 | mV | P
[5 empty] | pH 4 |  |  |  | mV | +180 ±50 mV
[5 empty] | pH Slope |  |  | 177 |  | > 155 mV |  | P |  |  | 175.9 |  | P
| pH 4 or 10 (ICV) | 10 |  |  |  | ✔ |  | 10.01 |  | ± 0.1 |  | P | 8:59
| BP at Depth Cal |  |  |  | 758.99 | mmHg | [8 empty] | 763.48 | mmHg
| Depth | -0.014 | Offset |  | -0.076 | m |  | -0.014 | m | ±0.004 m |  | P | 10:04 |  | 0.044 | m | P | 11:44
| Turbidity | 0 | FNU |  | 0 |  |  | 0 | FNU | 0.00 ±0.3 |  |  | 10:09 | ✔ | 0.01 | FNU | P | 12:12 |  |  | General Maintenance
| Turbidity | 124 | FNU |  | 123.28 | FNU |  | 123.99 | FNU | 124 ±2.48 |  |  | 10:13 | ✔ | 122.98 | FNU | P | 12:15 | Ram rod tube and lock grease on deployment and retrieval
| Turbidity 124 (ICV) | 124 | FNU |  |  |  | ✔ | 123.97 | FNU | 124 ±2.48 |  | P | 10:14
| BP at ODO Cal |  |  |  | 758.99 | mmHg | [8 empty] | 763.4 | mmHg
| Optical DO % |  |  |  | 100.6 | % |  | 100.2 | % | ± 1% |  | P | 10:20 |  | 98.8 | % | F | 11:37
[6 empty] | DO gain |  | 0.98354 | mg/L (table) | 0.87-1.25 |  | P |  |  | 98.7 | %
| Optical DO mg/L (ICV) | 8.12 | mg/L (DL) |  | 25.837 | °C |  | 8.113 |  | ± 0.1 mg/L |  | P | 10:21 |  | 7.41 | mg/L (DL)
| Battery Voltage | [6 empty] | 5.24 | V | ~6V (EXO 2) |  | P |  |  | 30.446 | °C | F | 11:38
[6 empty] | Optional Parameters | [8 empty] | 5.91 | V | P
| Chlorophyll | 0 | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| Chlorophyll |  | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [5 empty] | °C
| Chlorophyll | 0 | µg/L |  |  | µg/L |  |  | µg/L | [6 empty] | µg/L |  |  |  |  | Post-Deployment Notes
| Chlorophyll |  | µg/L |  |  | µg/L |  |  | µg/L | [6 empty] | µg/L |  |  | NST thermometer may be reading slightly high, all loggers reading same temps in water bath but the thermometer is consistently higher, keep eye on or purchase new thermometer.
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [5 empty] | °C
| BGA | 0 | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| BGA |  | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [5 empty] | °C
Field Log
Fouling Presence | [9 empty] | Deployment |  |  | Retrieval |  |  |  | Field Notes:
A=algae, AM=amphipods, B=barnacles, BR=bryozoans, C=crabs, E=eggs, F=fish, FA=filamentous algae, G=gastropod, H=hydroids, I=isopod, MD=mud, S=sponges, SI=silt, SL=shell, SP=shrimp, T=tunicates, W=worm, O=other | [8 empty] | Date deploy/retrieval | 2022-07-13 |  |  | 2022-08-03 |  |  | Deployment:
[9 empty] | Time deploy/retrieval | 7:36 |  |  | 12:02 |  |  | Turbidity: 9.9 FNU, sunny, wind SW 10-15mph, low tide at ~+1.5 feet.
[9 empty] | First/last sample time | 7:45 |  |  | 12:00
| H= heavy, M= moderate, L= light | [7 empty] | Technician initials | MP |  |  | MP
Mount: N/A | [9 empty] | Discrete Samples (Required)
[9 empty] | Depth (m) | 1.9 | m |  | 1.6 |  | m
Sonde Body: B/L, MD/L, T/L, Oysters/L | [8 empty] | Temp (°C) | 25.6 | °C |  | 28.8 |  | °C
[9 empty] | DO % | 76.5 | % |  | 91.5 |  | %
Field guard: T/H, C/M, MD/M | [8 empty] | DO mg/L | 5.26 | mg/L |  | 5.96 |  | mg/L
All Sesnors: | [8 empty] | SpCond (mS/cm) | 47.4 | mS/cm |  | 47.702 |  | mS/cm | Retrieval:
C/T: B/L, MD/M, T/L | [8 empty] | Salinity (psu) | 30.83 | psu |  | 30.96 |  | psu | Turbidity: 3.4 FNU, partly cloudy, calm wind 0-5mph SW, low tide rising.
pH: B/L, MD/M, T/L | [8 empty] | pH | 7.75 |  |  | 7.99
DO: B/L, MD/M, T/L | [11 empty] | Optional Samples
Turb: B/L, MD/M, T/L | [8 empty] | pHmV |  | mV |  |  |  | mV
Wiper: A/L, B/L, MD/M, T/L | [8 empty] | Conductivity |  | mS/cm |  |  |  | mS/cm
TAL-PE:
Other/notes:

## CPAH2_080322
Aquatic Preserve Continuous Water Quality Monitoring Calibration Log
|  | File Name | CPAH2_080322 | [5 empty] | .csv | [5 empty] | File Name | CPAH2_21E103418_080322_073000 |  |  |  |  | .bin
[7 empty] | Datasonde and Probe Identification Numbers | [12 empty] | Calibration Standards
|  | Sonde Code |  |  | Serial Number |  |  | Serial Number |  | Model Number | [5 empty] | Standard | Lot No. |  |  | Received | Opened | Exp. Date
| Datasonde | 2 |  |  | 21E103418 |  | pH | 21G105084 |  | 577602 |  |  |  |  | SpCond (mS/cm) | 20 | 4203H74 |  |  |  | 4/11/2022 | 3/2024
| Vented | N |  |  | Model |  | ODO | 21G100610 |  | 599100-01 |  |  |  |  | SpCond (mS/cm) | 50 | 2908A58 |  |  |  |  | 7/2021
| Nickname | AH2 |  |  | EXO2 |  | Turb | 21G100848 |  | 599101-01 |  |  |  |  | pH | 4
| Preserve | CPAP |  |  |  |  | Cond | 21G101301 |  | 599827 |  |  |  |  | pH | 7 | 21M2S |  |  |  | 7/13/2022 | 12/8/2023
| Station | CPAH2 |  |  |  |  | Chl | Chl Serial No. |  | Chl Model |  |  |  |  | pH | 10 | 22A1T |  |  |  | 8/2/2022 | 1/3/2024
[7 empty] | Wiper | 21H105435 |  | 599090-01 |  |  |  |  | Turb (FNU) | 124 | 22D22130237 |  |  |  |  | 4/2023
|  |  |  | Wiper Replaced | Y |  |  |  | Batteries Replaced | N | [5 empty] | Technician(s) | Michael Palandri
|  |  |  | Wiper Parked in Garage | Y |  |  |  | Date of Calibration | 2022-08-02 | [5 empty] | Date of Post-Calibration | 2022-08-23
| Pre-deployment Comments/ Failed-calibration values: | ICV 20mS/cm high, need new 50mS/cm std.
| Pre-Deployment (CAL & ICV) | [12 empty] | Post-Deployment (CCV) | [5 empty] | Programming
|  | Standards |  |  | Before Cal |  |  | Calibrated |  | Specs |  | P/ F | Time |  | Post value |  | P/ F | Time |  | Interval |  | 15
| Temp | 21.1 | °C (Therm) |  | 20.863 | °C (DL) |  |  |  | ±0.2(w) +0.1°C |  | P | 12:28 |  | 22.9 | °C (Therm) | P | 7:51 |  | Duration |  | 365 days
| SpCond in air |  |  |  | 0 |  |  |  |  | ≤ 0.002 mS/cm |  | P |  |  | 22.698 | °C (DL) |  |  |  | Free memory %
| SpCond | 50 | mS/cm |  | 50.065 | mS/cm |  | 50 | mS/cm | 50 ± 0.25 (n) |  |  | 12:30 | ✔ | 50.146 | mS/cm | P | 7:03 |  | Memory (days) |  | 24468.19
[6 empty] | Cell Constant |  | 0.46691 |  | 0.419-0.519 (w) |  | P | [7 empty] | Start time |  | 7:30
| SpCond 20 (ICV) | 20 | mS/cm |  |  | ✔ |  | 20.33 | mS/cm | 20 ± 0.2 (w) |  | F | 12:34 | [6 empty] | Battery % |  | 87.6
| pH 7 | 7 |  |  | 6.94 |  |  | 7.02 |  | ± 0.1 |  |  | 12:37 | ✔ | 7.04 |  | P | 7:08 |  | Set clock |  | Y
| pH 10 | 10 |  |  | 10.04 |  |  | 10.04 |  | ± 0.1 |  |  | 12:40 | ✔ | 10.09 |  | P | 7:13 |  | End time |  | 13:18
[5 empty] | pH 7 |  |  | -20.6 | mV | 0 ±50 mV |  | P |  |  | -21.9 | mV | P |  |  |  | Other/notes
[5 empty] | pH 10 |  |  | -197.2 | mV | -180 ±50 mV |  | P |  |  | -201.8 | mV | P
[5 empty] | pH 4 |  |  |  | mV | +180 ±50 mV
[5 empty] | pH Slope |  |  | 176.6 |  | > 155 mV |  | P |  |  | 179.9 |  | P
| pH 4 or 10 (ICV) | 10 |  |  |  | ✔ |  | 10.03 |  | ± 0.1 |  | P | 12:41
| BP at Depth Cal |  |  |  | 763.94 | mmHg | [8 empty] | 760.11 | mmHg
| Depth | 0.054 | Offset |  | 0.064 | m |  | 0.054 | m | ±0.004 m |  | P | 12:48 |  | 0.005 | m | F | 7:20
| Turbidity | 0 | FNU |  | 0.01 |  |  | 0 | FNU | 0.00 ±0.3 |  |  | 12:53 | ✔ | 0.02 | FNU | P | 7:26 |  |  | General Maintenance
| Turbidity | 124 | FNU |  | 121.59 | FNU |  | 123.99 | FNU | 124 ±2.48 |  |  | 12:56 | ✔ | 123.76 | FNU | P | 7:30 | Ram rod and lock grease on deployment and retrieval.
| Turbidity 124 (ICV) | 124 | FNU |  |  |  | ✔ | 124.2 | FNU | 124 ±2.48 |  | P | 12:57
| BP at ODO Cal |  |  |  | 763.41 | mmHg | [8 empty] | 760.04 | mmHg
| Optical DO % |  |  |  | 100.4 | % |  | 100.4 | % | ± 1% |  | P | 13:01 |  | 99.9 | % | P | 6:56
[6 empty] | DO gain |  | 1.0532 | mg/L (table) | 0.87-1.25 |  | P |  |  | 99.6 | %
| Optical DO mg/L (ICV) | 8.96 | mg/L (DL) |  | 20.926 | °C |  | 8.915 |  | ± 0.1 mg/L |  | P | 13:02 |  | 8.88 | mg/L (DL)
| Battery Voltage | [6 empty] | 5.55 | V | ~6V (EXO 2) |  | P |  |  | 21.075 | °C | P | 6:56
[6 empty] | Optional Parameters | [8 empty] | 5.41 | V | P
| Chlorophyll | 0 | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| Chlorophyll |  | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [5 empty] | °C
| Chlorophyll | 0 | µg/L |  |  | µg/L |  |  | µg/L | [6 empty] | µg/L |  |  |  |  | Post-Deployment Notes
| Chlorophyll |  | µg/L |  |  | µg/L |  |  | µg/L | [6 empty] | µg/L |  |  | Depth technically failed, but only by 1mm, should suffice as passing.
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [5 empty] | °C
| BGA | 0 | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| BGA |  | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [5 empty] | °C
Field Log
Fouling Presence | [9 empty] | Deployment |  |  | Retrieval |  |  |  | Field Notes:
A=algae, AM=amphipods, B=barnacles, BR=bryozoans, C=crabs, E=eggs, F=fish, FA=filamentous algae, G=gastropod, H=hydroids, I=isopod, MD=mud, S=sponges, SI=silt, SL=shell, SP=shrimp, T=tunicates, W=worm, O=other | [8 empty] | Date deploy/retrieval | 2022-08-03 |  |  | 2022-08-19 |  |  | Deployment:
[9 empty] | Time deploy/retrieval | 12:05 |  |  | 8:34 |  |  | Turbidity: 3.4 FNU, partly cloudy, calm wind 0-5mph SW, low tide rising
[9 empty] | First/last sample time | 12:15 |  |  | 8:30
| H= heavy, M= moderate, L= light | [7 empty] | Technician initials | MP |  |  | MP
Mount: MD/L | [9 empty] | Discrete Samples (Required)
[9 empty] | Depth (m) | 1.6 | m |  | 2 |  | m
Sonde Body: MD/L, T/L | [8 empty] | Temp (°C) | 28.8 | °C |  | 29.2 |  | °C
[9 empty] | DO % | 91.5 | % |  | 76.9 |  | %
Field guard: B/M, C/M, MD/L, T/H | [8 empty] | DO mg/L | 5.96 | mg/L |  | 4.97 |  | mg/L
All Sesnors: | [8 empty] | SpCond (mS/cm) | 47.702 | mS/cm |  | 47.143 |  | mS/cm | Retrieval:
C/T: A/M, MD/M, T/L | [8 empty] | Salinity (psu) | 30.96 | psu |  | 30.54 |  | psu | Turbidity: 5.73 NTU, calm, sunny, no wind, high tide
pH: A/M, B/L, MD/M, T/L | [8 empty] | pH | 7.99 |  |  | 7.91
DO: A/M, B/L, MD/M, T/L | [11 empty] | Optional Samples
Turb: A/M, B/L, MD/M, T/L | [8 empty] | pHmV |  | mV |  |  |  | mV
Wiper: A/M, B/M, MD/M | [8 empty] | Conductivity |  | mS/cm |  |  |  | mS/cm
TAL-PE:
Other/notes:

## CPAH2_081922
Aquatic Preserve Continuous Water Quality Monitoring Calibration Log
|  | File Name | [6 empty] | .csv | [5 empty] | File Name | [6 empty] | .bin
[7 empty] | Datasonde and Probe Identification Numbers | [13 empty] | Calibration Standards
|  | Sonde Code |  |  | Serial Number |  |  | Serial Number |  | Model Number | [5 empty] | Standard | Lot No. |  |  |  | Received | Opened | Exp. Date
| Datasonde | 1 |  |  | 21E103416 |  | pH | 21G105082 |  | 577602 |  |  |  |  | SpCond (mS/cm) | 20 | 4203H74 |  |  |  |  | 4/11/2022 | 3/2024
| Vented | Y/N |  |  | Model |  | ODO | 21G100611 |  | 599100-01 |  |  |  |  | SpCond (mS/cm) | 50 | 2908A58 | [5 empty] | 7/2021
| Nickname | AH1 |  |  | EXO2 |  | Turb | 21G100846 |  | 599101-01 |  |  |  |  | pH | 4
| Preserve | CPAP |  |  |  |  | Cond | 21G101304 |  | 599827 |  |  |  |  | pH | 7 | 21M2S | [5 empty] | 12/8/2023
| Station | CPAH2 |  |  |  |  | Chl | [7 empty] | pH | 10 | 22A1T | [5 empty] | 1/3/2024
[7 empty] | Wiper | 21H105437 |  | 599090-01 |  |  |  |  | Turb (FNU) | 124 | 22D22130237 | [5 empty] | 4/2023
|  |  |  | Wiper Replaced | Y |  |  |  | Batteries Replaced | N | [5 empty] | Technician(s) | Michael Palandri
|  |  |  | Wiper Parked in Garage | Y |  |  |  | Date of Calibration | 2022-08-18 | [5 empty] | Date of Post-Calibration
| Pre-deployment Comments/ Failed-calibration values: | ICV 20mS/cm high need spending authority for 50mS/cm standard.
| Pre-Deployment (CAL & ICV) | [12 empty] | Post-Deployment (CCV) | [6 empty] | Programming
|  | Standards |  |  | Before Cal |  |  | Calibrated |  | Specs |  | P/ F | Time |  | Post value |  | P/ F |  | Time |  | Interval |  | 15
| Temp | 27.8 | °C (Therm) |  | 27.548 | °C (DL) |  |  |  | ±0.2(w) +0.1°C |  | P | 9:40 |  | 22.1 | °C (Therm) | P |  | 6:59 |  | Duration |  | 365 days
| SpCond in air |  |  |  | 0 |  |  |  |  | ≤ 0.002 mS/cm |  | P |  |  | 21.88 | °C (DL) |  |  |  |  | Free memory %
| SpCond | 50 | mS/cm |  | 50.06 | mS/cm |  | 50 | mS/cm | 50 ± 0.25 (n) |  |  | 9:52 | ✔ | 51.145 | mS/cm | F |  | 7:02 |  | Memory (days) |  | 24394.09
[6 empty] | Cell Constant |  | 0.47012 |  | 0.419-0.519 (w) |  | P | [8 empty] | Start time |  | 6:27
| SpCond 20 (ICV) | 20 | mS/cm |  |  | ✔ |  | 20.443 | mS/cm | 20 ± 0.2 (w) |  | F | 10:01 | [7 empty] | Battery % |  | 91.7
| pH 7 | 7 |  |  | 6.97 |  |  | 7.01 |  | ± 0.1 |  |  | 10:07 | ✔ | 7.09 |  | P |  | 7:05 |  | Set clock |  | Y
| pH 10 | 10 |  |  | 10.04 |  |  | 10.03 |  | ± 0.1 |  |  | 10:10 | ✔ | 10.12 |  | P |  | 7:07 |  | End time
[5 empty] | pH 7 |  |  | -24.4 | mV | 0 ±50 mV |  | P |  |  | -42 | mV | P |  |  |  |  | Other/notes
[5 empty] | pH 10 |  |  | -204.9 | mV | -180 ±50 mV |  | P |  |  | -212.2 | mV | P
[5 empty] | pH 4 |  |  |  | mV | +180 ±50 mV
[5 empty] | pH Slope |  |  | 180.5 |  | > 155 mV |  | P |  |  | 170.2 |  | P
| pH 4 or 10 (ICV) | 10 |  |  |  | ✔ |  | 10.01 |  | ± 0.1 |  | P | 10:11
| BP at Depth Cal |  |  |  | 759.81 | mmHg | [8 empty] | 758.95 | mmHg
| Depth | 0 | Offset |  | 0.005 | m |  | 0 | m | ±0.004 m |  | P | 9:46 |  | -0.02 | m | P |  | 7:08
| Turbidity | 0 | FNU |  | 0.09 |  |  | 0 | FNU | 0.00 ±0.3 |  |  | 10:15 | ✔ | 0.53 | FNU | F |  | 7:09 |  |  | General Maintenance
| Turbidity | 124 | FNU |  | 121.7 | FNU |  | 124 | FNU | 124 ±2.48 |  |  | 10:19 | ✔ | 126.43 | FNU | P |  | 7:10 | Lock grease and ram rod on deployment, tube needs new top hose clamp
| Turbidity 124 (ICV) | 124 | FNU |  |  |  | ✔ | 123.99 | FNU | 124 ±2.48 |  | P | 10:23
| BP at ODO Cal |  |  |  | 760.04 | mmHg | [8 empty] | 758.95 | mmHg
| Optical DO % |  |  |  | 100.2 | % |  | 99.9 | % | ± 1% |  | P | 10:24 |  | 98.9 | % | P |  | 7:02
[6 empty] | DO gain |  | 0.098161 | mg/L (table) | 0.87-1.25 |  | P |  |  | 98.9 | %
| Optical DO mg/L (ICV) | 8.07 | mg/L (DL) |  | 26.28 | °C |  | 8.113 |  | ± 0.1 mg/L |  | P | 10:25 |  | 8.67 | mg/L (DL)
| Battery Voltage | [6 empty] | 5.95 | V | ~6V (EXO 2) |  | P |  |  | 22.1 | °C | P | 7:02 | 7:02
[6 empty] | Optional Parameters | [8 empty] | 5.49 | V
| Chlorophyll | 0 | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| Chlorophyll |  | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [5 empty] | °C
| Chlorophyll | 0 | µg/L |  |  | µg/L |  |  | µg/L | [6 empty] | µg/L | [5 empty] | Post-Deployment Notes
| Chlorophyll |  | µg/L |  |  | µg/L |  |  | µg/L | [6 empty] | µg/L |  |  |  | Updated firmware upon connecting to the computer when back in the lab.
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [5 empty] | °C
| BGA | 0 | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| BGA |  | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [5 empty] | °C
Field Log
Fouling Presence | [9 empty] | Deployment |  |  | Retrieval |  |  |  | Field Notes:
A=algae, AM=amphipods, B=barnacles, BR=bryozoans, C=crabs, E=eggs, F=fish, FA=filamentous algae, G=gastropod, H=hydroids, I=isopod, MD=mud, S=sponges, SI=silt, SL=shell, SP=shrimp, T=tunicates, W=worm, O=other | [8 empty] | Date deploy/retrieval | 2022-08-19 |  |  | 2022-09-20 |  |  | Deployment:
[9 empty] | Time deploy/retrieval | 8:36 |  |  | 9:01 |  |  | Turbidity: 5.73NTU, calm no wind, sunny, high tide.
[9 empty] | First/last sample time | 8:45 |  |  | 8:45
| H= heavy, M= moderate, L= light | [7 empty] | Technician initials | MP |  |  | JB
Mount: None | [9 empty] | Discrete Samples (Required)
[9 empty] | Depth (m) | 2 | m |  | 2.2 |  | m
Sonde Body: None | [8 empty] | Temp (°C) | 29.2 | °C |  | 28 |  | °C
[9 empty] | DO % | 76.9 | % |  | 95.4 |  | %
Field guard: L - B and other (goo) | [8 empty] | DO mg/L | 4.97 | mg/L |  | 6.33 |  | mg/L
All Sesnors: | [8 empty] | SpCond (mS/cm) | 47.143 | mS/cm |  | 45.767 |  | mS/cm | Retrieval: Turb = 7.89 FNU
C/T: L-O goo | [8 empty] | Salinity (psu) | 30.54 | psu |  | 29.98 |  | psu
pH: L-O (goo) | [8 empty] | pH | 7.91 |  |  | 8.11
DO: L-O (goo) | [11 empty] | Optional Samples
Turb: L-O (goo) | [8 empty] | pHmV |  | mV |  |  |  | mV
Wiper: None | [8 empty] | Conductivity |  | mS/cm |  |  |  | mS/cm
TAL-PE:
Other/notes:

## CPAH2_092022
Aquatic Preserve Continuous Water Quality Monitoring Calibration Log
|  | File Name | CPAH2_092022 | [5 empty] | .csv | [5 empty] | File Name | CPAH2_21E103418_092022_06000 |  |  |  | .bin
[7 empty] | Datasonde and Probe Identification Numbers | [11 empty] | Calibration Standards
|  | Sonde Code |  |  | Serial Number |  |  | Serial Number |  | Model Number | [5 empty] | Standard | Lot No. |  | Received | Opened | Exp. Date
| Datasonde | 2 |  |  | 21E103418 |  | pH | 21G105084 |  | 577602 |  |  |  |  | SpCond (mS/cm) | 20 | 4203H74 |  |  | 4/11/2022 | 3/2024
| Vented | N |  |  | Model |  | ODO | 21G100610 |  | 599100-01 |  |  |  |  | SpCond (mS/cm) | 50 | 2908A58 |  |  |  | 7/2021
| Nickname | AH2 |  |  | EXO2 |  | Turb | 21G100848 |  | 599101-01 |  |  |  |  | pH | 4
| Preserve | CPAP |  |  |  |  | Cond | 21G101301 |  | 599827 |  |  |  |  | pH | 7 | 21M2S |  |  | 7/13/2022 | 12/8/2023
| Station | CPAH2 |  |  |  |  | Chl | Chl Serial No. |  | Chl Model |  |  |  |  | pH | 10 | 22A1T |  |  | 8/2/2022 | 1/3/2024
[7 empty] | Wiper | 21H105435 |  | 599090-01 |  |  |  |  | Turb (FNU) | 124 | 22D22130237 |  |  |  | 4/2023
|  |  |  | Wiper Replaced | Y |  |  |  | Batteries Replaced | N | [5 empty] | Technician(s) | Jon Brucker
|  |  |  | Wiper Parked in Garage | Y |  |  |  | Date of Calibration | 2022-09-12 | [5 empty] | Date of Post-Calibration
| Pre-deployment Comments/ Failed-calibration values: | Temp off in calibration - thought to be due to failing lab thermometer. Sp cond still reading off in 20 mS/cm standard - not sure why. Depth off in terms of specs in cal, but not sure why.
| Pre-Deployment (CAL & ICV) | [12 empty] | Post-Deployment (CCV) |  |  |  |  | Programming
|  | Standards |  |  | Before Cal |  |  | Calibrated |  | Specs |  | P/ F | Time |  | Post value |  | P/ F | Time | Interval |  | 15
| Temp | 22.4 | °C (Therm) |  | 21.881 | °C (DL) |  |  |  | ±0.2(w) +0.1°C |  | F | 12:44 |  | 19.8 | °C (Therm) | P | 6:50 | Duration |  | 365 days
| SpCond in air |  |  |  | 0.001 |  |  |  |  | ≤ 0.002 mS/cm |  | P |  |  | 20.1 | °C (DL) |  |  | Free memory % |  | 99
| SpCond | 50 | mS/cm |  | 50.258 | mS/cm |  | 50 | mS/cm | 50 ± 0.25 (n) |  |  | 12:46 | ✔ | 49.34 | mS/cm | P | 6:52 | Memory (days) |  | 24452
[6 empty] | Cell Constant |  | 0.46 |  | 0.419-0.519 (w) |  | P | [6 empty] | Start time |  | 5:50
| SpCond 20 (ICV) | 20 | mS/cm |  |  | ✔ |  | 22.591 | mS/cm | 20 ± 0.2 (w) |  | F | 12:47 | [5 empty] | Battery % |  | 85
| pH 7 | 7 |  |  | 7.14 |  |  | 7 |  | ± 0.1 |  |  | 12:48 | ✔ | 7.26 |  | P | 6:53 | Set clock |  | Y
| pH 10 | 10 |  |  | 10.15 |  |  | 10 |  | ± 0.1 |  |  | 12:49 | ✔ | 10.11 |  | P | 6:54 | End time |  | 9:53
[5 empty] | pH 7 |  |  | -27.7 | mV | 0 ±50 mV |  | P |  |  | -41.4 | mV | P |  |  | Other/notes
[5 empty] | pH 10 |  |  | -203.9 | mV | -180 ±50 mV |  | P |  |  | -207.2 | mV | P
[5 empty] | pH 4 |  |  | N/A | mV | +180 ±50 mV
[5 empty] | pH Slope |  |  | 176.2 |  | > 155 mV |  | P |  |  | 165.8 |  | P
| pH 4 or 10 (ICV) | 10 |  |  |  | ✔ |  | 10.03 |  | ± 0.1 |  | P | 12:51
| BP at Depth Cal |  |  |  | 760.73 | mmHg | [8 empty] | 763.016 | mmHg
| Depth | 0.014 | Offset |  | -0.033 | m |  | 0.014 | m | ±0.004 m |  | F | 12:55 |  | 0.042 | m | P | 6:49
| Turbidity | 0 | FNU |  | 0.34 |  |  | 0 | FNU | 0.00 ±0.3 |  |  | 12:56 | ✔ | 0.2 | FNU | P | 6:56 |  | General Maintenance
| Turbidity | 124 | FNU |  | 123.86 | FNU |  | 124 | FNU | 124 ±2.48 |  |  | 12:58 | ✔ | 125.01 | FNU | P | 6:58
| Turbidity 124 (ICV) | 124 | FNU |  |  |  | ✔ | 123.99 | FNU | 124 ±2.48 |  | P | 12:59
| BP at ODO Cal |  |  |  | 760.73 | mmHg | [8 empty] | 763.016 | mmHg
| Optical DO % |  |  |  | 101.4 | % |  | 100.1 | % | ± 1% |  | P | 13:02 |  | 100.1 | % | P | 6:48
[6 empty] | DO gain |  | 1.03 | mg/L (table) | 0.87-1.25 |  | P |  |  | 100.1 | %
| Optical DO mg/L (ICV) | 8.77 | mg/L (DL) |  | 21.809 | °C |  | 8.743 |  | ± 0.1 mg/L |  | P | 13:03 |  | 9.11 | mg/L (DL)
| Battery Voltage | [7 empty] | V | ~6V (EXO 2) |  | P |  |  | 19.8 | °C | P | 6:48
[6 empty] | Optional Parameters | [9 empty] | V | P
| Chlorophyll | 0 | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| Chlorophyll |  | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [5 empty] | °C
| Chlorophyll | 0 | µg/L |  |  | µg/L |  |  | µg/L | [6 empty] | µg/L |  |  |  | Post-Deployment Notes
| Chlorophyll |  | µg/L |  |  | µg/L |  |  | µg/L | [6 empty] | µg/L
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [5 empty] | °C
| BGA | 0 | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| BGA |  | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [5 empty] | °C
Field Log
Fouling Presence | [9 empty] | Deployment |  |  | Retrieval |  |  |  | Field Notes:
A=algae, AM=amphipods, B=barnacles, BR=bryozoans, C=crabs, E=eggs, F=fish, FA=filamentous algae, G=gastropod, H=hydroids, I=isopod, MD=mud, S=sponges, SI=silt, SL=shell, SP=shrimp, T=tunicates, W=worm, O=other | [8 empty] | Date deploy/retrieval | 2022-09-20 |  |  | 2022-10-24 |  |  | Deployment:
[9 empty] | Time deploy/retrieval | 9:05 |  |  | 11:18
[9 empty] | First/last sample time | 9:15 |  |  | 11:15
| H= heavy, M= moderate, L= light | [7 empty] | Technician initials | JB |  |  | JB
Mount: | [9 empty] | Discrete Samples (Required)
[9 empty] | Depth (m) | 2.2 | m |  | 2.3 |  | m
Sonde Body: | [8 empty] | Temp (°C) | 28 | °C |  | 16.8 |  | °C
[9 empty] | DO % | 95.4 | % |  | 98 |  | %
Field guard: | [8 empty] | DO mg/L | 6.33 | mg/L |  | 7.79 |  | mg/L
All Sesnors: | [8 empty] | SpCond (mS/cm) | 45.767 | mS/cm |  | 50.191 |  | mS/cm | Retrieval:
C/T: | [8 empty] | Salinity (psu) | 29.58 | psu |  | 32.95 |  | psu | Turbidity:
pH: | [8 empty] | pH | 8.11 |  |  | 8.01
DO: | [11 empty] | Optional Samples
Turb: | [8 empty] | pHmV |  | mV |  |  |  | mV
Wiper: | [8 empty] | Conductivity |  | mS/cm |  |  |  | mS/cm
TAL-PE:
Other/notes:

## CPAH2_102422
Aquatic Preserve Continuous Water Quality Monitoring Calibration Log
|  | File Name | CPAH2_102422 | [5 empty] | .csv | [5 empty] | File Name | [6 empty] | .bin
[7 empty] | Datasonde and Probe Identification Numbers | [13 empty] | Calibration Standards
|  | Sonde Code |  |  | Serial Number |  |  | Serial Number |  | Model Number | [5 empty] | Standard | Lot No. |  |  |  | Received | Opened | Exp. Date
| Datasonde | 1 |  |  | 21E103416 |  | pH | 21G105082 |  | 577602 |  |  |  |  | SpCond (mS/cm) | 20 | 4203H74 |  |  |  |  | 4/11/2022 | 3/2024
| Vented | Y/N |  |  | Model |  | ODO | 21G100611 |  | 599100-01 |  |  |  |  | SpCond (mS/cm) | 50 | 2908A58 | [5 empty] | 7/2021
| Nickname | AH1 |  |  | EXO2 |  | Turb | 21G100846 |  | 599101-01 |  |  |  |  | pH | 4
| Preserve | CPAP |  |  |  |  | Cond | 21G101304 |  | 599827 |  |  |  |  | pH | 7 | 21M2S | [5 empty] | 12/8/2023
| Station | CPAH2 |  |  |  |  | Chl | [7 empty] | pH | 10 | 22A1T | [5 empty] | 1/3/2024
[7 empty] | Wiper | 21H105437 |  | 599090-01 |  |  |  |  | Turb (FNU) | 124 | 22D22130237 | [5 empty] | 4/2023
|  |  |  | Wiper Replaced | Y |  |  |  | Batteries Replaced | N | [5 empty] | Technician(s) | Jon Brucker
|  |  |  | Wiper Parked in Garage | Y |  |  |  | Date of Calibration | 2022-10-10 | [5 empty] | Date of Post-Calibration | 2022-11-30
| Pre-deployment Comments/ Failed-calibration values: | ICV 20mS/cm high need spending authority for 50mS/cm standard.
| Pre-Deployment (CAL & ICV) | [12 empty] | Post-Deployment (CCV) | [6 empty] | Programming
|  | Standards |  |  | Before Cal |  |  | Calibrated |  | Specs |  | P/ F | Time |  | Post value |  | P/ F |  | Time |  | Interval |  | 15
| Temp | 22.9 | °C (Therm) |  | 22.317 | °C (DL) |  |  |  | ±0.2(w) +0.1°C |  | P | 12:14 |  | 20.44 | °C (Therm) | P |  | 11:37 |  | Duration |  | 365 days
| SpCond in air |  |  |  | 0 |  |  |  |  | ≤ 0.002 mS/cm |  | P |  |  | 20.09 | °C (DL) |  |  |  |  | Free memory % |  | 100
| SpCond | 50 | mS/cm |  | 49.6 | mS/cm |  | 50 | mS/cm | 50 ± 0.25 (n) |  |  | 12:16 | ✔ | 49.1 | mS/cm | F |  | 11:48 |  | Memory (days) |  | 24363
[6 empty] | Cell Constant |  | 0.47 |  | 0.419-0.519 (w) |  | P | [8 empty] | Start time |  | 7:45
| SpCond 20 (ICV) | 20 | mS/cm |  |  | ✔ |  | 21.2 | mS/cm | 20 ± 0.2 (w) |  | F | 12:18 | [7 empty] | Battery % |  | 100
| pH 7 | 7 |  |  | 7.09 |  |  | 7.02 |  | ± 0.1 |  |  | 12:20 | ✔ | 7.1 |  | P |  | 11:52 |  | Set clock |  | Y
| pH 10 | 10 |  |  | 10.04 |  |  | 10.05 |  | ± 0.1 |  |  | 12:22 | ✔ | 10.05 |  | P |  | 11:53 |  | End time |  | 15:30
[5 empty] | pH 7 |  |  | -29.1 | mV | 0 ±50 mV |  | P |  |  | -33.5 | mV | P |  |  |  |  | Other/notes
[5 empty] | pH 10 |  |  | -204.7 | mV | -180 ±50 mV |  | P |  |  | -203.7 | mV | P
[5 empty] | pH 4 |  |  |  | mV | +180 ±50 mV
[5 empty] | pH Slope |  |  | 175.6 |  | > 155 mV |  | P |  |  | 170.2 |  | P
| pH 4 or 10 (ICV) | 10 |  |  |  | ✔ |  | 10.02 |  | ± 0.1 |  | P | 12:24
| BP at Depth Cal |  |  |  | 766.572 | mmHg | [8 empty] | 763.27 | mmHg
| Depth | 0 | Offset |  | 0.145 | m |  | 0 | m | ±0.004 m |  | P | 12:26 |  | 0.056 | m | P |  | 12:29
| Turbidity | 0 | FNU |  | 0.66 |  |  | 0 | FNU | 0.00 ±0.3 |  |  | 12:28 | ✔ | 1.97 | FNU | F |  | 12:30 |  |  | General Maintenance
| Turbidity | 124 | FNU |  | 121.88 | FNU |  | 124 | FNU | 124 ±2.48 |  |  | 12:34 | ✔ | 133.36 | FNU | P |  | 12:35
| Turbidity 124 (ICV) | 124 | FNU |  |  |  | ✔ | 123.59 | FNU | 124 ±2.48 |  | P | 12:35
| BP at ODO Cal |  |  |  | 766.572 | mmHg | [8 empty] | 765.81 | mmHg
| Optical DO % |  |  |  | 99.4 | % |  | 100.6 | % | ± 1% |  | P | 12:36 |  | 100.6 | % | P |  | 11:39
[6 empty] | DO gain |  | 1 | mg/L (table) | 0.87-1.25 |  | P |  |  | 100.7 | %
| Optical DO mg/L (ICV) | 8.81 |  |  | 22.014 | °C |  | 8.743 |  | ± 0.1 mg/L |  | P | 12:38 |  | 9.3 | mg/L (DL)
| Battery Voltage | [6 empty] | 6.37 | V | ~6V (EXO 2) |  | P |  |  | 19.4 | °C | P | 7:02 | 11:36
[6 empty] | Optional Parameters | [8 empty] | 5.7 | V
| Chlorophyll | 0 | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| Chlorophyll |  | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [5 empty] | °C
| Chlorophyll | 0 | µg/L |  |  | µg/L |  |  | µg/L | [6 empty] | µg/L | [5 empty] | Post-Deployment Notes
| Chlorophyll |  | µg/L |  |  | µg/L |  |  | µg/L | [6 empty] | µg/L |  |  |  | BP= 763.27 ON 11/30/22 AND BP= 765.81 ON 11/14/22
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [5 empty] | °C
| BGA | 0 | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| BGA |  | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [5 empty] | °C
Field Log
Fouling Presence | [9 empty] | Deployment |  |  | Retrieval |  |  |  | Field Notes:
A=algae, AM=amphipods, B=barnacles, BR=bryozoans, C=crabs, E=eggs, F=fish, FA=filamentous algae, G=gastropod, H=hydroids, I=isopod, MD=mud, S=sponges, SI=silt, SL=shell, SP=shrimp, T=tunicates, W=worm, O=other | [8 empty] | Date deploy/retrieval | 2022-10-24 |  |  | 2022-11-14 |  |  | Deployment:
[9 empty] | Time deploy/retrieval | 11:20 |  |  | 15:32 |  |  | Turbidity:.
[9 empty] | First/last sample time | 11:15 |  |  | 15:30
| H= heavy, M= moderate, L= light | [7 empty] | Technician initials | JB |  |  | AL
Mount: MD-L
[9 empty] | Depth (m) | 2.3 |  |  | 2.2 |  | m
Sonde Body: MD-L | [8 empty] | Temp (°C) | 16.8 |  |  | 16.5 |  | °C
[9 empty] | DO % | 98 |  |  | 117.3 |  | %
Field guard: T-L, B-L | [8 empty] | DO mg/L | 7.79 |  |  | 9.38 |  | mg/L
All Sesnors: | [8 empty] | SpCond (mS/cm) | 50.191 |  |  | 50.361 |  | mS/cm | Retrieval:
C/T: T-L | [8 empty] | Salinity (psu) | 32.95 |  |  | 33.05 |  | psu
pH: T-L | [8 empty] | pH | 8.01 |  |  | 8.11
DO: T-L | [11 empty] | Optional Samples
Turb: T-L | [8 empty] | pHmV |  | mV |  |  |  | mV
Wiper: T-L | [8 empty] | Conductivity |  | mS/cm |  |  |  | mS/cm
TAL-PE: T-L
Other/notes:

## CPAH2_111422
Aquatic Preserve Continuous Water Quality Monitoring Calibration Log
|  | File Name | CPAH2_111422 | [5 empty] | .csv | [5 empty] | File Name | CPAH2_21E103418_111422_0131500 | [5 empty] | .bin
[7 empty] | Datasonde and Probe Identification Numbers | [13 empty] | Calibration Standards
|  | Sonde Code |  |  | Serial Number |  |  | Serial Number |  | Model Number | [5 empty] | Standard | Lot No. |  |  |  | Received | Opened | Exp. Date
| Datasonde | 2 |  |  | 21E103416 |  | pH | 21G105082 |  | 577602 |  |  |  |  | SpCond (mS/cm) | 20 | 4203H74 |  |  |  |  | 4/11/2022 | 3/2024
| Vented | N |  |  | Model |  | ODO | 21G100611 |  | 599100-01 |  |  |  |  | SpCond (mS/cm) | 50 | 2908A58 | [5 empty] | 7/2021
| Nickname | AH2 |  |  | EXO2 |  | Turb | 21G100846 |  | 599101-01 |  |  |  |  | pH | 4
| Preserve | CPAP |  |  |  |  | Cond | 21G101304 |  | 599827 |  |  |  |  | pH | 7 | 21M2S | [5 empty] | 12/8/2023
| Station | CPAH2 |  |  |  |  | Chl | [7 empty] | pH | 10 | 22A1T | [5 empty] | 1/3/2024
[7 empty] | Wiper | 21H105437 |  | 599090-01 |  |  |  |  | Turb (FNU) | 124 | 22D22130237 | [5 empty] | 4/2023
|  |  |  | Wiper Replaced | Y |  |  |  | Batteries Replaced | N | [5 empty] | Technician(s) | Jon Brucker
|  |  |  | Wiper Parked in Garage | Y |  |  |  | Date of Calibration | 2022-11-09 | [5 empty] | Date of Post-Calibration | 2022-12-08
Pre-deployment Comments/ Failed-calibration values:
| Pre-Deployment (CAL & ICV) | [12 empty] | Post-Deployment (CCV) | [6 empty] | Programming
|  | Standards |  |  | Before Cal |  |  | Calibrated |  | Specs |  | P/ F | Time |  | Post value |  | P/ F |  | Time |  | Interval |  | 15
| Temp | 22.1 | °C (Therm) |  | 21.954 | °C (DL) |  |  |  | ±0.2(w) +0.1°C |  | F | 13:36 |  | 20.48 | °C (Therm) | P |  | 11:09 |  | Duration |  | 365 days
| SpCond in air |  |  |  | 0 |  |  |  |  | ≤ 0.002 mS/cm |  | P |  |  | 20.503 | °C (DL) |  |  |  |  | Free memory % |  | 100
| SpCond | 50 | mS/cm |  | 49.152 | mS/cm |  | 50 | mS/cm | 50 ± 0.25 (n) |  |  | 13:45 | ✔ | 47.863 | mS/cm | F |  | 11:12 |  | Memory (days) |  | 24417
[6 empty] | Cell Constant |  | 0.47 |  | 0.419-0.519 (w) |  | P | [8 empty] | Start time |  | 13:15
| SpCond 20 (ICV) | 20 | mS/cm |  |  | ✔ |  | 20.534 | mS/cm | 20 ± 0.2 (w) |  | F | 13:48 | [7 empty] | Battery % |  | 79.9
| pH 7 | 7 |  |  | 7.02 |  |  | 7.05 |  | ± 0.1 |  |  | 13:53 | ✔ | 7.15 |  | P |  | 11:14 |  | Set clock |  | Y
| pH 10 | 10 |  |  | 10.05 |  |  | 10.13 |  | ± 0.1 |  |  | 13:54 | ✔ | 10.08 |  | P |  | 11:16 |  | End time |  | 15:56
[8 empty] | -29.8 | mV | 0 ±50 mV |  | P |  |  | -37.4 | mV | P |  |  |  |  | Other/notes
[8 empty] | -208.5 | mV | -180 ±50 mV |  | P |  |  | -208.9 | mV | P
[9 empty] | mV | +180 ±50 mV
[8 empty] | 178.7 |  | > 155 mV |  | P |  |  | 171.5 |  | P
| pH 4 or 10 (ICV) | 10 |  |  |  | ✔ |  | 10.04 |  | ± 0.1 |  | F | 13:58
| BP at Depth Cal |  |  |  | 766.04 | mmHg | [8 empty] | 755.25 | mmHg
| Depth | 0.082 | Offset |  | 0.13 | m |  | 0.082 | m | ±0.004 m |  | P | 14:02 |  | 0.129 | m | P |  | 11:23
| Turbidity | 0 | FNU |  | -0.04 |  |  | 0 | FNU | 0.00 ±0.3 |  |  | 14:05 | ✔ | 1.61 | FNU | F |  | 11:19 |  |  | General Maintenance
| Turbidity | 124 | FNU |  | 125.57 | FNU |  | 124 | FNU | 124 ±2.48 |  |  | 14:09 | ✔ | 124.25 | FNU | P |  | 11:21 | PVC lid replaced on 12/07/2022
| Turbidity 124 (ICV) | 124 | FNU |  |  |  | ✔ | 123.98 | FNU | 124 ±2.48 |  | P | 14:10
| BP at ODO Cal |  |  |  | 765.81 | mmHg | [8 empty] | 765.81 | mmHg
| Optical DO % |  |  |  | 101.1 | % |  | 100.8 | % | ± 1% |  | P | 13:09 |  | 100.5 | % | P |  | 16:04
[6 empty] | DO gain |  | 1.03 | mg/L (table) | 0.87-1.25 |  | P |  |  | 100.2 | %
| Optical DO mg/L (ICV) | 9.63 |  |  | 17.459 | °C |  | 9.665 |  | ± 0.1 mg/L |  | P | 13:09 |  | 9.42 | mg/L (DL)
| Battery Voltage | [6 empty] | 5.27 | V | ~6V (EXO 2) |  | P |  |  | 18.958 | °C | P | 7:02 | 16:00
[6 empty] | Optional Parameters | [8 empty] | 4.92 | V | P
| Chlorophyll | 0 | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| Chlorophyll |  | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [5 empty] | °C
| Chlorophyll | 0 | µg/L |  |  | µg/L |  |  | µg/L | [6 empty] | µg/L | [5 empty] | Post-Deployment Notes
| Chlorophyll |  | µg/L |  |  | µg/L |  |  | µg/L | [6 empty] | µg/L
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [5 empty] | °C
| BGA | 0 | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| BGA |  | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [5 empty] | °C
Field Log
Fouling Presence | [9 empty] | Deployment |  |  | Retrieval |  |  |  | Field Notes:
A=algae, AM=amphipods, B=barnacles, BR=bryozoans, C=crabs, E=eggs, F=fish, FA=filamentous algae, G=gastropod, H=hydroids, I=isopod, MD=mud, S=sponges, SI=silt, SL=shell, SP=shrimp, T=tunicates, W=worm, O=other | [8 empty] | Date deploy/retrieval | 2022-11-14 |  |  | 2022-12-07 |  |  | Deployment:
[9 empty] | Time deploy/retrieval | 15:40 |  |  | 13:18 |  |  | Turbidity:.3.71
[9 empty] | First/last sample time | 15:45 |  |  | 13:15
| H= heavy, M= moderate, L= light | [7 empty] | Technician initials | AL |  |  | AL
Mount: MD-M
[9 empty] | Depth (m) | 1 |  |  | 2.3 |  | m
Sonde Body: MD-M | [8 empty] | Temp (°C) | 14.1 |  |  | 17.7 |  | °C
[9 empty] | DO % | 93.5 |  |  | 107.5 |  | %
Field guard: MD-M, AM-L, T-M | [8 empty] | DO mg/L | 7.81 |  |  | 8.4 |  | mg/L
All Sesnors: | [8 empty] | SpCond (mS/cm) | 51401 |  |  | 50457 |  | mS/cm | Retrieval:
C/T: MD-M, T-M | [8 empty] | Salinity (psu) | 33.75 |  |  | 33.14 |  | psu | Turbidity:2.77
pH: MD-M, T-M | [8 empty] | pH | 8.09 |  |  | 7.96
DO: MD-M, T-M | [11 empty] | Optional Samples
Turb: MD-M, T-M | [8 empty] | pHmV |  | mV |  |  |  | mV
Wiper: MD-M, T-M | [8 empty] | Conductivity |  | mS/cm |  |  |  | mS/cm
TAL-PE: MD-M, T-M
Other/notes:

## CPAH2_120722
Aquatic Preserve Continuous Water Quality Monitoring Calibration Log
|  | File Name | CPAH2_120722 | [5 empty] | .csv | [5 empty] | File Name | CPAH212072221E103416_120722_093000 |  |  |  |  | .bin
[7 empty] | Datasonde and Probe Identification Numbers | [12 empty] | Calibration Standards
|  | Sonde Code |  |  | Serial Number |  |  | Serial Number |  | Model Number | [5 empty] | Standard | Lot No. |  |  | Received | Opened | Exp. Date
| Datasonde | 1 |  |  | 21E103416 |  | pH | 21G105082 |  | 577602 |  |  |  |  | SpCond (mS/cm) | 20 | 4203H74 |  |  |  | 4/11/2022 | 3/2024
| Vented | Y/N |  |  | Model |  | ODO | 21G100611 |  | 599100-01 |  |  |  |  | SpCond (mS/cm) | 50 | 2908A58 |  |  |  |  | 7/2021
| Nickname | AH1 |  |  | EXO2 |  | Turb | 21G100846 |  | 599101-01 |  |  |  |  | pH | 4
| Preserve | CPAP |  |  |  |  | Cond | 21G101304 |  | 599827 |  |  |  |  | pH | 7 | 21M2S |  |  |  |  | 12/8/2023
| Station | CPAH2 |  |  |  |  | Chl | [7 empty] | pH | 10 | 22A1T |  |  |  |  | 1/3/2024
[7 empty] | Wiper | 21H105437 |  | 599090-01 |  |  |  |  | Turb (FNU) | 124 | 22D22130237 |  |  |  |  | 4/2023
|  |  |  | Wiper Replaced | Y |  |  |  | Batteries Replaced | N | [5 empty] | Technician(s) | Jon Brucker
|  |  |  | Wiper Parked in Garage | Y |  |  |  | Date of Calibration | 2022-12-01 | [5 empty] | Date of Post-Calibration
| Pre-deployment Comments/ Failed-calibration values: | ICV 20mS/cm high need spending authority for 50mS/cm standard.
| Pre-Deployment (CAL & ICV) | [12 empty] | Post-Deployment (CCV) | [5 empty] | Programming
|  | Standards |  |  | Before Cal |  |  | Calibrated |  | Specs |  | P/ F | Time |  | Post value |  | P/ F |  | Time | Interval |  | 15
| Temp | 18.83 | °C (Therm) |  | 18.443 | °C (DL) |  |  |  | ±0.2(w) +0.1°C |  | F | 12:23 |  |  | °C (Therm) |  |  |  | Duration |  | 365 days
| SpCond in air |  |  |  | 0 |  |  |  |  | ≤ 0.002 mS/cm |  | P |  |  |  | °C (DL) |  |  |  | Free memory % |  | 98%
| SpCond | 50 | mS/cm |  | 49.6612 | mS/cm |  | 50 | mS/cm | 50 ± 0.25 (n) |  |  | 12:57 | ✔ |  | mS/cm |  |  |  | Memory (days) |  | 24341.23
[6 empty] | Cell Constant |  | 0.48 |  | 0.419-0.519 (w) |  | P | [7 empty] | Start time |  | 9:30
| SpCond 20 (ICV) | 20 | mS/cm |  |  | ✔ |  |  | mS/cm | 20 ± 0.2 (w) | [9 empty] | Battery % |  | 97.8
| pH 7 | 7 |  |  | 7.05 |  |  | 7.03 |  | ± 0.1 |  |  | 13:16 | ✔ | [5 empty] | Set clock |  | Y
| pH 10 | 10 |  |  | 10.18 |  |  | 10.08 |  | ± 0.1 |  |  | 13:19 | ✔ | [5 empty] | End time
[5 empty] | pH 7 |  |  | -30.7 | mV | 0 ±50 mV |  | P |  |  |  | mV |  |  |  |  | Other/notes
[5 empty] | pH 10 |  |  | -211.1 | mV | -180 ±50 mV |  | P |  |  |  | mV
[5 empty] | pH 4 |  |  |  | mV | +180 ±50 mV
[5 empty] | pH Slope |  |  | 180.4 |  | > 155 mV |  | P
| pH 4 or 10 (ICV) | 10 |  |  |  | ✔ |  | 10.04 |  | ± 0.1 |  | F | 13:20
| BP at Depth Cal |  |  |  | 771.14 | mmHg | [9 empty] | mmHg
| Depth | 0.15 | Offset |  | 0.203 | m |  | 0.15 | m | ±0.004 m |  | P | 13:32 |  |  | m
| Turbidity | 0 | FNU |  | -0.54 |  |  | -0.01 | FNU | 0.00 ±0.3 |  |  | 13:37 | ✔ |  | FNU |  |  |  |  | General Maintenance
| Turbidity | 124 | FNU |  | 122.51 | FNU |  | 124 | FNU | 124 ±2.48 |  |  | 13:41 | ✔ |  | FNU
| Turbidity 124 (ICV) | 124 | FNU |  |  |  | ✔ | 124.06 | FNU | 124 ±2.48 |  | P | 13:42
| BP at ODO Cal |  |  |  | 755.25 | mmHg | [9 empty] | mmHg
| Optical DO % |  |  |  | 102.8 | % |  | 99.5 | % | ± 1% |  | P | 9:22 |  |  | %
[6 empty] | DO gain |  | 0.96 | mg/L (table) | 0.87-1.25 |  | P |  |  |  | %
| Optical DO mg/L (ICV) | 9 |  |  | 20.131 | °C |  | 9.092 |  | ± 0.1 mg/L |  | P | 9:24 |  |  | mg/L (DL)
| Battery Voltage | [6 empty] | 5.91 | V | ~6V (EXO 2) |  | P |  |  |  | °C |  | 7:02
[6 empty] | Optional Parameters | [9 empty] | V
| Chlorophyll | 0 | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| Chlorophyll |  | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [5 empty] | °C
| Chlorophyll | 0 | µg/L |  |  | µg/L |  |  | µg/L | [6 empty] | µg/L |  |  |  |  | Post-Deployment Notes
| Chlorophyll |  | µg/L |  |  | µg/L |  |  | µg/L | [6 empty] | µg/L
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [5 empty] | °C
| BGA | 0 | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| BGA |  | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [5 empty] | °C
Field Log
Fouling Presence | [9 empty] | Deployment |  |  | Retrieval |  |  |  | Field Notes:
A=algae, AM=amphipods, B=barnacles, BR=bryozoans, C=crabs, E=eggs, F=fish, FA=filamentous algae, G=gastropod, H=hydroids, I=isopod, MD=mud, S=sponges, SI=silt, SL=shell, SP=shrimp, T=tunicates, W=worm, O=other | [8 empty] | Date deploy/retrieval | 2022-12-07 | [5 empty] | Deployment:
[9 empty] | Time deploy/retrieval | 13:21 | [5 empty] | Turbidity:2.77
[9 empty] | First/last sample time | 13:30
| H= heavy, M= moderate, L= light | [7 empty] | Technician initials | AL
Mount:
[9 empty] | Depth (m) | 2.3 |  |  |  |  | m
Sonde Body: | [8 empty] | Temp (°C) | 17.7 |  |  |  |  | °C
[9 empty] | DO % | 107.5 |  |  |  |  | %
Field guard: | [8 empty] | DO mg/L | 8.4 |  |  |  |  | mg/L
All Sesnors: | [8 empty] | SpCond (mS/cm) | 50.457 |  |  |  |  | mS/cm | Retrieval:
C/T: | [8 empty] | Salinity (psu) | 33.14 |  |  |  |  | psu
pH: | [8 empty] | pH | 7.96
DO: | [11 empty] | Optional Samples
Turb: | [8 empty] | pHmV |  | mV |  |  |  | mV
Wiper: | [8 empty] | Conductivity |  | mS/cm |  |  |  | mS/cm
TAL-PE:
Other/notes: