## Instructions
Instructions
● | Do not move, edit, swap, or shift any columns, rows or cells in the calibration log.
● | In the Drop-downs tab, edit the white columns to match your station, probe, and sonde information. Gray columns should not be edited, only selected in the log drop downs.
● | Within the log, only in white boxes labeled notes/other, you may type in extra items you wish the calibration log to prompt (see example).
● | If a calibration or ICV fails, place a note with the failed value in the pre-deployment comments section. Passing recalibrated value should go in the actual log.
● | General maintenance section is for any maintenance (in-field maintenance, probe repair date/ reason, etc.).
● | Parameters should be calibrated in the order of the calibration log from top to bottom; this matters especially for temperature and SpCond.
● | Your NIST device should be at least as accurate as your sensor. The accuracy of the NIST device must be included in your temperature accuracy determination, meaning that the temperature probe specification should be added to the accuracy of your thermometer.
● | Submit calibration log in .xls format named by AP station code and year, for example, BBSCH_2021_CalLog
● | Tabs at the bottom should be named by the station code and deployment date, for example, BBSCH_010222. There should be approximately 12 tabs in each yearly calibration log file.
Other information
● | Optional parameters will not be uploaded to the AP Data Portal at this time but will remain in the dataset for possible future upload.
● | Check marks represent DEP SOP requirement for ICV (initial calibration verification) and CCV (continuing calibration verification) steps.
● | The DL next to a parameter's units indicates that the reading comes from the datalogger's probe.

## FT1500
Solubility of Oxygen in Water at Atmospheric Pressure
Temperature | Oxygen Solubility | Use this table to DEP verify that DO mg/L are reading within range based on the temperature reading.
°C | mg/L | DEP SOP FT 1500
0 | 14.621
1 | 14.216
2 | 13.829
3 | 13.46
4 | 13.107
5 | 12.77
6 | 12.447
7 | 12.139
8 | 11.843
9 | 11.559
10 | 11.288
11 | 11.027
12 | 10.777
13 | 10.537
14 | 10.306
15 | 10.084
16 | 9.87
17 | 9.665
18 | 9.467
19 | 9.276
20 | 9.092
21 | 8.915
22 | 8.743
23 | 8.578
24 | 8.418
25 | 8.263
26 | 8.113
27 | 7.968
28 | 7.827
29 | 7.691
30 | 7.559
31 | 7.43
32 | 7.305
33 | 7.183
34 | 7.065
35 | 9.65
36 | 6.837
37 | 6.727
38 | 6.62
39 | 6.515
40 | 6.412
41 | 6.312
42 | 6.213
43 | 6.116
44 | 6.021
45 | 5.927
46 | 8.835
47 | 5.744
48 | 5.654
49 | 5.565
50 | 5.477

## Drop-downs
AP | Station | Datasonde | Sonde No. | pH Serial No. | ODO Serial No. | Turb Serial No. | Cond Serial No. | Chl Serial No. | Wiper Serial No. | Y/N | Pass/Fail | pH Model | ODO Model | Turb Model | Cond Model | Chl Model | Wiper Model | Sonde Model | Lot No. | Technician | Cell Constant |  | Battery Voltage | DO Gain |  | Sp Cond (mS/cm) |  | Temperature |  | DO |  | Depth |  | Turbidity
CPAP | CPAH2 | 1 | 21E103418 | 21G105084 | 21G100610 | 21G100848 | 21G101301 |  | 21H105435 | Y | P | 577602 | 599100-01 | 599101-01 | 599827 | 599103-01 | 599090-01 | EXO2 | 2712749 | Jon Brucker | EXO Wiped CT | 0.419-0.519 (w) | ~6V (EXO 2) | EXO Optical | 0.87-1.25 | Non-wiped (10 std.) | 10 ± 0.05 (n) | Non-wiped (-5 to 35°C) | ±0.01(n) +0.1°C | Optical (0-200% air sat.) | ± 1% (op) | EXO (10m) | ±0.004 m | 0.00 ±0.3
| CPFS | 2 | 21E103415 | 21G105934 | 21G100609 | 21G100849 | 22F100242 |  | 21H105436 | N | F | 6151 | 6150 | 6136 | 6560 |  |  | EXO3 | 2906F57 | Michael Palandri | EXO Non-Wiped CT | 5.05-5.95 (n) | ~3V (EXO 3) | 6600 | 0.7 - 1.40 | Non-wiped (20 std.) | 20 ± 0.01 (n) | Wiped sensor | ±0.2(w) +0.1°C |  |  | 6600 (10m) | ±0.06 ft (0.018 m) | 124 ±2.48
| CPWD |  | 21E103417 | 21G105935 | 21G100608 | 21G100845 | 21G101300 |  | 21H105434 | Auto | [7 empty] | 6600 | 106779 | [6 empty] | Non-wiped (50 std.) | 50 ± 0.25 (n)
|  |  | 21E103418 | 21G105083 | 21G100607 | 21G100847 | 21G101303 |  | 21H105438 | [9 empty] | 21H2S | [6 empty] | Wiped Sensor (10 std.) | 10 ± 0.1 (w)
|  | 145153 | 21E103416 | 21G105082 | 21G100611 | 21G100846 | 21G101304 |  | 21H105497 | [9 empty] | 21G3T | [6 empty] | Wiped Sensor (20 std.) | 20 ± 0.2 (w)
|  | 140387 | 10B100491 | 19E | 13C100357 | 15B104385 | 16G100119 | [11 empty] | 21G21280150 | [6 empty] | Wiped Sensor (50 std.) | 50 ± 0.5 (w)
|  | 145459 | 13C100686 | 20C | 15L103284 | 12C100448 | 13C100687
|  | 120803 | 02A0876AI | 21A | 15L103286 | 12C100447 | 15G100321
|  | 145154 | 12H102535 | 19F | 12H101596 | 14C103217 | 16G100618
|  | 145458 | 13C100687 |  | 14D100738 | 11H101089 | 14J100146
[6 empty] | 16C102423 | 13C100054
[6 empty] | 14D102113 | 15F100557
Note on AA1: Howe, Samantha: 0 to 100: ± 0.5% of reading or 0.001 mS/cm 100 to 200: ± 1% of reading Wiped Sensor: ±1% of the reading or 0.002 mS/cm (whichever is greater)
Note on AD2: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: +0.1 represents the accuracy of your thermometer. The total accuracy should include the addition of your NIST device. Your NIST should be at least this accurate.

## Example
Aquatic Preserve Continuous Water Quality Monitoring Calibration Log
|  | File Name | BBSCH_110221 | [5 empty] | .csv | [5 empty] | File Name | BBSCH19E101881_110221_071500 |  |  |  |  | .bin
[7 empty] | Datasonde and Probe Identification Numbers | [12 empty] | Calibration Standards
|  | Sonde Code |  |  | Serial Number |  |  | Serial Number |  | Model Number | [5 empty] | Standard | Lot No. |  |  | Received | Opened | Exp. Date
| Datasonde | 156019 |  |  | 19E101880 |  | pH | 19K101465 |  | 577602 |  |  |  |  | SpCond (mS/cm) | 20 | 2008A95 |  |  | 2021-07-12 | 08/23/2021 | 07/2022
| Vented | Y/N |  |  | Model |  | ODO | 20A103173 |  | 599100-01 |  |  |  |  | SpCond (mS/cm) | 50 | 2105829 |  |  | 2021-07-01 | 09/16/2021 | 04/2023
| Nickname | CH |  |  | EXO2 |  | Turb | 20A103291 |  | 599101-01 |  |  |  |  | pH | 4 | 2004922 |  |  | 2021-07-12 | 08/27/2021 | 03/2022
| Preserve | BBSAP |  |  |  |  | Cond | 20A103494 |  | 599827 |  |  |  |  | pH | 7 | 2103C60 |  |  | 2021-05-03 | 08/04/2021 | 02/2023
| Station | BBSCH (Chassahowitzka) |  |  |  |  | Chl | 19J104091 |  | 599103-01 |  |  |  |  | pH | 10 | 4102C02 |  |  | 2021-04-30 | 10/06/2021 | 07/2022
[7 empty] | Wiper | 19J105081 |  | 599090-01 |  |  |  |  | Turb (FNU) | 124 | 21H21280150 |  |  | 2021-08-24 | 08/27/2021 | 08/2022
|  |  |  | Wiper Replaced | N |  |  |  | Batteries Replaced | N | [5 empty] | Technician(s) | Trisha Green
|  |  |  | Wiper Parked in Garage | Y |  |  |  | Date of Calibration | 2021-11-02 | [5 empty] | Date of Post-Calibration | 2021-11-23
| Pre-deployment Comments/ Failed-calibration values: | If a calibration QC score or ICV fails, place a note with the failed value in the pre-deployment comments section. Passing recalibrated value should go in the actual log.
| Pre-Deployment (CAL & ICV) | [12 empty] | Post-Deployment (CCV) | [5 empty] | Programming
|  | Standards |  |  | Before Cal |  |  | Calibrated |  | Specs |  | P/ F | Time |  | Post value |  | P/ F | Time |  | Interval |  | 15
| Temp | 21.75 | °C (Therm) |  | 21.693 | °C (DL) |  |  |  | ±0.2(w) +0.1°C |  | P | 7:10 |  | 20.5 | °C (Therm) | P | 12:56 |  | Duration |  | 365 days
| SpCond in air |  |  |  | 0 |  |  |  |  | ≤ 0.002 mS/cm |  | P |  |  | 20.46 | °C (DL) |  |  |  | Free memory % |  | 99.55
| SpCond | 50 | mS/cm |  | 49.767 | mS/cm |  | 50 | mS/cm | 50 ± 0.25 (n) |  |  | 7:08 | ✔ | 50.02 | mS/cm | P | 13:00 |  | Memory (days) |  | 24432.81
[6 empty] | Cell Constant |  | 0.47 |  | 0.419-0.519 (w) |  | P | [7 empty] | Start time |  | 9:30
| SpCond 20 (ICV) | 20 | mS/cm |  |  | ✔ |  | 20.112 | mS/cm | 20 ± 0.2 (w) |  | P | 7:14 | [6 empty] | Battery % |  | 65
| pH 7 | 7 |  |  | 7.05 |  |  | 7.02 |  | ± 0.2 |  |  | 7:11 | ✔ | 7.1 |  | P | 13:04 |  | Set clock |  | Y/N
| pH 10 | 10 |  |  | 10.08 |  |  | 10.05 |  | ± 0.2 |  |  | 7:11 | ✔ | 10.06 |  | P | 13:05 |  | End time |  | 10:00
[5 empty] | pH 7 |  |  | -30.7 | mV | 0 ±50 mV |  | P |  |  | -35.5 | mV | P |  |  |  | Other/notes
[5 empty] | pH 10 |  |  | -204.1 | mV | -180 ±50 mV |  | P |  |  | -203.9 | mV | P |  |  | Change locks |  | Y
[5 empty] | pH 4 |  |  | 148.8 | mV | +180 ±50 mV |  | P | [7 empty] | Check for DST |  | Y
[5 empty] | pH Slope |  |  | 173.4 |  | > 155 mV |  | P |  |  | 168.4 |  | P
| pH 4 or 10 (ICV) | 4 |  |  |  | ✔ |  | 3.9 |  | ± 0.1 |  | P | 7:16
| BP at Depth Cal |  |  |  | 765.556 | mmHg | [8 empty] | 765.56 | mmHg
| Depth | 0.075534 | Offset |  | 0.062 | m |  | 0.076 | m | ±0.004 m |  | P | 7:24 |  | 0.072 | m | P | 13:10
| Turbidity | 0 | FNU |  | -0.1 |  |  | 0 | FNU | 0.00 ±0.3 |  |  | 7:27 | ✔ | 0.04 | FNU | P | 13:15 |  |  | General Maintenance
| Turbidity | 124 | FNU |  | 123.24 | FNU |  | 124 | FNU | 124 ±2.48 |  |  | 7:31 | ✔ | 117.72 | FNU | P | 13:17 | This section is for any maintenance (in-field maintenance, probe repair etc.). Please include dates, probe numbers, and any information that is needed to track the incident.
| Turbidity 124 (ICV) | 124 | FNU |  |  |  | ✔ | 123.89 | FNU | 124 ±2.48 |  | P | 7:36
| BP at ODO Cal |  |  |  | 765.556 | mmHg | [8 empty] | 765.556 | mmHg
| Optical DO % |  |  |  | 101.7 | % |  | 100.9 | % | ± 1% |  | P | 7:38 |  | 100.5 | % | P | 13:22
[6 empty] | DO gain |  | 1.02 | mg/L (table) | 0.87-1.25 |  | P |  |  | 100.5 | %
| Optical DO mg/L (ICV) | 8.85 | mg/L (DL) |  | 21.75 | °C |  | 8.743 |  | ± 0.1 mg/L |  | P | 7:43 |  | 9.05 | mg/L (DL)
| Battery Voltage | [6 empty] | 3.2 | V | ~3V (EXO 3) |  | P |  |  | 20.5 | °C | P | 13:29
[6 empty] | Optional Parameters | [8 empty] | 2.24 | V | P
| Chlorophyll | 0 | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| Chlorophyll |  | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [5 empty] | °C
| Chlorophyll | 0 | µg/L |  |  | µg/L |  |  | µg/L | [6 empty] | µg/L |  |  |  |  | Post-Deployment Notes
| Chlorophyll |  | µg/L |  |  | µg/L |  |  | µg/L | [6 empty] | µg/L |  |  | Turbidity post-cal kind of low. Watch this sensor. May need to be replaced soon or just needed to be cleaned well.
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [5 empty] | °C
| BGA | 0 | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| BGA |  | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [5 empty] | °C
Field Log
Fouling Presence | [9 empty] | Deployment |  |  | Retrieval |  |  |  | Field Notes:
A=algae, AM=amphipods, B=barnacles, BR=bryozoans, C=crabs, E=eggs, F=fish, FA=filamentous algae, G=gastropod, H=hydroids, I=isopod, MD=mud, S=sponges, SI=silt, SL=shell, SP=shrimp, T=tunicates, W=worm, O=other | [8 empty] | Date deploy/retrieval | 2021-11-02 |  |  | 2021-11-23 |  |  | Deployment:
[9 empty] | Time deploy/retrieval | 8:52 |  |  | 10:17 |  |  | Low tide, calm, light winds.
[9 empty] | First/last sample time | 9:00 |  |  | 10:00
| H= heavy, M= moderate, L= light | [7 empty] | Technician initials | TG, KS |  |  | TG, KS, TJ
Mount: 2 Fish; A/H, B/M, Si/M | [9 empty] | Discrete Samples (Required)
[9 empty] | Depth (m) | 0.0634 | m |  | 0.0787 |  | m
Sonde Body: Fish in guard; E/H, Si/M, T/L, W/L | [8 empty] | Temp (°C) | 19.3 | °C |  | 17.5 |  | °C
[9 empty] | DO % | 103.9 | % |  | 94.2 |  | %
Field guard: | [8 empty] | DO mg/L | 8.21 | mg/L |  | 7.94 |  | mg/L
All Sesnors: Si/M, EH, W/L | [8 empty] | SpCond (mS/cm) | 41.065 | mS/cm |  | 33.298 |  | mS/cm | Retrieval:
C/T: G/L | [8 empty] | Salinity (psu) | 26.35 | psu |  | 20.91 |  | psu | Low tide, cold, windy, choppy, brown water, clear skies.
pH: | [8 empty] | pH | 8.2 |  |  | 8.03
DO: | [11 empty] | Optional Samples
Turb: | [8 empty] | pHmV | -76.9 | mV |  | -74.1 |  | mV
Wiper: Bad; fell off | [8 empty] | Conductivity |  | mS/cm |  |  |  | mS/cm
TAL-PE:
Other/Notes:

## APCODE_MMDDYY
Aquatic Preserve Continuous Water Quality Monitoring Calibration Log
|  | File Name | [6 empty] | .csv | [5 empty] | File Name |  |  |  |  | .bin
[7 empty] | Datasonde and Probe Identification Numbers | [11 empty] | Calibration Standards
|  | Sonde Code |  |  | Serial Number |  |  | Serial Number |  | Model Number | [5 empty] | Standard | Lot No. |  | Received | Opened | Exp. Date
| Datasonde | Datasonde |  |  | Sonde No. |  | pH | pH Serial No. |  | pH Model |  |  |  |  | SpCond (mS/cm) | 20
| Vented | Y/N |  |  | Model |  | ODO | ODO Serial No. |  | ODO Model |  |  |  |  | SpCond (mS/cm) | 50
| Nickname |  |  |  | Sonde Model |  | Turb | Turb Serial No. |  | Turb Model |  |  |  |  | pH | 4
| Preserve | AP |  |  |  |  | Cond | Cond Serial No. |  | Cond Model |  |  |  |  | pH | 7
| Station | Station |  |  |  |  | Chl | Chl Serial No. |  | Chl Model |  |  |  |  | pH | 10
[7 empty] | Wiper | Wiper Serial No. |  | Wiper Model |  |  |  |  | Turb (FNU) | 124
|  |  |  | Wiper Replaced |  |  |  |  | Batteries Replaced | [6 empty] | Technician(s) | Technician
|  |  |  | Wiper Parked in Garage |  |  |  |  | Date of Calibration | [6 empty] | Date of Post-Calibration
Pre-deployment Comments/ Failed-calibration values:
| Pre-Deployment (CAL & ICV) | [12 empty] | Post-Deployment (CCV) |  |  |  |  | Programming
|  | Standards |  |  | Before Cal |  |  | Calibrated |  | Specs |  | P/ F | Time |  | Post value |  | P/ F | Time | Interval |  | 15
| Temp |  | °C (Therm) |  |  | °C (DL) |  |  |  | ±0.2(w) +0.1°C | [5 empty] | °C (Therm) |  |  | Duration |  | 365 days
| SpCond in air | [8 empty] | ≤ 0.002 mS/cm | [5 empty] | °C (DL) |  |  | Free memory %
| SpCond | 50 | mS/cm |  |  | mS/cm |  |  | mS/cm | 50 ± 0.25 (n) |  |  |  | ✔ |  | mS/cm |  |  | Memory (days)
[6 empty] | Cell Constant |  |  |  | 0.419-0.519 (w) | [8 empty] | Start time
| SpCond 20 (ICV) | 20 | mS/cm |  |  | ✔ |  |  | mS/cm | 20 ± 0.2 (w) | [8 empty] | Battery %
| pH 7 | 7 | [7 empty] | ± 0.1 |  |  |  | ✔ |  |  |  |  | Set clock |  | Y/N
| pH 10 | 10 | [7 empty] | ± 0.1 |  |  |  | ✔ |  |  |  |  | End time
[5 empty] | pH 7 |  |  |  | mV | 0 ±50 mV | [5 empty] | mV |  |  |  | Other/notes
[5 empty] | pH 10 |  |  |  | mV | -180 ±50 mV | [5 empty] | mV
[5 empty] | pH 4 |  |  |  | mV | +180 ±50 mV
[5 empty] | pH Slope |  |  | 0 |  | > 155 mV |  |  |  |  | 0
| pH 4 or 10 (ICV) | 10 |  |  |  | ✔ |  |  |  | ± 0.1
| BP at Depth Cal |  |  |  |  | mmHg | [9 empty] | mmHg
| Depth |  | Offset |  |  | m |  |  | m | ±0.004 m | [5 empty] | m
| Turbidity | 0 | FNU | [5 empty] | FNU | 0.00 ±0.3 |  |  |  | ✔ |  | FNU |  |  |  | General Maintenance
| Turbidity | 124 | FNU |  |  | FNU |  |  | FNU | 124 ±2.48 |  |  |  | ✔ |  | FNU
| Turbidity 124 (ICV) | 124 | FNU |  |  |  | ✔ |  | FNU | 124 ±2.48
| BP at ODO Cal |  |  |  |  | mmHg | [9 empty] | mmHg
| Optical DO % |  |  |  |  | % |  |  | % | ± 1% | [5 empty] | %
[6 empty] | DO gain |  |  | mg/L (table) | 0.87-1.25 | [5 empty] | %
| Optical DO mg/L (ICV) |  | mg/L (DL) |  |  | °C |  |  |  | ± 0.1 mg/L | [5 empty] | mg/L (DL)
| Battery Voltage | [7 empty] | V | ~6V (EXO 2) | [5 empty] | °C
[6 empty] | Optional Parameters | [9 empty] | V
| Chlorophyll | 0 | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| Chlorophyll |  | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [5 empty] | °C
| Chlorophyll | 0 | µg/L |  |  | µg/L |  |  | µg/L | [6 empty] | µg/L |  |  |  | Post-Deployment Notes
| Chlorophyll |  | µg/L |  |  | µg/L |  |  | µg/L | [6 empty] | µg/L
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [5 empty] | °C
| BGA | 0 | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| BGA |  | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [5 empty] | °C
Field Log
Fouling Presence | [9 empty] | Deployment |  |  | Retrieval |  |  |  | Field Notes:
A=algae, AM=amphipods, B=barnacles, BR=bryozoans, C=crabs, E=eggs, F=fish, FA=filamentous algae, G=gastropod, H=hydroids, I=isopod, MD=mud, S=sponges, SI=silt, SL=shell, SP=shrimp, T=tunicates, W=worm, O=other | [8 empty] | Date deploy/retrieval | [6 empty] | Deployment:
Time deploy/retrieval
First/last sample time
| H= heavy, M= moderate, L= light | [7 empty] | Technician initials
Mount: | [9 empty] | Discrete Samples (Required)
[9 empty] | Depth (m) |  | m |  |  |  | m
Sonde Body: | [8 empty] | Temp (°C) |  | °C |  |  |  | °C
[9 empty] | DO % |  | % |  |  |  | %
Field guard: | [8 empty] | DO mg/L |  | mg/L |  |  |  | mg/L
All Sesnors: | [8 empty] | SpCond (mS/cm) |  | mS/cm |  |  |  | mS/cm | Retrieval:
C/T: | [8 empty] | Salinity (psu) |  | psu |  |  |  | psu
pH: | [8 empty] | pH
DO: | [11 empty] | Optional Samples
Turb: | [8 empty] | pHmV |  | mV |  |  |  | mV
Wiper: | [8 empty] | Conductivity |  | mS/cm |  |  |  | mS/cm
TAL-PE:
Other/notes:

## CPFS_011422
Aquatic Preserve Continuous Water Quality Monitoring Calibration Log
|  | File Name | CPFS_011422 | [5 empty] | .csv | [5 empty] | File Name | CPFS_011422.DAT |  |  |  | .bin
[7 empty] | Datasonde and Probe Identification Numbers | [11 empty] | Calibration Standards
|  | Sonde Code |  |  | Serial Number |  |  | Serial Number |  | Model Number | [5 empty] | Standard | Lot No. |  | Received | Opened | Exp. Date
| Datasonde | 145459 |  |  | 13C100687 |  | pH | 19F |  | 6151 |  |  |  |  | SpCond (mS/cm) | 20
| Vented | N |  |  | Model |  | ODO | 15L103284 |  | 6150 |  |  |  |  | SpCond (mS/cm) | 50 | 2906F57 |  |  |  | 7/2021
| Nickname | FS |  |  | 6600 |  | Turb | 12C100448 |  | 6136 |  |  |  |  | pH | 4 | 160779 |  |  |  | 7/13/23
| Preserve | CPAP |  |  |  |  | Cond | 13C100687 |  | 6560 |  |  |  |  | pH | 7 | 21H2S |  |  |  | 8/11/23
| Station | CPFS |  |  |  |  | Chl | Chl Serial No. |  | Chl Model |  |  |  |  | pH | 10 | 21G3T |  |  |  | 7/13/23
[7 empty] | Wiper | Wiper Serial No. |  | Wiper Model |  |  |  |  | Turb (FNU) | 124 | 21G21280150 |  |  | 11/1/21 | 7/2022
|  |  |  | Wiper Replaced | N |  |  |  | Batteries Replaced | N | [5 empty] | Technician(s) | Jon Brucker
|  |  |  | Wiper Parked in Garage | Y |  |  |  | Date of Calibration | 2022-01-13 | [5 empty] | Date of Post-Calibration | 3/7/2022 but attempted 2/21/22
| Pre-deployment Comments/ Failed-calibration values: | Sonde would not connect properly when brought back to the lab. Got the # sign but program would not respond. Changed batteries and same result on 2/21. Connected sonde back up and was able to connect, post cal, and download 3/7. Sonde kept running 2/21 - 3/7. Delete tails!
| Pre-Deployment (CAL & ICV) | [12 empty] | Post-Deployment (CCV) |  |  |  |  | Programming
|  | Standards |  |  | Before Cal |  |  | Calibrated |  | Specs |  | P/ F | Time |  | Post value |  | P/ F | Time | Interval |  | 15
| Temp | 20.2 | °C (Therm) |  | 20.4 | °C (DL) |  |  |  | ±0.2(w) +0.1°C |  | P |  |  | 20.2 | °C (Therm) | P |  | Duration |  | 365 days
| SpCond in air |  |  |  | 0 |  |  |  |  | ≤ 0.002 mS/cm |  | P |  |  | 20.4 | °C (DL) |  |  | Free memory % |  | N/A
| SpCond | 50 | mS/cm |  | 49.12 | mS/cm |  | 50 | mS/cm | 50 ± 0.25 (n) |  |  |  | ✔ | 50.07 | mS/cm | P |  | Memory (days) |  | 185
[6 empty] | Cell Constant |  | 5.0439 |  | 0.419-0.519 (w) |  | P | [6 empty] | Start time |  | 10:14
| SpCond 20 (ICV) | 20 | mS/cm |  |  | ✔ |  | N/A | mS/cm | 20 ± 0.2 (w) | [8 empty] | Battery % |  | N/A
| pH 7 | 7 |  |  | 7.06 |  |  | 7 |  | ± 0.2 |  |  |  | ✔ | 6.88 |  | P |  | Set clock |  | Y
| pH 10 | 10 |  |  | 9.93 |  |  | 10 |  | ± 0.2 |  |  |  | ✔ | N/A |  |  |  | End time |  | 10:45
[5 empty] | pH 7 |  |  | -5.9 | mV | 0 ±50 mV |  | P |  |  | 0.8 | mV | P |  |  | Other/notes
[5 empty] | pH 10 |  |  | -173.3 | mV | -180 ±50 mV |  | P |  |  | N/A | mV
[5 empty] | pH 4 |  |  |  | mV | +180 ±50 mV |  | P
[5 empty] | pH Slope |  |  | 167.4 |  | > 155 mV |  | P |  |  | #VALUE!
| pH 4 or 10 (ICV) | 4 |  |  |  | ✔ |  | N/A |  | ± 0.1
| BP at Depth Cal |  |  |  | 764.54 | mmHg | [8 empty] | 767.842 | mmHg
| Depth | 0.062 | Offset |  | 0.048 | m |  | 0.062 | m | ±0.004 m |  | P |  |  | 0.071 | m
| Turbidity | 0 | FNU |  | 1.4 |  |  | 0 | FNU | 0.00 ±0.3 |  |  |  | ✔ | -2.2 | FNU | P |  |  | General Maintenance
| Turbidity | 126 | FNU |  | 122 | FNU |  | 126 | FNU | 124 ±2.48 |  |  |  | ✔ | N/A | FNU
| Turbidity 124 (ICV) | 126 | FNU |  |  |  | ✔ | N/A | FNU | 124 ±2.48
| BP at ODO Cal |  |  |  | 764.54 | mmHg | [8 empty] | 766.572 | mmHg
| Optical DO % |  |  |  | 97.7 | % |  | 100.6 | % | ± 1% |  | P |  |  | 100.7 | % | P
[6 empty] | DO gain |  | 1.0522 | mg/L (table) | 0.87-1.25 |  | P |  |  | 100.7 | %
| Optical DO mg/L (ICV) | N/A | mg/L (DL) |  | N/A | °C |  | n/a |  | ± 0.1 mg/L |  |  |  |  | N/A | mg/L (DL)
| Battery Voltage | [6 empty] | 10.8 | V | ~6V (EXO 2) |  | P |  |  | n/A | °C
[6 empty] | Optional Parameters | [8 empty] | 11.8 | V
| Chlorophyll | 0 | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| Chlorophyll |  | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [5 empty] | °C
| Chlorophyll | 0 | µg/L |  |  | µg/L |  |  | µg/L | [6 empty] | µg/L |  |  |  | Post-Deployment Notes
| Chlorophyll |  | µg/L |  |  | µg/L |  |  | µg/L | [6 empty] | µg/L
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [5 empty] | °C
| BGA | 0 | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| BGA |  | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [5 empty] | °C
Field Log
Fouling Presence | [9 empty] | Deployment |  |  | Retrieval |  |  |  | Field Notes:
A=algae, AM=amphipods, B=barnacles, BR=bryozoans, C=crabs, E=eggs, F=fish, FA=filamentous algae, G=gastropod, H=hydroids, I=isopod, MD=mud, S=sponges, SI=silt, SL=shell, SP=shrimp, T=tunicates, W=worm, O=other | [8 empty] | Date deploy/retrieval | 2022-01-14 |  |  | 2022-02-18 |  |  | Deployment:
[9 empty] | Time deploy/retrieval | 11:27 |  |  | 12:40 |  |  | Staff did not collect in field measurements. Sunny, low tide rising, calm
[9 empty] | First/last sample time | 11:30 |  |  | 12:30
| H= heavy, M= moderate, L= light | [7 empty] | Technician initials | JB |  |  | JB
Mount: N | [9 empty] | Discrete Samples (Required)
[9 empty] | Depth (m) |  | m |  |  |  | m
Sonde Body: N | [8 empty] | Temp (°C) |  | °C |  |  |  | °C
[9 empty] | DO % |  | % |  |  |  | %
Field guard: N | [8 empty] | DO mg/L |  | mg/L |  |  |  | mg/L
All Sesnors: | [8 empty] | SpCond (mS/cm) |  | mS/cm |  |  |  | mS/cm | Retrieval:
C/T: N | [8 empty] | Salinity (psu) |  | psu |  |  |  | psu | Staff did not collect in field measurements. Calm, mid tide rising, NW wind, some clouds
pH: N | [8 empty] | pH
DO: N | [11 empty] | Optional Samples
Turb: N | [8 empty] | pHmV |  | mV |  |  |  | mV
Wiper: N | [8 empty] | Conductivity |  | mS/cm |  |  |  | mS/cm
TAL-PE:
Other/notes:

## CPFS_021822
Aquatic Preserve Continuous Water Quality Monitoring Calibration Log
|  | File Name | CPFS_021822 | [5 empty] | .csv | [5 empty] | File Name | CPFS_021822.DAT |  |  |  |  | .bin
[7 empty] | Datasonde and Probe Identification Numbers | [12 empty] | Calibration Standards
|  | Sonde Code |  |  | Serial Number |  |  | Serial Number |  | Model Number | [5 empty] | Standard | Lot No. |  |  | Received | Opened | Exp. Date
| Datasonde | 145154 |  |  | 12H102535 |  | pH | 20C |  | 6151 |  |  |  |  | SpCond (mS/cm) | 20
| Vented | N |  |  | Model |  | ODO | 14D100738 |  | 6150 |  |  |  |  | SpCond (mS/cm) | 50 | 2906F57 |  |  |  |  | 7/2021
| Nickname | FS |  |  | 6600 |  | Turb | 14C103217 |  | 6136 |  |  |  |  | pH | 4 | 160779 |  |  |  |  | 7/13/23
| Preserve | CPAP |  |  |  |  | Cond | 16G100618 |  | 6560 |  |  |  |  | pH | 7 | 21H2S |  |  |  |  | 8/11/23
| Station | CPFS |  |  |  |  | Chl | Chl Serial No. |  | Chl Model |  |  |  |  | pH | 10 | 21G3T |  |  |  |  | 7/13/23
[7 empty] | Wiper | Wiper Serial No. |  | Wiper Model |  |  |  |  | Turb (FNU) | 124 | 21G21280150 |  |  |  | 11/1/21 | 7/2022
|  |  |  | Wiper Replaced | N |  |  |  | Batteries Replaced | Y | [5 empty] | Technician(s) | Jon Brucker
|  |  |  | Wiper Parked in Garage | Y |  |  |  | Date of Calibration | 2022-02-08 | [5 empty] | Date of Post-Calibration | 2022-03-18
| Pre-deployment Comments/ Failed-calibration values: | New batteries 2/8, New DO probe 2/8
| Pre-Deployment (CAL & ICV) | [12 empty] | Post-Deployment (CCV) | [5 empty] | Programming
|  | Standards |  |  | Before Cal |  |  | Calibrated |  | Specs |  | P/ F | Time |  | Post value |  | P/ F | Time |  | Interval |  | 15
| Temp | 20.2 | °C (Therm) |  | 20.4 | °C (DL) |  |  |  | ±0.2(w) +0.1°C |  | P | 11:00 |  |  | °C (Therm) |  |  |  | Duration |  | 365 days
| SpCond in air |  |  |  | 0 |  |  |  |  | ≤ 0.002 mS/cm |  | P |  |  |  | °C (DL) |  |  |  | Free memory % |  | N/A
| SpCond | 50 | mS/cm |  | 49.55 | mS/cm |  | 50 | mS/cm | 50 ± 0.25 (n) |  |  | 11:02 | ✔ | 50.311 | mS/cm | P |  |  | Memory (days)
[6 empty] | Cell Constant |  | 5.1805 |  | 0.419-0.519 (w) |  | P | [7 empty] | Start time |  | 10:44
| SpCond 20 (ICV) | 20 | mS/cm |  |  | ✔ |  | N/A | mS/cm | 20 ± 0.2 (w) | [9 empty] | Battery % |  | 62.8 days
| pH 7 | 7 |  |  | 6.61 |  |  | 7 |  | ± 0.1 |  |  | 11:05 | ✔ | 7.18 |  | F |  |  | Set clock |  | Y
| pH 10 | 10 |  |  | 10.28 |  |  | 10 |  | ± 0.1 |  |  | 11:08 | ✔ | [5 empty] | End time |  | 10:15
[5 empty] | pH 7 |  |  | 51.9 | mV | 0 ±50 mV |  | P |  |  | 42.2 | mV | P |  |  |  | Other/notes
[5 empty] | pH 10 |  |  | -136 | mV | -180 ±50 mV |  | P |  |  |  | mV
[5 empty] | pH 4 |  |  | N/A | mV | +180 ±50 mV |  | P
[5 empty] | pH Slope |  |  | y |  | > 155 mV |  | P |  |  | 42.2
| pH 4 or 10 (ICV) | 4 |  |  |  | ✔ |  | N/A |  | ± 0.1
| BP at Depth Cal |  |  |  | 766.572 | mmHg | [8 empty] | 761.746 | mmHg
| Depth | 0.095 | Offset |  | 0.074 | m |  | 0.095 | m | ±0.004 m |  | P | 11:15 |  | 0.032 | m | F
| Turbidity | 0 | FNU |  | 2.7 |  |  | 0 | FNU | 0.00 ±0.3 |  |  | 11:21 | ✔ | -2.8 | FNU | F |  |  |  | General Maintenance
| Turbidity | 126 | FNU |  | 119.6 | FNU |  | 126 | FNU | 124 ±2.48 |  |  | 11:23 | ✔ |  | FNU |  |  | Greased lock.
| Turbidity 124 (ICV) | 126 | FNU |  |  |  | ✔ | N/A | FNU | 124 ±2.48
| BP at ODO Cal |  |  |  | 766.572 | mmHg | [8 empty] | 761.746 | mmHg
| Optical DO % |  |  |  | 100.2 | % |  | 100.8 | % | ± 1% |  | P | 11:30 |  | 101.3 | % | F
[6 empty] | DO gain |  | 0.97856 | mg/L (table) | 0.87-1.25 |  | P |  |  | 101.5 | %
| Optical DO mg/L (ICV) | N/A | mg/L (DL) |  | N/A | °C |  | n/a |  | ± 0.1 mg/L | [5 empty] | mg/L (DL)
| Battery Voltage | [6 empty] | 13 | V | ~6V (EXO 2) |  | P |  |  |  | °C
[6 empty] | Optional Parameters | [8 empty] | 9.8 | V | P
| Chlorophyll | 0 | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| Chlorophyll |  | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [5 empty] | °C
| Chlorophyll | 0 | µg/L |  |  | µg/L |  |  | µg/L | [6 empty] | µg/L |  |  |  |  | Post-Deployment Notes
| Chlorophyll |  | µg/L |  |  | µg/L |  |  | µg/L | [6 empty] | µg/L
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [5 empty] | °C
| BGA | 0 | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| BGA |  | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [5 empty] | °C
Field Log
Fouling Presence | [9 empty] | Deployment |  |  | Retrieval |  |  |  | Field Notes:
A=algae, AM=amphipods, B=barnacles, BR=bryozoans, C=crabs, E=eggs, F=fish, FA=filamentous algae, G=gastropod, H=hydroids, I=isopod, MD=mud, S=sponges, SI=silt, SL=shell, SP=shrimp, T=tunicates, W=worm, O=other | [8 empty] | Date deploy/retrieval | 2022-02-18 |  |  | 2022-03-16 |  |  | Deployment:
[9 empty] | Time deploy/retrieval | 12:42 |  |  | 10:20 |  |  | Staff did not record in field measurements. Warm, calm, mid tide rising
[9 empty] | First/last sample time | 12:45 |  |  | 10:15
| H= heavy, M= moderate, L= light | [7 empty] | Technician initials | JB |  |  | MP
Mount: N | [9 empty] | Discrete Samples (Required)
[9 empty] | Depth (m) |  | m |  | 2.3 |  | m
Sonde Body: C/L, A/L, B/L | [8 empty] | Temp (°C) |  | °C |  | 15.5 |  | °C
[9 empty] | DO % |  | % |  | 98.4 |  | %
Field guard: C/L, A/L, B/L | [8 empty] | DO mg/L |  | mg/L |  | 8.19 |  | mg/L
All Sesnors: | [8 empty] | SpCond (mS/cm) |  | mS/cm |  | 46441 |  | mS/cm | Retrieval:
C/T: N | [8 empty] | Salinity (psu) |  | psu |  | 30.19 |  | psu | Sunny, calm, low tide rising.
pH: N | [8 empty] | pH |  |  |  | 8.34
DO: A/L | [11 empty] | Optional Samples
Turb: A/L | [8 empty] | pHmV |  | mV |  |  |  | mV
Wiper: | [8 empty] | Conductivity |  | mS/cm |  |  |  | mS/cm
TAL-PE: | [12 empty] | 0.9 NTU Turb.
Other/notes:

## CPFS_031622
Aquatic Preserve Continuous Water Quality Monitoring Calibration Log
|  | File Name | CPFS_031622 | [5 empty] | .csv | [5 empty] | File Name | CPFS_031622.DAT |  |  |  |  | .bin
[7 empty] | Datasonde and Probe Identification Numbers | [12 empty] | Calibration Standards
|  | Sonde Code |  |  | Serial Number |  |  | Serial Number |  | Model Number | [5 empty] | Standard | Lot No. |  |  | Received | Opened | Exp. Date
| Datasonde | 145459 |  |  | 13C100687 |  | pH | 19F |  | 6151 |  |  |  |  | SpCond (mS/cm) | 20
| Vented | N |  |  | Model |  | ODO | 15L103284 |  | 6150 |  |  |  |  | SpCond (mS/cm) | 50 | 2906F57 |  |  |  |  | 7/2021
| Nickname | FS |  |  | 6600 |  | Turb | 12C100448 |  | 6136 |  |  |  |  | pH | 4 | 160779 |  |  |  |  | 7/13/23
| Preserve | CPAP |  |  |  |  | Cond | 14J100146 |  | 6560 |  |  |  |  | pH | 7 | 21H2S |  |  |  |  | 8/11/23
| Station | CPFS |  |  |  |  | Chl | Chl Serial No. |  | Chl Model |  |  |  |  | pH | 10 | 21G3T |  |  |  |  | 7/13/23
[7 empty] | Wiper | Wiper Serial No. |  | Wiper Model |  |  |  |  | Turb (FNU) | 124 | 21G21280150 |  |  |  | 11/1/21 | 7/2022
|  |  |  | Wiper Replaced | N |  |  |  | Batteries Replaced | Y | [5 empty] | Technician(s) | Jon Brucker
|  |  |  | Wiper Parked in Garage | Y |  |  |  | Date of Calibration | 2022-03-07 | [5 empty] | Date of Post-Calibration | 2022-04-18
Pre-deployment Comments/ Failed-calibration values:
| Pre-Deployment (CAL & ICV) | [12 empty] | Post-Deployment (CCV) | [5 empty] | Programming
|  | Standards |  |  | Before Cal |  |  | Calibrated |  | Specs |  | P/ F | Time |  | Post value |  | P/ F | Time |  | Interval |  | 15
| Temp |  | °C (Therm) |  |  | °C (DL) |  |  |  | ±0.2(w) +0.1°C | [5 empty] | °C (Therm) | P |  |  | Duration |  | 365 days
| SpCond in air | [8 empty] | ≤ 0.002 mS/cm | [5 empty] | °C (DL) |  |  |  | Free memory % |  | N/A
| SpCond | 50 | mS/cm |  | 50.13 | mS/cm |  | 50 | mS/cm | 50 ± 0.25 (n) |  |  |  | ✔ | 49.83 | mS/cm | P |  |  | Memory (days) |  | 132
[6 empty] | Cell Constant |  | 5.031 |  | 0.419-0.519 (w) |  | P | [7 empty] | Start time |  | 7:29
| SpCond 20 (ICV) | 20 | mS/cm |  |  | ✔ |  | N/A | mS/cm | 20 ± 0.2 (w) | [9 empty] | Battery % |  | 60.7 days
| pH 7 | 7 |  |  | 6.86 |  |  | 7 |  | ± 0.1 |  |  |  | ✔ | 7 |  | P |  |  | Set clock |  | Y
| pH 10 | 10 |  |  | 9.95 |  |  | 10 |  | ± 0.1 |  |  |  | ✔ | N/A |  |  |  |  | End time |  | 8:30
[5 empty] | pH 7 |  |  | 2.1 | mV | 0 ±50 mV |  | P |  |  | 1.9 | mV | P |  |  |  | Other/notes
[5 empty] | pH 10 |  |  | -164.7 | mV | -180 ±50 mV |  | P |  |  | N/A | mV
[5 empty] | pH 4 |  |  |  | mV | +180 ±50 mV |  | P
[5 empty] | pH Slope |  |  | 166.8 |  | > 155 mV |  | P |  |  | #VALUE!
| pH 4 or 10 (ICV) | 4 |  |  |  | ✔ |  | N/A |  | ± 0.1
| BP at Depth Cal |  |  |  | 765.81 | mmHg | [8 empty] | 760.2 | mmHg
| Depth | 0.079 | Offset |  | 0.072 | m |  | 0.079 | m | ±0.004 m |  | P |  |  | 0.007 | m | P
| Turbidity | 0 | FNU |  | -2 |  |  | 0 | FNU | 0.00 ±0.3 |  |  |  | ✔ | -0.1 | FNU | P |  |  |  | General Maintenance
| Turbidity | 126 | FNU |  | 125 | FNU |  | 126 | FNU | 124 ±2.48 |  |  |  | ✔ | N/A | FNU |  |  | Sonde tube cleaning, lock greased.
| Turbidity 124 (ICV) | 126 | FNU |  |  |  | ✔ | N/A | FNU | 124 ±2.48
| BP at ODO Cal |  |  |  | 765.81 | mmHg | [8 empty] | 760.2 | mmHg
| Optical DO % |  |  |  | 103.3 | % |  | 100.7 | % | ± 1% |  | P |  |  | 88.7 | % | F
[6 empty] | DO gain |  | 1.0261 | mg/L (table) | 0.87-1.25 |  | P |  |  | 88.4 | %
| Optical DO mg/L (ICV) | N/A | mg/L (DL) |  | N/A | °C |  | n/a |  | ± 0.1 mg/L |  |  |  |  | N/A | mg/L (DL)
| Battery Voltage | [6 empty] | 11.5 | V | ~6V (EXO 2) |  | P |  |  | n/A | °C
[6 empty] | Optional Parameters | [8 empty] | 10.7 | V | P
| Chlorophyll | 0 | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| Chlorophyll |  | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [5 empty] | °C
| Chlorophyll | 0 | µg/L |  |  | µg/L |  |  | µg/L | [6 empty] | µg/L |  |  |  |  | Post-Deployment Notes
| Chlorophyll |  | µg/L |  |  | µg/L |  |  | µg/L | [6 empty] | µg/L |  |  | DO % values from the last two recorded before stopping the sonde. However, readings were around 100% previous to these, may possibly have read while towel was removed as was reading fine in saturated water bath.
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [5 empty] | °C
| BGA | 0 | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| BGA |  | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [5 empty] | °C
Field Log
Fouling Presence | [9 empty] | Deployment |  |  | Retrieval |  |  |  | Field Notes:
A=algae, AM=amphipods, B=barnacles, BR=bryozoans, C=crabs, E=eggs, F=fish, FA=filamentous algae, G=gastropod, H=hydroids, I=isopod, MD=mud, S=sponges, SI=silt, SL=shell, SP=shrimp, T=tunicates, W=worm, O=other | [8 empty] | Date deploy/retrieval | 2022-03-16 |  |  | 2022-04-15 |  |  | Deployment:
[9 empty] | Time deploy/retrieval | 10:23 |  |  | 8:47 |  |  | Turbidity: 0.9 NTU; Foggy, calm, low tide rising.
[9 empty] | First/last sample time | 10:30 |  |  | 8:30
| H= heavy, M= moderate, L= light | [7 empty] | Technician initials | MP |  |  | MP
Mount: N | [9 empty] | Discrete Samples (Required)
[9 empty] | Depth (m) | 2.3 | m |  | 2.5 |  | m
Sonde Body: B/L, MD/L | [8 empty] | Temp (°C) | 15.5 | °C |  | 19.8 |  | °C
[9 empty] | DO % | 98.4 | % |  | 85 |  | %
Field guard: B/L, MD/L, C/M | [8 empty] | DO mg/L | 8.19 | mg/L |  | 6.64 |  | mg/L
All Sesnors: | [8 empty] | SpCond (mS/cm) | 46441 | mS/cm |  | 41.745 |  | mS/cm | Retrieval:
C/T: MD/L, T/L | [8 empty] | Salinity (psu) | 30.19 | psu |  | 26.84 |  | psu | Turbidity: 0.7 NTU; Sunny, wind NE 10-15, low tide rising.
pH: MD/L, B/L | [8 empty] | pH | 8.34 |  |  | 7.94
DO: MD/L, B/L | [11 empty] | Optional Samples
Turb: MD/L, B/L | [8 empty] | pHmV |  | mV |  |  |  | mV
Wiper: | [8 empty] | Conductivity |  | mS/cm |  |  |  | mS/cm
TAL-PE:
Other/notes:

## CPFS_041522
Aquatic Preserve Continuous Water Quality Monitoring Calibration Log
|  | File Name | CPFS_041522 | [5 empty] | .csv | [5 empty] | File Name | CPFS_041522.DAT |  |  |  |  | .bin
[7 empty] | Datasonde and Probe Identification Numbers | [12 empty] | Calibration Standards
|  | Sonde Code |  |  | Serial Number |  |  | Serial Number |  | Model Number | [5 empty] | Standard | Lot No. |  |  | Received | Opened | Exp. Date
| Datasonde | 145154 |  |  | 12H102535 |  | pH | 21A |  | 6151 |  |  |  |  | SpCond (mS/cm) | 20 | 4203H74 |  |  |  | 4/11/2022 | 3/2024
| Vented | N |  |  | Model |  | ODO | 14D100738 |  | 6150 |  |  |  |  | SpCond (mS/cm) | 50 | 2906F57 |  |  |  |  | 6/2021
| Nickname | FS #17 |  |  | 6600 |  | Turb | 14C103217 |  | 6136 |  |  |  |  | pH | 4
| Preserve | CPAP |  |  |  |  | Cond | 16G100618 |  | 6560 |  |  |  |  | pH | 7 | 21H2S |  |  |  | 3/23/2022 | 8/11/2023
| Station | CPFS |  |  |  |  | Chl | Chl Serial No. |  | Chl Model |  |  |  |  | pH | 10 | 21G3T |  |  |  |  | 7/13/2023
[7 empty] | Wiper | Wiper Serial No. |  | Wiper Model |  |  |  |  | Turb (FNU) | 124 | 607300 |  |  |  | 3/23/2022 | 11/2022
|  |  |  | Wiper Replaced | N |  |  |  | Batteries Replaced | Y | [5 empty] | Technician(s) | Michael Palandri
|  |  |  | Wiper Parked in Garage |  |  |  |  | Date of Calibration | 2022-04-13 | [5 empty] | Date of Post-Calibration | 2022-05-19
| Pre-deployment Comments/ Failed-calibration values: | pH 7 reading low for calibration, ph 10 millivolts were off, however, best probe tested out of 5 available. Turned on 4/14/2022, but was deployed on 4/15/2022 due to time constraints for deployments.Sonde is a 6600, different post-cal specifications.
| Pre-Deployment (CAL & ICV) | [12 empty] | Post-Deployment (CCV) | [5 empty] | Programming
|  | Standards |  |  | Before Cal |  |  | Calibrated |  | Specs |  | P/ F | Time |  | Post value |  | P/ F | Time |  | Interval |  | 15
| Temp |  | °C (Therm) |  |  | °C (DL) |  |  |  | ±0.2(w) +0.1°C | [5 empty] | °C (Therm) |  |  |  | Duration |  | 365 days
| SpCond in air | [8 empty] | ≤ 0.002 mS/cm | [5 empty] | °C (DL) |  |  |  | Free memory %
| SpCond | 50 | mS/cm |  | 49.1 | mS/cm |  | 50 | mS/cm | 50 ± 0.25 (n) |  |  | 12:18 | ✔ | 48.61 | mS/cm | F |  |  | Memory (days) |  | 185
[6 empty] | Cell Constant |  | 5.0925 |  | 0.419-0.519 (w) |  | P | [7 empty] | Start time |  | 8:44
| SpCond 20 (ICV) | 20 | mS/cm |  |  | ✔ |  | 20.63 | mS/cm | 20 ± 0.2 (w) |  | F | [7 empty] | Battery % |  | 62.8 days
| pH 7 | 7 |  |  | 5.73 |  |  | 7 |  | ± 0.1 |  |  |  | ✔ | 7.08 |  | P |  |  | Set clock |  | Y
| pH 10 | 10 |  |  | 10.08 |  |  | 10 |  | ± 0.1 |  |  |  | ✔ | [5 empty] | End time |  | 15:47
[5 empty] | pH 7 |  |  | 45.4 | mV | 0 ±50 mV |  | P |  |  | 41.7 | mV | P |  |  |  | Other/notes
[5 empty] | pH 10 |  |  | -110 | mV | -180 ±50 mV |  | P |  |  |  | mV
[5 empty] | pH 4 |  |  |  | mV | +180 ±50 mV
[5 empty] | pH Slope |  |  | 155.4 |  | > 155 mV |  | P |  |  | 41.7
| pH 4 or 10 (ICV) | 4 |  |  |  | ✔ |  |  |  | ± 0.1
| BP at Depth Cal |  |  |  | 762.89 | mmHg | [8 empty] | 760.9 | mmHg
| Depth | 0.041 | Offset |  | 0.066 | m |  | 0.041 | m | ±0.004 m |  | P |  |  | 0.01 | m | F
| Turbidity | 0 | FNU |  | -3.3 |  |  | 0 | FNU | 0.00 ±0.3 |  |  |  | ✔ | 0.5 | FNU | F |  |  |  | General Maintenance
| Turbidity | 124 | FNU |  | 131.6 | FNU |  | 126 | FNU | 124 ±2.48 |  |  |  | ✔ |  | FNU |  |  | Cleaned sonde tube, greased locks.
| Turbidity 124 (ICV) | 124 | FNU |  |  |  | ✔ |  | FNU | 124 ±2.48
| BP at ODO Cal |  |  |  | 762.89 | mmHg | [8 empty] | 760.9 | mmHg
| Optical DO % |  |  |  | 102.6 | % |  | 100.4 | % | ± 1% |  | P |  |  | 99.7 | % | P
[6 empty] | DO gain |  | 0.95713 | mg/L (table) | 0.87-1.25 |  | P |  |  | 99.6 | %
| Optical DO mg/L (ICV) |  | mg/L (DL) |  |  | °C |  |  |  | ± 0.1 mg/L | [5 empty] | mg/L (DL)
| Battery Voltage | [6 empty] | 12.3 | V | ~6V (EXO 2) |  | P |  |  |  | °C
[6 empty] | Optional Parameters | [9 empty] | V
| Chlorophyll | 0 | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| Chlorophyll |  | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [5 empty] | °C
| Chlorophyll | 0 | µg/L |  |  | µg/L |  |  | µg/L | [6 empty] | µg/L |  |  |  |  | Post-Deployment Notes
| Chlorophyll |  | µg/L |  |  | µg/L |  |  | µg/L | [6 empty] | µg/L |  |  | Ram rod sonde tube, greased locks. Hole in mesh over field guard, probably from goby.
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [5 empty] | °C
| BGA | 0 | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| BGA |  | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [5 empty] | °C
Field Log
Fouling Presence | [9 empty] | Deployment |  |  | Retrieval |  |  |  | Field Notes:
A=algae, AM=amphipods, B=barnacles, BR=bryozoans, C=crabs, E=eggs, F=fish, FA=filamentous algae, G=gastropod, H=hydroids, I=isopod, MD=mud, S=sponges, SI=silt, SL=shell, SP=shrimp, T=tunicates, W=worm, O=other | [8 empty] | Date deploy/retrieval | 2022-04-15 |  |  | 2022-05-10 |  |  | Deployment:
[9 empty] | Time deploy/retrieval | 8:50 |  |  | 13:09 |  |  | Turbidity: 0.7 NTU; low tide rising, wind NE 10-15 mph.
[9 empty] | First/last sample time | 9:00 |  |  | 13:00
| H= heavy, M= moderate, L= light | [7 empty] | Technician initials | MP |  |  | MP
Mount: N | [9 empty] | Discrete Samples (Required)
[9 empty] | Depth (m) | 2.5 | m |  | 2.8 |  | m
Sonde Body: MD/L, E/M | [8 empty] | Temp (°C) | 19.8 | °C |  | 23.3 |  | °C
[9 empty] | DO % | 85 | % |  | 107.8 |  | %
Field guard: F/L (one goby), A/M, B/L, E/L | [8 empty] | DO mg/L | 6.64 | mg/L |  | 7.86 |  | mg/L
All Sesnors: | [8 empty] | SpCond (mS/cm) | 41.745 | mS/cm |  | 44218 |  | mS/cm | Retrieval:
C/T: A/M | [8 empty] | Salinity (psu) | 26.84 | psu |  | 28.57 |  | psu | Turbidity: 1.3 FNU, sunny, warm, wind east 10-15mph, mid-tide rising.
pH: A/M | [8 empty] | pH | 7.94 |  |  | 7.06
DO: A/M | [11 empty] | Optional Samples
Turb: E/M, A/M | [8 empty] | pHmV |  | mV |  |  |  | mV
Wiper: | [8 empty] | Conductivity |  | mS/cm |  |  |  | mS/cm
TAL-PE:
Other/notes:

## CPFS_051022
Aquatic Preserve Continuous Water Quality Monitoring Calibration Log
|  | File Name | CPFS_051022 | [5 empty] | .csv | [5 empty] | File Name | CPFS_051022.DAT |  |  |  |  | .bin
[7 empty] | Datasonde and Probe Identification Numbers | [12 empty] | Calibration Standards
|  | Sonde Code |  |  | Serial Number |  |  | Serial Number |  | Model Number | [5 empty] | Standard | Lot No. |  |  | Received | Opened | Exp. Date
| Datasonde | 145459 |  |  | 13C100687 |  | pH | 19E |  | 6151 |  |  |  |  | SpCond (mS/cm) | 20 | 4203H74 |  |  |  | 4/11/2022 | 3/2024
| Vented | N |  |  | Model |  | ODO | 15L103284 |  | 6150 |  |  |  |  | SpCond (mS/cm) | 50 | 2906F57 |  |  |  |  | 6/2021
| Nickname | FS#19 |  |  | 6600 |  | Turb | 12C100448 |  | 6136 |  |  |  |  | pH | 4
| Preserve | CPAP |  |  |  |  | Cond | 14J100146 |  | 6560 |  |  |  |  | pH | 7 | 21M2S |  |  |  |  | 12/8/2023
| Station | CPFS |  |  |  |  | Chl | Chl Serial No. |  | Chl Model |  |  |  |  | pH | 10 | 22A1T |  |  |  |  | 1/3/2024
[7 empty] | Wiper | Wiper Serial No. |  | Wiper Model |  |  |  |  | Turb (FNU) | 124 | 21L21460003 |  |  |  | 3/23/2022 | 11/2022
|  |  |  | Wiper Replaced | Y |  |  |  | Batteries Replaced | N | [5 empty] | Technician(s) | Michael Palandri
|  |  |  | Wiper Parked in Garage |  |  |  |  | Date of Calibration | 2022-05-09 | [5 empty] | Date of Post-Calibration | 2022-06-09
Pre-deployment Comments/ Failed-calibration values:
| Pre-Deployment (CAL & ICV) | [12 empty] | Post-Deployment (CCV) | [5 empty] | Programming
|  | Standards |  |  | Before Cal |  |  | Calibrated |  | Specs |  | P/ F | Time |  | Post value |  | P/ F | Time |  | Interval |  | 15
| Temp |  | °C (Therm) |  |  | °C (DL) |  |  |  | ±0.2(w) +0.1°C |  |  |  |  | 21.4 | °C (Therm) | F | 10:56 |  | Duration |  | 365 days
| SpCond in air | [8 empty] | ≤ 0.002 mS/cm |  |  |  |  | 20.95 | °C (DL) |  |  |  | Free memory % |  | 391424 bytes
| SpCond | 50 | mS/cm |  | 49.8 | mS/cm |  | 50 | mS/cm | 50 ± 0.25 (n) |  |  | 13:37 | ✔ | 50.93 | mS/cm | F | 11:05 |  | Memory (days) |  | 185
[6 empty] | Cell Constant |  | 5.0511 |  | 0.419-0.519 (w) |  | P | [7 empty] | Start time |  | 13:45
| SpCond 20 (ICV) | 20 | mS/cm |  |  | ✔ |  |  | mS/cm | 20 ± 0.2 (w) | [9 empty] | Battery % |  | 10.9 V
| pH 7 | 7 |  |  | 6.96 |  |  | 7 |  | ± 0.1 |  |  | 13:43 | ✔ | 7.45 |  | F | 11:19 |  | Set clock |  | Y
| pH 10 | 10 |  |  | 9.86 |  |  | 10 |  | ± 0.1 |  |  | 13:57 | ✔ | 10.55 |  | F | 11:22 |  | End time |  | 12:15
[5 empty] | pH 7 |  |  | 18 | mV | 0 ±50 mV |  | P |  |  | -5.6 | mV | P |  |  |  | Other/notes
[5 empty] | pH 10 |  |  | -141.1 | mV | -180 ±50 mV |  | P |  |  | -170.3 | mV | P |  | Sonde turned off on 6/3/2022. Sonde was accidentally programmed to start after it was already deployed. No "head" in data.
[5 empty] | pH 4 |  |  |  | mV | +180 ±50 mV
[5 empty] | pH Slope |  |  | 159.1 |  | > 155 mV |  | P |  |  | 164.7 |  | P
| pH 4 or 10 (ICV) | 10 |  |  |  | ✔ |  | 10.03 |  | ± 0.1 |  | P | 14:02
| BP at Depth Cal |  |  |  | 762.51 | mmHg | [8 empty] | 761.01 | mmHg
| Depth | 0.041 | Offset |  | 0.041 | m |  | 0.041 | m | ±0.004 m |  | P | 14:24 |  | 0.022 | m | F | 11:05
| Turbidity | 0 | FNU |  | -0.4 |  |  | 0 | FNU | 0.00 ±0.3 |  |  | 14:09 | ✔ | 0.8 | FNU | F | 11:26 |  |  | General Maintenance
| Turbidity | 124 | FNU |  | 127.9 | FNU |  | 126 | FNU | 124 ±2.48 |  |  | 14:13 | ✔ | 127.8 | FNU | F | 11:28 | Ram rod sonde tube and greased lock on deployment.
| Turbidity 124 (ICV) | 124 | FNU |  |  |  | ✔ | 126 | FNU | 124 ±2.48 |  | P | 14:14
| BP at ODO Cal |  |  |  | 765.5 | mmHg | [8 empty] | 759.8 | mmHg
| Optical DO % |  |  |  | 101 | % |  | 100.5 | % | ± 1% |  | P | 9:03 |  | 98.8 | % | F | 10:50
[6 empty] | DO gain |  | 1.0243 | mg/L (table) | 0.87-1.25 |  | P |  |  | 98.9 | %
| Optical DO mg/L (ICV) | 8.87 | mg/L (DL) |  | 21.43 | °C |  | 8.915 |  | ± 0.1 mg/L |  | P | 9:08 |  | 8.87 | mg/L (DL)
| Battery Voltage | [6 empty] | 10.9 | V | ~6V (EXO 2) |  | P |  |  | 20.95 | °C | P | 13:30
[6 empty] | Optional Parameters | [8 empty] | 11.4 | V | P
| Chlorophyll | 0 | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| Chlorophyll |  | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [5 empty] | °C
| Chlorophyll | 0 | µg/L |  |  | µg/L |  |  | µg/L | [6 empty] | µg/L |  |  |  |  | Post-Deployment Notes
| Chlorophyll |  | µg/L |  |  | µg/L |  |  | µg/L | [6 empty] | µg/L |  |  | DO probe would quit reading during post-cal, giving 0 for concentration and saturation. Did not do this during deployment, however. Sonde had connection issues on retrieval, wouldn't load menu on 6/2 but was fine on 6/3.
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [5 empty] | °C
| BGA | 0 | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| BGA |  | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [5 empty] | °C
Field Log
Fouling Presence | [9 empty] | Deployment |  |  | Retrieval |  |  |  | Field Notes:
A=algae, AM=amphipods, B=barnacles, BR=bryozoans, C=crabs, E=eggs, F=fish, FA=filamentous algae, G=gastropod, H=hydroids, I=isopod, MD=mud, S=sponges, SI=silt, SL=shell, SP=shrimp, T=tunicates, W=worm, O=other | [8 empty] | Date deploy/retrieval | 2022-05-10 |  |  | 2022-06-02 |  |  | Deployment:
[9 empty] | Time deploy/retrieval | 13:12 |  |  | 10:51 |  |  | Turbidity: 1.3 FNU, wind east 10-15mph, sunny, warm, mid-tide rising
[9 empty] | First/last sample time | 14:00 |  |  | 10:45
| H= heavy, M= moderate, L= light | [7 empty] | Technician initials | MP |  |  | MC/LC
Mount: | [9 empty] | Discrete Samples (Required)
[9 empty] | Depth (m) | 2.8 | m |  |  |  | m
Sonde Body: | [8 empty] | Temp (°C) | 23.3 | °C |  | 26.3 |  | °C
[9 empty] | DO % | 107.8 | % |  | 107.4 |  | %
Field guard: A/L, C/L | [8 empty] | DO mg/L | 7.86 | mg/L |  | 7.29 |  | mg/L
All Sesnors: | [8 empty] | SpCond (mS/cm) | 44.218 | mS/cm |  | 46.823 |  | mS/cm | Retrieval:
C/T: | [8 empty] | Salinity (psu) | 28.57 | psu |  | 30.36 |  | psu | Turbidity: 3.3 FNU, low tide rising, sunny
pH: | [8 empty] | pH | 7.06 |  |  | 7.77
DO: | [11 empty] | Optional Samples
Turb: | [8 empty] | pHmV |  | mV |  |  |  | mV
Wiper: | [8 empty] | Conductivity |  | mS/cm |  |  |  | mS/cm
TAL-PE:
Other/notes:

## CPFS_060222
Aquatic Preserve Continuous Water Quality Monitoring Calibration Log
|  | File Name | CPFS_060222 | [5 empty] | .csv | [5 empty] | File Name | CPFS_060222.DAT |  |  |  |  | .bin
[7 empty] | Datasonde and Probe Identification Numbers | [12 empty] | Calibration Standards
|  | Sonde Code |  |  | Serial Number |  |  | Serial Number |  | Model Number | [5 empty] | Standard | Lot No. |  |  | Received | Opened | Exp. Date
| Datasonde | 145458 |  |  | 13C100687 |  | pH | 21A |  | 6151 |  |  |  |  | SpCond (mS/cm) | 20 | 4203H74 |  |  |  | 4/11/2022 | 3/2024
| Vented | N |  |  | Model |  | ODO | 14D100738 |  | 6150 |  |  |  |  | SpCond (mS/cm) | 50 | 2906F57 |  |  |  |  | 6/2021
| Nickname | AH18 |  |  | 6600 |  | Turb | 14C103217 |  | 6136 |  |  |  |  | pH | 4
| Preserve | CPAP |  |  |  |  | Cond | 16G100618 |  | 6560 |  |  |  |  | pH | 7 | 21M2S |  |  |  |  | 12/8/2023
| Station | CPFS |  |  |  |  | Chl | Chl Serial No. |  | Chl Model |  |  |  |  | pH | 10 | 22A1T |  |  |  |  | 1/3/2024
[7 empty] | Wiper | Wiper Serial No. |  | Wiper Model |  |  |  |  | Turb (FNU) | 124 | 21L21460003 |  |  |  | 3/23/2022 | 11/2022
|  |  |  | Wiper Replaced | N |  |  |  | Batteries Replaced | N | [5 empty] | Technician(s) | Michael Palandri
|  |  |  | Wiper Parked in Garage | N |  |  |  | Date of Calibration | 2022-06-01 | [5 empty] | Date of Post-Calibration | 2022-06-24
| Pre-deployment Comments/ Failed-calibration values: | DO was not calibrated, pay special attention to during post-cal.Turbidity was reading low pre-cal in the 126 NTU standard.
| Pre-Deployment (CAL & ICV) | [12 empty] | Post-Deployment (CCV) | [5 empty] | Programming
|  | Standards |  |  | Before Cal |  |  | Calibrated |  | Specs |  | P/ F | Time |  | Post value |  | P/ F | Time |  | Interval |  | 15
| Temp |  | °C (Therm) |  |  | °C (DL) |  |  |  | ±0.2(w) +0.1°C |  |  |  |  | 20.6 | °C (Therm) | P | 9:37 |  | Duration |  | 365 days
| SpCond in air |  |  |  | 0 |  |  |  |  | ≤ 0.002 mS/cm |  | P |  |  | 20.41 | °C (DL) |  |  |  | Free memory %
| SpCond | 50 | mS/cm |  | 48.24 | mS/cm |  | 50.01 | mS/cm | 50 ± 0.25 (n) |  |  | 12:33 | ✔ | 50.25 | mS/cm | P | 9:50 |  | Memory (days) |  | 185
[6 empty] | Cell Constant |  | 4.9947 |  | 0.419-0.519 (w) |  | P | [7 empty] | Start time |  | 9:15
| SpCond 20 (ICV) | 20 | mS/cm |  |  | ✔ |  |  | mS/cm | 20 ± 0.2 (w) | [9 empty] | Battery % |  | 61.9 days
| pH 7 | 7 |  |  | 7.04 |  |  | 7 |  | ± 0.1 |  |  | 12:40 | ✔ | 7.17 |  | F | 9:54 |  | Set clock |  | Y
| pH 10 | 10 |  |  | 10.07 |  |  | 10 |  | ± 0.1 |  |  | 12:44 | ✔ | 10.18 |  | F | 10:04 |  | End time |  | 12:36
[5 empty] | pH 7 |  |  | 15.8 | mV | 0 ±50 mV |  | P |  |  | 5.1 | mV | P |  |  |  | Other/notes
[5 empty] | pH 10 |  |  | -155 | mV | -180 ±50 mV |  | P |  |  | -164.7 | mV | P
[5 empty] | pH 4 |  |  |  | mV | +180 ±50 mV
[5 empty] | pH Slope |  |  | 170.8 |  | > 155 mV |  | P |  |  | 169.8 |  | P
| pH 4 or 10 (ICV) | 10 |  |  |  | ✔ |  | 9.99 |  | ± 0.1 |  | P | 12:45
| BP at Depth Cal |  |  |  | 761.2 | mmHg | [8 empty] | 758.91 | mmHg
| Depth | 0.014 | Offset |  | 0.021 | m |  | 0.014 | m | ±0.004 m |  | P | 13:00 |  | -0.015 | m | P | 9:42
| Turbidity | 0 | FNU |  | 0 |  |  | 0 | FNU | 0.00 ±0.3 |  |  | 12:51 | ✔ | 0.6 | FNU | F | 10:08 |  |  | General Maintenance
| Turbidity | 124 | FNU |  | 113.6 | FNU |  | 126 | FNU | 124 ±2.48 |  |  | 12:54 | ✔ | 124.2 | FNU | P | 10:15 | Ram rod and greased locks on retrieval.
| Turbidity 124 (ICV) | 124 | FNU |  |  |  | ✔ | 125.8 | FNU | 124 ±2.48 |  | P | 12:56
| BP at ODO Cal |  |  |  |  | mmHg | [8 empty] | 758.91 | mmHg
| Optical DO % |  |  |  |  | % |  |  | % | ± 1% |  |  |  |  | 97.5 | % | F | 9:39
[6 empty] | DO gain |  |  | mg/L (table) | 0.87-1.25 |  |  |  |  | 94.9 | %
| Optical DO mg/L (ICV) |  | mg/L (DL) |  |  | °C |  |  |  | ± 0.1 mg/L |  |  |  |  | 8.48 | mg/L (DL)
| Battery Voltage | [6 empty] | 11.7 | V | ~6V (EXO 2) |  | P |  |  | 20.4 | °C | F | 9:40
[6 empty] | Optional Parameters | [8 empty] | 10.9 | V | P
| Chlorophyll | 0 | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| Chlorophyll |  | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [5 empty] | °C
| Chlorophyll | 0 | µg/L |  |  | µg/L |  |  | µg/L | [6 empty] | µg/L |  |  |  |  | Post-Deployment Notes
| Chlorophyll |  | µg/L |  |  | µg/L |  |  | µg/L | [6 empty] | µg/L |  |  | Damage to field guard screen. No wiper pads on DO or turbidity. Turbidity probe took long to stabilize in post-cal.
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [5 empty] | °C
| BGA | 0 | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| BGA |  | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [5 empty] | °C
Field Log
Fouling Presence | [9 empty] | Deployment |  |  | Retrieval |  |  |  | Field Notes:
A=algae, AM=amphipods, B=barnacles, BR=bryozoans, C=crabs, E=eggs, F=fish, FA=filamentous algae, G=gastropod, H=hydroids, I=isopod, MD=mud, S=sponges, SI=silt, SL=shell, SP=shrimp, T=tunicates, W=worm, O=other | [8 empty] | Date deploy/retrieval | 2022-06-02 |  |  | 2022-06-23 |  |  | Deployment:
[9 empty] | Time deploy/retrieval | 10:56 |  |  | 10:10 |  |  | Turbidity: 3.3 FNU, sunny, low-tide rising
[9 empty] | First/last sample time | 11:00 |  |  | 10:15
| H= heavy, M= moderate, L= light | [7 empty] | Technician initials | MC/LC |  |  | MP
Mount: N | [9 empty] | Discrete Samples (Required)
[9 empty] | Depth (m) |  | m |  | 2.7 |  | m
Sonde Body: N | [8 empty] | Temp (°C) | 26.3 | °C |  | 28.4 |  | °C
[9 empty] | DO % | 107.4 | % |  |  |  | %
Field guard: C/H, A/L, SL/L, B/L | [8 empty] | DO mg/L | 7.29 | mg/L |  | 6.53 |  | mg/L
All Sesnors: | [8 empty] | SpCond (mS/cm) | 46.823 | mS/cm |  | 48.371 |  | mS/cm | Retrieval:
C/T: B/L, MD/L, A/L | [8 empty] | Salinity (psu) | 30.36 | psu |  | 31.45 |  | psu | pH probe malfunctioning, value likely higher than actual. Turbidity:0.9 FNU, sunny wind N @ 5mph, high-tide falling
pH: B/L, MD/L, A/L | [8 empty] | pH | 7.77 |  |  | 9.06*
DO: B/L, MD/L, A/L | [11 empty] | Optional Samples
Turb:B/L, MD/L, A/L | [8 empty] | pHmV |  | mV |  |  |  | mV
Wiper: | [8 empty] | Conductivity |  | mS/cm |  |  |  | mS/cm
TAL-PE:
Other/notes:

## CPFS_062322
Aquatic Preserve Continuous Water Quality Monitoring Calibration Log
|  | File Name | CPFS_062322 | [5 empty] | .csv | [5 empty] | File Name | CPFS_062322.DAT |  |  |  |  | .bin
[7 empty] | Datasonde and Probe Identification Numbers | [12 empty] | Calibration Standards
|  | Sonde Code |  |  | Serial Number |  |  | Serial Number |  | Model Number | [5 empty] | Standard | Lot No. |  |  | Received | Opened | Exp. Date
| Datasonde | Datasonde |  |  | Sonde No. |  | pH | 19E |  | 6151 |  |  |  |  | SpCond (mS/cm) | 20 | 4203H74 |  |  |  | 4/11/2022 | 3/2024
| Vented | N |  |  | Model |  | ODO | 13C100357 |  | 6150 |  |  |  |  | SpCond (mS/cm) | 50 | 2906F57 |  |  |  |  | 6/2021
| Nickname | FS19 |  |  | Sonde Model |  | Turb | 12C100448 |  | 6136 |  |  |  |  | pH | 4
| Preserve | CPAP |  |  |  |  | Cond | 14J100146 |  | 6560 |  |  |  |  | pH | 7 | 21M2S |  |  |  |  | 12/8/2023
| Station | CPFS |  |  |  |  | Chl | Chl Serial No. |  | Chl Model |  |  |  |  | pH | 10 | 22A1T |  |  |  |  | 1/3/2024
[7 empty] | Wiper | Wiper Serial No. |  | Wiper Model |  |  |  |  | Turb (FNU) | 124 | 22D22130237 |  |  |  |  | 4/2023
|  |  |  | Wiper Replaced | Y |  |  |  | Batteries Replaced | N | [5 empty] | Technician(s) | Michael Palandri
|  |  |  | Wiper Parked in Garage |  |  |  |  | Date of Calibration | 2022-06-22 | [5 empty] | Date of Post-Calibration | 2022-07-15
| Pre-deployment Comments/ Failed-calibration values: | ICV 20mS/cm reading high due to old 50mS/cm standard.
| Pre-Deployment (CAL & ICV) | [12 empty] | Post-Deployment (CCV) | [5 empty] | Programming
|  | Standards |  |  | Before Cal |  |  | Calibrated |  | Specs |  | P/ F | Time |  | Post value |  | P/ F | Time |  | Interval |  | 15
| Temp | 23.1 | °C (Therm) |  | 22.91 | °C (DL) |  |  |  | ±0.2(w) +0.1°C |  | P | 10:52 |  | 21.3 | °C (Therm) | P | 12:07 |  | Duration |  | 365 days
| SpCond in air |  |  |  | 0.002 |  |  |  |  | ≤ 0.002 mS/cm |  | P |  |  | 21.02 | °C (DL) |  |  |  | Free memory %
| SpCond | 50 | mS/cm |  | 51.27 | mS/cm |  | 50 | mS/cm | 50 ± 0.25 (n) |  |  | 11:04 | ✔ | 49.89 | mS/cm | P | 12:10 |  | Memory (days) |  | 185
[6 empty] | Cell Constant |  | 4.9248 |  | 0.419-0.519 (w) |  | P | [7 empty] | Start time |  | 6:29
| SpCond 20 (ICV) | 20 | mS/cm |  |  | ✔ |  | 20.4 | mS/cm | 20 ± 0.2 (w) |  | F | 11:09 | [6 empty] | Battery % |  | 58.9
| pH 7 | 7 |  |  | 7.3 |  |  | 7 |  | ± 0.1 |  |  | 11:21 | ✔ | 7.31 |  | F | 12:17 |  | Set clock |  | Y
| pH 10 | 10 |  |  | 10.13 |  |  | 10 |  | ± 0.1 |  |  | 11:29 | ✔ | 10.26 |  | F | 12:21 |  | End time |  | 12:58
[5 empty] | pH 7 |  |  | 1.7 | mV | 0 ±50 mV |  | P |  |  | -14.6 | mV | P |  |  |  | Other/notes
[5 empty] | pH 10 |  |  | -163.2 | mV | -180 ±50 mV |  | P |  |  | -177.6 | mV | P |  | Free Bytes: 391424
[5 empty] | pH 4 |  |  |  | mV | +180 ±50 mV
[5 empty] | pH Slope |  |  | 164.9 |  | > 155 mV |  | P |  |  | 163 |  | P
| pH 4 or 10 (ICV) | 10 |  |  |  | ✔ |  | 9.99 |  | ± 0.1 |  | P | 11:31
| BP at Depth Cal |  |  |  | 761.76 | mmHg | [8 empty] | 764.23 | mmHg
| Depth | 0.027 | Offset |  | 0.031 | m |  | 0.027 | m | ±0.004 m |  | P | 11:44 |  | 0.06 | m | F | 12:35
| Turbidity | 0 | NTU |  | 0.4 |  |  | 0 | NTU | 0.00 ±0.3 |  |  | 11:35 | ✔ | 0 | FNU | P | 12:27 |  |  | General Maintenance
| Turbidity | 126 | NTU |  | 126.8 | NTU |  | 126 | NTU | 124 ±2.48 |  |  | 11:40 | ✔ | 128.7 | FNU | F | 12:30 | Ram Rod and lock grease on deployment and retrieval.
| Turbidity 124 (ICV) | 126 | NTU |  |  |  | ✔ | 126.2 | NTU | 124 ±2.48 |  | P | 11:41
| BP at ODO Cal |  |  |  | 758.9 | mmHg | [8 empty] | - | mmHg
| Optical DO % |  |  |  | 101.8 | % |  | 99.8 | % | ± 1% |  | P | 6:25 |  | - | %
[6 empty] | DO gain |  | 1.0214 | mg/L (table) | 0.87-1.25 |  | P |  |  | - | %
| Optical DO mg/L (ICV) | 9.01 | mg/L (DL) |  | 20.2 | °C |  | 9.09 |  | ± 0.1 mg/L |  | P | 6:25 |  | - | mg/L (DL)
| Battery Voltage | [6 empty] | 11.4 | V | ~6V (EXO 2) |  | P |  |  | - | °C
[6 empty] | Optional Parameters | [9 empty] | V
| Chlorophyll | 0 | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| Chlorophyll |  | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [5 empty] | °C
| Chlorophyll | 0 | µg/L |  |  | µg/L |  |  | µg/L | [6 empty] | µg/L |  |  |  |  | Post-Deployment Notes
| Chlorophyll |  | µg/L |  |  | µg/L |  |  | µg/L | [6 empty] | µg/L |  |  | DO probe had malfunctioned and needed to be removed for post-cal. While attached, all probes were reading very incorrectly, then read normally once DO probe was removed. Sonde quit recording data after 7/12/2022 at 14:30 due to DO probe malfunction. Wiper was not parked 180 degrees to DO sensor, piece of plastic slightly over sensor face, indicated probe failure. Large gaps and incorrect probe readings will be in data.
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [5 empty] | °C
| BGA | 0 | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| BGA |  | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [5 empty] | °C
Field Log
Fouling Presence | [9 empty] | Deployment |  |  | Retrieval |  |  |  | Field Notes:
A=algae, AM=amphipods, B=barnacles, BR=bryozoans, C=crabs, E=eggs, F=fish, FA=filamentous algae, G=gastropod, H=hydroids, I=isopod, MD=mud, S=sponges, SI=silt, SL=shell, SP=shrimp, T=tunicates, W=worm, O=other | [8 empty] | Date deploy/retrieval | 2022-06-23 |  |  | 2022-07-13 |  |  | Deployment:
[9 empty] | Time deploy/retrieval | 10:10 |  |  | 7:05 |  |  | Turbidity: 0.9 FNU, sunny, wind N @ 5mph, high-tide falling. *pH probe for discrete samples malfunctioning, value higher than actual.
[9 empty] | First/last sample time | 10:15 |  |  | 7:00
| H= heavy, M= moderate, L= light | [7 empty] | Technician initials | MP |  |  | MP
Mount: N | [9 empty] | Discrete Samples (Required)
[9 empty] | Depth (m) | 2.7 | m |  | 2.5 |  | m
Sonde Body: N | [8 empty] | Temp (°C) | 28.4 | °C |  | 25.8 |  | °C
[9 empty] | DO % |  | % |  | 85.1 |  | %
Field guard: A/L, B/L, C/L, MD/L | [8 empty] | DO mg/L | 6.53 | mg/L |  | 5.83 |  | mg/L
All Sesnors: | [8 empty] | SpCond (mS/cm) | 48.371 | mS/cm |  | 47.305 |  | mS/cm | Retrieval:
C/T: A/M | [8 empty] | Salinity (psu) | 31.45 | psu |  | 30.76 |  | psu | Turbidity: 9.6 FNU, sunny, wind SW 5-10 mph, choppy, low tide ~+1.5 feet.
pH: A/M, MD/L | [8 empty] | pH | 9.06* |  |  | 7.84
DO: A/M, B/L | [11 empty] | Optional Samples
Turb: A/M, B/L | [8 empty] | pHmV |  | mV |  |  |  | mV
Wiper: | [8 empty] | Conductivity |  | mS/cm |  |  |  | mS/cm
TAL-PE:
Other/notes:

## CPFS_071322
Aquatic Preserve Continuous Water Quality Monitoring Calibration Log
|  | File Name | CPFS_071322 | [5 empty] | .csv | [5 empty] | File Name | CPFS_071322.DAT |  |  |  |  | .bin
[7 empty] | Datasonde and Probe Identification Numbers | [12 empty] | Calibration Standards
|  | Sonde Code |  |  | Serial Number |  |  | Serial Number |  | Model Number | [5 empty] | Standard | Lot No. |  |  | Received | Opened | Exp. Date
| Datasonde | 145458 |  |  | 13C100686 |  | pH | 21A |  | 6151 |  |  |  |  | SpCond (mS/cm) | 20 | 4203H74 |  |  |  | 4/11/2022 | 3/2024
| Vented | N |  |  | Model |  | ODO | 14D100738 |  | 6150 |  |  |  |  | SpCond (mS/cm) | 50 | 2908A58 |  |  |  |  | 6/2021
| Nickname | AH18 |  |  | 6600 |  | Turb | 14C103217 |  | 6136 |  |  |  |  | pH | 4
| Preserve | CPAP |  |  |  |  | Cond | 16G100618 |  | 6560 |  |  |  |  | pH | 7 | 21M2S |  |  |  | 7/11/2022 | 12/8/2023
| Station | CPFS |  |  |  |  | Chl | Chl Serial No. |  | Chl Model |  |  |  |  | pH | 10 | 22A1T |  |  |  | 7/11/2022 | 1/3/2024
[7 empty] | Wiper | Wiper Serial No. |  | Wiper Model |  |  |  |  | Turb (FNU) | 124 | 22D22130237 |  |  |  |  | 4/2023
|  |  |  | Wiper Replaced | Y |  |  |  | Batteries Replaced | Y | [5 empty] | Technician(s) | Michael Palandri
|  |  |  | Wiper Parked in Garage |  |  |  |  | Date of Calibration | 2022-07-11 | [5 empty] | Date of Post-Calibration | 2022-08-05
| Pre-deployment Comments/ Failed-calibration values: | SpCond failed in air, 0.006, is this requirement for 6600? SpCond 20mS failed need new 50mS/cm. Keep eye on turbidity probe, was reading off before calibrated.
| Pre-Deployment (CAL & ICV) | [12 empty] | Post-Deployment (CCV) | [5 empty] | Programming
|  | Standards |  |  | Before Cal |  |  | Calibrated |  | Specs |  | P/ F | Time |  | Post value |  | P/ F | Time |  | Interval |  | 15
| Temp | 23.6 | °C (Therm) |  | 23.35 | °C (DL) |  |  |  | ±0.2(w) +0.1°C |  | P | 12:49 |  | 21.4 | °C (Therm) | P | 11:23 |  | Duration |  | 365 days
| SpCond in air |  |  |  | 0.006 |  |  |  |  | ≤ 0.002 mS/cm |  | F |  |  | 21.21 | °C (DL) |  |  |  | Free memory %
| SpCond | 50 | mS/cm |  | 50.5 | mS/cm |  | 50 | mS/cm | 50 ± 0.25 (n) |  |  | 12:56 | ✔ | 49.25 | mS/cm | F | 11:43 |  | Memory (days) |  | 185
[6 empty] | Cell Constant |  | 4.945 |  | 0.419-0.519 (w) |  | P | [7 empty] | Start time |  | 6:14
| SpCond 20 (ICV) | 20 | mS/cm |  |  | ✔ |  | 20.4 | mS/cm | 20 ± 0.2 (w) |  | F | 12:59 | [6 empty] | Battery % |  | 58.3 days
| pH 7 | 7 |  |  | 7 |  |  | 7 |  | ± 0.1 |  |  | 13:02 | ✔ | 7.14 |  | F | 11:49 |  | Set clock |  | Y
| pH 10 | 10 |  |  | 10.07 |  |  | 10 |  | ± 0.1 |  |  | 13:05 | ✔ | 10.02 |  | P | 11:53 |  | End time |  | 10:20
[5 empty] | pH 7 |  |  | 14.6 | mV | 0 ±50 mV |  | P |  |  | 6.8 | mV | P |  |  |  | Other/notes
[5 empty] | pH 10 |  |  | -158.7 | mV | -180 ±50 mV |  | P |  |  | -160 | mV | P
[5 empty] | pH 4 |  |  |  | mV | +180 ±50 mV
[5 empty] | pH Slope |  |  | 173.3 |  | > 155 mV |  | P |  |  | 166.8 |  | P
| pH 4 or 10 (ICV) | 10 |  |  |  | ✔ |  | 9.99 |  | ± 0.1 |  | P | 13:12
| BP at Depth Cal |  |  |  | 758.54 | mmHg | [8 empty] | 764.4 | mmHg
| Depth | -0.014 | Offset |  | -0.023 | m |  | -0.014 | m | ±0.004 m |  | P | 13:32 |  | 0.071 | m | P | 11:30
| Turbidity | 0 | NTU |  | 5.5 |  |  | -0.1 | NTU | 0.00 ±0.3 |  |  | 13:24 | ✔ | -5.8 | FNU | F | 11:57 |  |  | General Maintenance
| Turbidity | 126 | NTU |  | 119.3 | NTU |  | 126 | NTU | 124 ±2.48 |  |  | 13:28 | ✔ | 125.8 | FNU | P | 11:59 | Greased locks and ram rod tube on deployment and retrieval.
| Turbidity 124 (ICV) | 126 | NTU |  |  |  | ✔ | 126.2 | NTU | 124 ±2.48 |  | P | 13:29
| BP at ODO Cal |  |  |  | 761.61 | mmHg | [8 empty] | 763.4 | mmHg
| Optical DO % |  |  |  | 98.9 | % |  | 100.2 | % | ± 1% |  | P | 5:56 |  | 98.2 | % | F | 11:58
[6 empty] | DO gain |  | 0.97036 | mg/L (table) | 0.87-1.25 |  | P |  |  | 98.4 | %
| Optical DO mg/L (ICV) | 8.92 | mg/L (DL) |  | 21.01 | °C |  | 8.91 |  | ± 0.1 mg/L |  | P | 5:56 |  | 7.75 | mg/L (DL)
| Battery Voltage | [6 empty] | 11.2 | V | ~6V (EXO 2) |  | P |  |  | 27.48 | °C | P | 11:58
[6 empty] | Optional Parameters | [8 empty] | 10 | V | P
| Chlorophyll | 0 | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| Chlorophyll |  | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [5 empty] | °C
| Chlorophyll | 0 | µg/L |  |  | µg/L |  |  | µg/L | [6 empty] | µg/L |  |  |  |  | Post-Deployment Notes
| Chlorophyll |  | µg/L |  |  | µg/L |  |  | µg/L | [6 empty] | µg/L |  |  | pH bulb very fouled. Turbidity in 0NTU reading ~5 during precal, which could explain the very negative post-cal reading? Possible bad 0NTU standard on calibration? Depth would pass at 765mmHg so called it a passing reading.
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [5 empty] | °C
| BGA | 0 | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| BGA |  | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [5 empty] | °C
Field Log
Fouling Presence | [9 empty] | Deployment |  |  | Retrieval |  |  |  | Field Notes:
A=algae, AM=amphipods, B=barnacles, BR=bryozoans, C=crabs, E=eggs, F=fish, FA=filamentous algae, G=gastropod, H=hydroids, I=isopod, MD=mud, S=sponges, SI=silt, SL=shell, SP=shrimp, T=tunicates, W=worm, O=other | [8 empty] | Date deploy/retrieval | 2022-07-13 |  |  | 2022-08-03 |  |  | Deployment:
[9 empty] | Time deploy/retrieval | 7:05 |  |  | 8:52 |  |  | Turbidity: 9.6 FNU, sunny, wind SW 10-15, choppy, high tide.
[9 empty] | First/last sample time | 7:15 |  |  | 8:45
| H= heavy, M= moderate, L= light | [7 empty] | Technician initials | MP |  |  | MP
Mount: N/A | [9 empty] | Discrete Samples (Required)
[9 empty] | Depth (m) | 2.5 | m |  | 2.7 |  | m
Sonde Body: B/L, MD/L | [8 empty] | Temp (°C) | 25.8 | °C |  | 27.5 |  | °C
[9 empty] | DO % | 85.1 | % |  | 68.1 |  | %
Field guard: A/L, B/L, C/M, E/L, MD/L, SP/M | [8 empty] | DO mg/L | 5.83 | mg/L |  | 4.6 |  | mg/L
All Sesnors: | [8 empty] | SpCond (mS/cm) | 47.305 | mS/cm |  | 45.837 |  | mS/cm | Retrieval:
C/T: A/L, B/L, MD/L | [8 empty] | Salinity (psu) | 30.76 | psu |  | 29.64 |  | psu | Turbidity: 2.4FNU, sunny, wind calm 0-5mph NW, high tide falling.
pH: A/L, B/L, MD/L | [8 empty] | pH | 7.84 |  |  | 8.08
DO: A/L, B/M, MD/L | [11 empty] | Optional Samples
Turb: A/L, B/M, MD/L | [8 empty] | pHmV |  | mV |  |  |  | mV
Wiper: | [8 empty] | Conductivity |  | mS/cm |  |  |  | mS/cm
TAL-PE:
Other/notes:

## CPFS_080322
Aquatic Preserve Continuous Water Quality Monitoring Calibration Log
|  | File Name | CPFS_080322 | [5 empty] | .csv | [5 empty] | File Name | CPFS_080322.DAT |  |  |  |  | .bin
[7 empty] | Datasonde and Probe Identification Numbers | [12 empty] | Calibration Standards
|  | Sonde Code |  |  | Serial Number |  |  | Serial Number |  | Model Number | [5 empty] | Standard | Lot No. |  |  | Received | Opened | Exp. Date
| Datasonde | 145459 |  |  | 13C100687 |  | pH | 19E |  | 6151 |  |  |  |  | SpCond (mS/cm) | 20 | 4203H74 |  |  |  | 4/11/2022 | 3/2024
| Vented | N |  |  | Model |  | ODO | 15L103284 |  | 6150 |  |  |  |  | SpCond (mS/cm) | 50 | 2908A58 |  |  |  |  | 6/2021
| Nickname | FS19 |  |  | 6600 |  | Turb | 11H101089 |  | 6136 |  |  |  |  | pH | 4
| Preserve | CPAP |  |  |  |  | Cond | 14J100146 |  | 6560 |  |  |  |  | pH | 7 | 21M2S |  |  |  | 7/11/2022 | 12/8/2023
| Station | CPFS |  |  |  |  | Chl | [7 empty] | pH | 10 | 22A1T |  |  |  | 7/11/2022 | 1/3/2024
[7 empty] | Wiper | [7 empty] | Turb (FNU) | 124 | 22D22130237 |  |  |  |  | 4/2023
|  |  |  | Wiper Replaced | Y |  |  |  | Batteries Replaced | N | [5 empty] | Technician(s) | Michael Palandri
|  |  |  | Wiper Parked in Garage |  |  |  |  | Date of Calibration | 2022-08-02 | [5 empty] | Date of Post-Calibration | 2022-08-23
| Pre-deployment Comments/ Failed-calibration values: | DO probe from WD10 placed on sonde. ICV 20mS/cm high need new 50mS/cm standard. *****Turbidity probe not working well, several issues with all probes on multiple sondes, couldn't get anything to work on any sondes with any probes. This turb probe was sent out anyways.
| Pre-Deployment (CAL & ICV) | [12 empty] | Post-Deployment (CCV) | [5 empty] | Programming
|  | Standards |  |  | Before Cal |  |  | Calibrated |  | Specs |  | P/ F | Time |  | Post value |  | P/ F | Time |  | Interval |  | 15
| Temp | 21.1 | °C (Therm) |  | 20.8 | °C (DL) |  |  |  | ±0.2(w) +0.1°C |  | P | 13:18 |  | 22.6 | °C (Therm) | F | 9:11 |  | Duration |  | 365 days
| SpCond in air |  |  |  | 0.002 |  |  |  |  | ≤ 0.002 mS/cm |  | P |  |  | 22.04 | °C (DL) |  |  |  | Free memory %
| SpCond | 50 | mS/cm |  | 49.86 | mS/cm |  | 50 | mS/cm | 50 ± 0.25 (n) |  |  | 13:25 | ✔ | 49.75 | mS/cm | P | 8:35 |  | Memory (days) |  | 185
[6 empty] | Cell Constant |  | 4.9387 |  | 0.419-0.519 (w) |  | P | [7 empty] | Start time |  | 7:30
| SpCond 20 (ICV) | 20 | mS/cm |  |  | ✔ |  | 20.48 | mS/cm | 20 ± 0.2 (w) |  | F | 13:29 | [6 empty] | Battery % |  | 56.8 days
| pH 7 | 7 |  |  | 6.96 |  |  | 7 |  | ± 0.1 |  |  | 13:36 | ✔ | 7.24 |  | F | 8:38 |  | Set clock |  | Y
| pH 10 | 10 |  |  | 9.89 |  |  | 10 |  | ± 0.1 |  |  | 13:39 | ✔ | 10.25 |  | F | 8:49 |  | End time |  | 13:26
[5 empty] | pH 7 |  |  | 10 | mV | 0 ±50 mV |  | P |  |  | -8.4 | mV | P |  |  |  | Other/notes
[5 empty] | pH 10 |  |  | -155 | mV | -180 ±50 mV |  | P |  |  | -169.3 | mV | P
[5 empty] | pH 4 |  |  |  | mV | +180 ±50 mV
[5 empty] | pH Slope |  |  | 165 |  | > 155 mV |  | P |  |  | 160.9 |  | P
| pH 4 or 10 (ICV) | 10 |  |  |  | ✔ |  | 9.97 |  | ± 0.1 |  | P | 13:43
| BP at Depth Cal |  |  |  | 763.19 | mmHg | [8 empty] | 760.26 | mmHg
| Depth | 0.041 | Offset |  | 0.05 | m |  | 0.041 | m | ±0.004 m |  | P | 13:52 |  | 0 | m | P | 8:50
| Turbidity | 0 | FNU | [5 empty] | FNU | 0.00 ±0.3 |  |  |  | ✔ |  | FNU |  |  |  |  | General Maintenance
| Turbidity | 124 | FNU |  |  | FNU |  |  | FNU | 124 ±2.48 |  |  |  | ✔ |  | FNU |  |  | Ram rod and lock grease on deployment and retrieval.
| Turbidity 124 (ICV) | 124 | FNU |  |  |  | ✔ |  | FNU | 124 ±2.48
| BP at ODO Cal |  |  |  | 762.44 | mmHg | [8 empty] | 760.11 | mmHg
| Optical DO % |  |  |  | 96.2 | % |  | 100.3 | % | ± 1% |  | P | 7:21 |  | 99.7 | % | P | 7:57
[6 empty] | DO gain |  | 1.0207 | mg/L (table) | 0.87-1.25 |  | P |  |  | 99.6 | %
| Optical DO mg/L (ICV) | 8.9 | mg/L (DL) |  | 21.18 | °C |  | 8.915 |  | ± 0.1 mg/L |  | P | 7:22 |  | 8.89 | mg/L (DL)
| Battery Voltage | [6 empty] | 11.3 | V | ~6V (EXO 2) |  | P |  |  | 20.93 | °C | P | 7:57
[6 empty] | Optional Parameters | [8 empty] | 10.7 | V | P
| Chlorophyll | 0 | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| Chlorophyll |  | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [5 empty] | °C
| Chlorophyll | 0 | µg/L |  |  | µg/L |  |  | µg/L | [6 empty] | µg/L |  |  |  |  | Post-Deployment Notes
| Chlorophyll |  | µg/L |  |  | µg/L |  |  | µg/L | [6 empty] | µg/L |  |  | Took post-cal readings for turbidity to see what probe was doing: 0 NTU Standard: 36, then with new water 18.3. 126 NTU Standard: 124.3 (passing value). C/T probe also reading low temp, even after cleaning.
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [5 empty] | °C
| BGA | 0 | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| BGA |  | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [5 empty] | °C
Field Log
Fouling Presence | [9 empty] | Deployment |  |  | Retrieval |  |  |  | Field Notes:
A=algae, AM=amphipods, B=barnacles, BR=bryozoans, C=crabs, E=eggs, F=fish, FA=filamentous algae, G=gastropod, H=hydroids, I=isopod, MD=mud, S=sponges, SI=silt, SL=shell, SP=shrimp, T=tunicates, W=worm, O=other | [8 empty] | Date deploy/retrieval | 2022-08-03 |  |  | 2022-08-19 |  |  | Deployment:
[9 empty] | Time deploy/retrieval | 8:56 |  |  | 8:07 |  |  | Turbidity: 2.4 FNU, sunny, wind calm 0-5mph NW, high tide falling.
[9 empty] | First/last sample time | 9:00 |  |  | 8:00
| H= heavy, M= moderate, L= light | [7 empty] | Technician initials | MP |  |  | MP
Mount: N/A | [9 empty] | Discrete Samples (Required)
[9 empty] | Depth (m) | 2.7 | m |  | 2.6 |  | m
Sonde Body: N/A | [8 empty] | Temp (°C) | 27.5 | °C |  | 29 |  | °C
[9 empty] | DO % | 68.6 | % |  | 86.9 |  | %
Field guard: A/M, B/L, C/M, SL/L | [8 empty] | DO mg/L | 4.6 | mg/L |  | 5.66 |  | mg/L
All Sesnors: | [8 empty] | SpCond (mS/cm) | 45.837 | mS/cm |  | 46.333 |  | mS/cm | Retrieval:
C/T: A/M, B/L | [8 empty] | Salinity (psu) | 29.64 | psu |  | 29.96 |  | psu | Turbidity: 2.6 NTU, calm, no wind, sunny, high tide
pH: A/M, B/L | [8 empty] | pH | 8.08 |  |  | 7.9
DO: A/L, B/L | [11 empty] | Optional Samples
Turb: A/L, B/M | [8 empty] | pHmV |  | mV |  |  |  | mV
Wiper: | [8 empty] | Conductivity |  | mS/cm |  |  |  | mS/cm
TAL-PE:
Other/notes:

## CPFS_081922
Aquatic Preserve Continuous Water Quality Monitoring Calibration Log
|  | File Name | CPFS_081922 | [5 empty] | .csv | [5 empty] | File Name | CPFS_081922.DAT |  |  |  |  | .bin
[7 empty] | Datasonde and Probe Identification Numbers | [12 empty] | Calibration Standards
|  | Sonde Code |  |  | Serial Number |  |  | Serial Number |  | Model Number | [5 empty] | Standard | Lot No. |  |  | Received | Opened | Exp. Date
| Datasonde | 145458 |  |  | 13C100686 |  | pH | 21A |  | 6151 |  |  |  |  | SpCond (mS/cm) | 20 | 4203H74 |  |  |  | 4/11/2022 | 3/2024
| Vented | N |  |  | Model |  | ODO | 14D100738 |  | 6150 |  |  |  |  | SpCond (mS/cm) | 50 | 2908A58 |  |  |  |  | 6/2021
| Nickname | AH18 |  |  | 6600 |  | Turb | 14C103217 |  | 6136 |  |  |  |  | pH | 4
| Preserve | CPAP |  |  |  |  | Cond | 16G100618 |  | 6560 |  |  |  |  | pH | 7 | 21M2S |  |  |  |  | 12/8/2023
| Station | CPFS |  |  |  |  | Chl | [7 empty] | pH | 10 | 22A1T |  |  |  |  | 1/3/2024
[7 empty] | Wiper | [7 empty] | Turb (FNU) | 124 | 22D22130237 |  |  |  |  | 4/2023
|  |  |  | Wiper Replaced | Y |  |  |  | Batteries Replaced | Y | [5 empty] | Technician(s) | Michael Palandri
|  |  |  | Wiper Parked in Garage |  |  |  |  | Date of Calibration | 2022-08-18 | [5 empty] | Date of Post-Calibration | 2022-09-26
| Pre-deployment Comments/ Failed-calibration values: | 0.000mS/cm in air same for 6600s? Probe is reading fine although read higher. 20mS/ICV failed because need new 50mS/cm standard, has been ordered.
| Pre-Deployment (CAL & ICV) | [12 empty] | Post-Deployment (CCV) | [5 empty] | Programming
|  | Standards |  |  | Before Cal |  |  | Calibrated |  | Specs |  | P/ F | Time |  | Post value |  | P/ F | Time |  | Interval |  | 15
| Temp | 25 | °C (Therm) |  | 24.8 | °C (DL) |  |  |  | ±0.2(w) +0.1°C |  | P | 11:37 |  | 22.7 | °C (Therm) | P | 6:07 |  | Duration |  | 365 days
| SpCond in air |  |  |  | 0.006 |  |  |  |  | ≤ 0.002 mS/cm |  | F |  |  | 22.28 | °C (DL) |  |  |  | Free memory %
| SpCond | 50 | mS/cm |  | 50.14 | mS/cm |  | 49.99 | mS/cm | 50 ± 0.25 (n) |  |  | 11:47 | ✔ | 50.04 | mS/cm | P | 6:11 |  | Memory (days) |  | 162
[6 empty] | Cell Constant |  | 4.9312 |  | 0.419-0.519 (w) |  | P | [7 empty] | Start time |  | 6:46
| SpCond 20 (ICV) | 20 | mS/cm |  |  | ✔ |  | 20.4 | mS/cm | 20 ± 0.2 (w) |  | F | 11:50 | [6 empty] | Battery % |  | 62.8 days
| pH 7 | 7 |  |  | 6.92 |  |  | 7 |  | ± 0.1 |  |  | 11:54 | ✔ | 7.04 |  | P | 6:13 |  | Set clock |  | Y
| pH 10 | 10 |  |  | 9.98 |  |  | 10 |  | ± 0.1 |  |  | 11:59 | ✔ | 9.95 |  | P | 6:15 |  | End time
[5 empty] | pH 7 |  |  | 19.5 | mV | 0 ±50 mV |  | P |  |  | 17.4 | mV | P |  |  |  | Other/notes
[5 empty] | pH 10 |  |  | -153 | mV | -180 ±50 mV |  | P |  |  | -150.4 | mV | P
[5 empty] | pH 4 |  |  |  | mV | +180 ±50 mV
[5 empty] | pH Slope |  |  | 172.5 |  | > 155 mV |  | P |  |  | 167.8 |  | P
| pH 4 or 10 (ICV) | 10 |  |  |  | ✔ |  | 10 |  | ± 0.1 |  | P | 12:02
| BP at Depth Cal |  |  |  | 761.01 | mmHg | [8 empty] | 758.95 | mmHg
| Depth | 0.014 | Offset |  | 0.025 | m |  | 0.014 | m | ±0.004 m |  | P | 12:03 |  | -0.02 | m | P | 6:11
| Turbidity | 0 | FNU |  | -5.5 |  |  | 0 | FNU | 0.00 ±0.3 |  |  | 12:14 | ✔ | 0.2 | FNU | P | 6:17 |  |  | General Maintenance
| Turbidity | 124 | FNU |  | 129.7 | FNU |  | 126 | FNU | 124 ±2.48 |  |  | 12:17 | ✔ | 127.1 | FNU | P | 6:18 | Lock grease and ram rod on deployment
| Turbidity 124 (ICV) | 124 | FNU |  |  |  | ✔ | 125.8 | FNU | 124 ±2.48 |  | P | 12:19
| BP at ODO Cal |  |  |  | 760.56 | mmHg | [8 empty] | 758.95 | mmHg
| Optical DO % |  |  |  | 99.5 | % |  | 100 | % | ± 1% |  | P | 6:38 |  | 99.7 | % | P | 6:07
[6 empty] | DO gain |  | 0.97588 | mg/L (table) | 0.87-1.25 |  | P |  |  | 99.6 | %
| Optical DO mg/L (ICV) | 8.59 | mg/L (DL) |  | 22.9 | °C |  | 8.578 |  | ± 0.1 mg/L |  | P | 6:43 |  | 8.64 | mg/L (DL)
| Battery Voltage | [6 empty] | 12.9 | V | ~6V (EXO 2) |  | P |  |  | 22.3 | °C | P | 6:08
[6 empty] | Optional Parameters | [8 empty] | 11.1 | V | P
| Chlorophyll | 0 | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| Chlorophyll |  | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [5 empty] | °C
| Chlorophyll | 0 | µg/L |  |  | µg/L |  |  | µg/L | [6 empty] | µg/L |  |  |  |  | Post-Deployment Notes
| Chlorophyll |  | µg/L |  |  | µg/L |  |  | µg/L | [6 empty] | µg/L |  |  | Guard slid up slightly so barnacles grew at the base and did not impact sensors.
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [5 empty] | °C
| BGA | 0 | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| BGA |  | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [5 empty] | °C
Field Log
Fouling Presence | [9 empty] | Deployment |  |  | Retrieval |  |  |  | Field Notes:
A=algae, AM=amphipods, B=barnacles, BR=bryozoans, C=crabs, E=eggs, F=fish, FA=filamentous algae, G=gastropod, H=hydroids, I=isopod, MD=mud, S=sponges, SI=silt, SL=shell, SP=shrimp, T=tunicates, W=worm, O=other | [8 empty] | Date deploy/retrieval | 2022-08-19 |  |  | 2022-09-20 |  |  | Deployment:
[9 empty] | Time deploy/retrieval | 8:12 |  |  | 7:55 |  |  | Turbidity: 2.0 NTU, calm no wind, sunny, high tide.
[9 empty] | First/last sample time | 8:15 |  |  | 7:45
| H= heavy, M= moderate, L= light | [7 empty] | Technician initials | MP |  |  | JB
Mount: None | [9 empty] | Discrete Samples (Required)
[9 empty] | Depth (m) | 2.6 | m |  | 2.9 |  | m
Sonde Body: None | [8 empty] | Temp (°C) | 29 | °C |  | 27.4 |  | °C
[9 empty] | DO % | 86.9 | % |  | 89.9 |  | %
Field guard: L-B on bottom of guard where it slid up | [8 empty] | DO mg/L | 5.66 | mg/L |  | 6.04 |  | mg/L
All Sesnors: | [8 empty] | SpCond (mS/cm) | 46.333 | mS/cm |  | 45.193 |  | mS/cm | Retrieval: Turb 2.94 FNU
C/T: None | [8 empty] | Salinity (psu) | 29.96 | psu |  | 29.18 |  | psu
pH:None | [8 empty] | pH | 7.9 |  |  | 8.06
DO:None | [11 empty] | Optional Samples
Turb:None | [8 empty] | pHmV |  | mV |  |  |  | mV
Wiper: None | [8 empty] | Conductivity |  | mS/cm |  |  |  | mS/cm
TAL-PE:
Other/notes:

## CPFS_092022
Aquatic Preserve Continuous Water Quality Monitoring Calibration Log
|  | File Name | CPFS_0920 | [5 empty] | .csv | [5 empty] | File Name | CPFS_092022.DAT |  |  |  |  | .bin
[7 empty] | Datasonde and Probe Identification Numbers | [12 empty] | Calibration Standards
|  | Sonde Code |  |  | Serial Number |  |  | Serial Number |  | Model Number | [5 empty] | Standard | Lot No. |  |  | Received | Opened | Exp. Date
| Datasonde | 145459 |  |  | 13C100687 |  | pH | 19F |  | 6151 |  |  |  |  | SpCond (mS/cm) | 20 | 4203H74 |  |  |  | 4/11/2022 | 3/2024
| Vented | N |  |  | Model |  | ODO | 15L103284 |  | 6150 |  |  |  |  | SpCond (mS/cm) | 50 | 2908A58 |  |  |  |  | 6/2021
| Nickname | FS19 |  |  | 6600 |  | Turb | 11H101089 |  | 6136 |  |  |  |  | pH | 4
| Preserve | CPAP |  |  |  |  | Cond | 14J100146 |  | 6560 |  |  |  |  | pH | 7 | 21M2S |  |  |  | 7/11/2022 | 12/8/2023
| Station | CPFS |  |  |  |  | Chl | [7 empty] | pH | 10 | 22A1T |  |  |  | 7/11/2022 | 1/3/2024
[7 empty] | Wiper | [7 empty] | Turb (FNU) | 124 | 22D22130237 |  |  |  |  | 4/2023
|  |  |  | Wiper Replaced | Y |  |  |  | Batteries Replaced | N | [5 empty] | Technician(s) | Jon Brucker
|  |  |  | Wiper Parked in Garage |  |  |  |  | Date of Calibration | 2022-09-12 | [5 empty] | Date of Post-Calibration
| Pre-deployment Comments/ Failed-calibration values: | Temp probe may be off. No DO probe deployed with this sonde. New pH probe (19F) installed - old one (21A) out of range (millivolts) in 10 standard
| Pre-Deployment (CAL & ICV) | [12 empty] | Post-Deployment (CCV) | [5 empty] | Programming
|  | Standards |  |  | Before Cal |  |  | Calibrated |  | Specs |  | P/ F | Time |  | Post value |  | P/ F | Time |  | Interval |  | 15
| Temp | 23 | °C (Therm) |  | 22.3 | °C (DL) |  |  |  | ±0.2(w) +0.1°C |  | P | 11:22 |  | 19.8 | °C (Therm) | F | 6:29 |  | Duration |  | 365 days
| SpCond in air |  |  |  | 0.001 |  |  |  |  | ≤ 0.002 mS/cm |  | P |  |  | 19.92 | °C (DL) |  |  |  | Free memory % |  | 100
| SpCond | 50 | mS/cm |  | 50.58 | mS/cm |  | 50 | mS/cm | 50 ± 0.25 (n) |  |  | 11:24 | ✔ | 49.1 | mS/cm | P | 6:29 |  | Memory (days) |  | 185
[6 empty] | Cell Constant |  | 50.0983 |  | 0.419-0.519 (w) |  | P | [7 empty] | Start time |  | 5:49
| SpCond 20 (ICV) | 20 | mS/cm |  |  | ✔ |  | 20.27 | mS/cm | 20 ± 0.2 (w) |  | F | 11:33 | [6 empty] | Battery % |  | 100
| pH 7 | 7 |  |  | 6.26 |  |  | 7 |  | ± 0.1 |  |  | 11:45 | ✔ | 7.01 |  | P | 6:34 |  | Set clock |  | Y
| pH 10 | 10 |  |  | 9.55 |  |  | 10 |  | ± 0.1 |  |  | 11:47 | ✔ | 9.26 |  | F | 6:37 |  | End time
[5 empty] | pH 7 |  |  | 47 | mV | 0 ±50 mV |  | P |  |  | 24.1 | mV | P |  |  |  | Other/notes
[5 empty] | pH 10 |  |  | -142.5 | mV | -180 ±50 mV |  | P |  |  | -86.4 | mV | P
[5 empty] | pH 4 |  |  |  | mV | +180 ±50 mV
[5 empty] | pH Slope |  |  | 189.5 |  | > 155 mV |  | P |  |  | 110.5 |  | P
| pH 4 or 10 (ICV) | 10 |  |  |  | ✔ |  | 9.93 |  | ± 0.1 |  | P | 11:49
| BP at Depth Cal |  |  |  | 760.73 | mmHg | [8 empty] | 763.016 | mmHg
| Depth | 0.014 | Offset |  | 0.061 | m |  | 0.014 | m | ±0.004 m |  | P | 11:50 |  | 0.036 | m | P | 6:30
| Turbidity | 0 | FNU |  | 0.1 |  |  | 0 | FNU | 0.00 ±0.3 |  |  | 11:52 | ✔ | 0 | FNU |  | 6:37 |  |  | General Maintenance
| Turbidity | 126 | FNU |  | 133.8 | FNU |  | 126 | FNU | 124 ±2.48 |  |  | 11:53 | ✔ | 128.2 | FNU |  | 6:39 | Ram rod and lock grease on deployment and retrieval.
| Turbidity 124 (ICV) | 126 | FNU |  |  |  | ✔ | 125.7 | FNU | 124 ±2.48 |  | P | 11:54
| BP at ODO Cal |  |  |  |  | mmHg | [9 empty] | mmHg
| Optical DO % |  |  |  |  | % |  |  | % | ± 1% |  | P |  |  |  | % | P
[6 empty] | DO gain |  |  | mg/L (table) | 0.87-1.25 |  | P |  |  |  | %
| Optical DO mg/L (ICV) |  | mg/L (DL) |  |  | °C |  |  |  | ± 0.1 mg/L |  | P |  |  |  | mg/L (DL)
| Battery Voltage | [7 empty] | V | ~6V (EXO 2) |  | P |  |  |  | °C | P
[6 empty] | Optional Parameters | [9 empty] | V | P
| Chlorophyll | 0 | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| Chlorophyll |  | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [5 empty] | °C
| Chlorophyll | 0 | µg/L |  |  | µg/L |  |  | µg/L | [6 empty] | µg/L |  |  |  |  | Post-Deployment Notes
| Chlorophyll |  | µg/L |  |  | µg/L |  |  | µg/L | [6 empty] | µg/L |  |  | pH 10 o.o.r. - not sure why. Will see what happens with next cal process. May need to recondition.
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [5 empty] | °C
| BGA | 0 | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| BGA |  | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [5 empty] | °C
Field Log
Fouling Presence | [9 empty] | Deployment |  |  | Retrieval |  |  |  | Field Notes:
A=algae, AM=amphipods, B=barnacles, BR=bryozoans, C=crabs, E=eggs, F=fish, FA=filamentous algae, G=gastropod, H=hydroids, I=isopod, MD=mud, S=sponges, SI=silt, SL=shell, SP=shrimp, T=tunicates, W=worm, O=other | [8 empty] | Date deploy/retrieval | 2022-09-20 |  |  | 2022-10-24 |  |  | Deployment:
[9 empty] | Time deploy/retrieval | 7:59 |  |  | 9:26 |  |  | Turbidity: 2.94
[9 empty] | First/last sample time | 8:00 |  |  | 9:45
| H= heavy, M= moderate, L= light | [7 empty] | Technician initials | JN |  |  | AL
Mount: N/A | [9 empty] | Discrete Samples (Required)
[9 empty] | Depth (m) | 2.9 | m |  | 3.1 |  | m
Sonde Body: N/A | [8 empty] | Temp (°C) | 29.18 | °C |  | 16.1 |  | °C
[9 empty] | DO % | 89.9 | % |  | 101.1 |  | %
Field guard: | [8 empty] | DO mg/L | 6.04 | mg/L |  | 8.13 |  | mg/L
All Sesnors: | [8 empty] | SpCond (mS/cm) | 45.193 | mS/cm |  | 50.836 |  | mS/cm | Retrieval:
C/T: | [8 empty] | Salinity (psu) | 29.18 | psu |  | 33.4 |  | psu | Turbidity:
pH: | [8 empty] | pH | 8.06 |  |  | 8.15
DO: | [11 empty] | Optional Samples
Turb: | [8 empty] | pHmV |  | mV |  |  |  | mV
Wiper: | [8 empty] | Conductivity |  | mS/cm |  |  |  | mS/cm
TAL-PE:
Other/notes:

## CPFS_102422
Aquatic Preserve Continuous Water Quality Monitoring Calibration Log
|  | File Name | CPFS_102422 | [5 empty] | .csv | [5 empty] | File Name | CPFS_102422.DAT |  |  |  |  | .bin
[7 empty] | Datasonde and Probe Identification Numbers | [12 empty] | Calibration Standards
|  | Sonde Code |  |  | Serial Number |  |  | Serial Number |  | Model Number | [5 empty] | Standard | Lot No. |  |  | Received | Opened | Exp. Date
| Datasonde | 145458 |  |  | 13C100686 |  | pH | 21A |  | 6151 |  |  |  |  | SpCond (mS/cm) | 20 | 4203H74 |  |  |  | 4/11/2022 | 3/2024
| Vented | N |  |  | Model |  | ODO | 14D100738 |  | 6150 |  |  |  |  | SpCond (mS/cm) | 50 | 2908A58 |  |  |  |  | 6/2021
| Nickname | AH18 |  |  | 6600 |  | Turb | 14C103217 |  | 6136 |  |  |  |  | pH | 4
| Preserve | CPAP |  |  |  |  | Cond | 16G100618 |  | 6560 |  |  |  |  | pH | 7 | 21M2S |  |  |  |  | 12/8/2023
| Station | CPFS |  |  |  |  | Chl | [7 empty] | pH | 10 | 22A1T |  |  |  |  | 1/3/2024
[7 empty] | Wiper | [7 empty] | Turb (FNU) | 124 | 22D22130237 |  |  |  |  | 4/2023
|  |  |  | Wiper Replaced | Y |  |  |  | Batteries Replaced | Y | [5 empty] | Technician(s) | Jon Brucker
|  |  |  | Wiper Parked in Garage | Y |  |  |  | Date of Calibration | 2022-10-10 | [5 empty] | Date of Post-Calibration | 2022-11-30
| Pre-deployment Comments/ Failed-calibration values: | Temperature probe may be off. 20mS/ICV failed because need new 50mS/cm standard, has been ordered.
| Pre-Deployment (CAL & ICV) | [12 empty] | Post-Deployment (CCV) | [5 empty] | Programming
|  | Standards |  |  | Before Cal |  |  | Calibrated |  | Specs |  | P/ F | Time |  | Post value |  | P/ F | Time |  | Interval |  | 15
| Temp | 22.5 | °C (Therm) |  | 22.02 | °C (DL) |  |  |  | ±0.2(w) +0.1°C |  | P | 12:41 |  | 20.69 | °C (Therm) | P | 11:24 |  | Duration |  | 365 days
| SpCond in air |  |  |  | 0.002 |  |  |  |  | ≤ 0.002 mS/cm |  | F |  |  | 20.34 | °C (DL) |  |  |  | Free memory % |  | 274432
| SpCond | 50 | mS/cm |  | 49.11 | mS/cm |  | 50 | mS/cm | 50 ± 0.25 (n) |  |  | 12:44 | ✔ | 48.3 | mS/cm | F | 10:24 |  | Memory (days) |  | 130
[6 empty] | Cell Constant |  | 5.0208 |  | 0.419-0.519 (w) |  | P | [7 empty] | Start time |  | 8:14
| SpCond 20 (ICV) | 20 | mS/cm |  |  | ✔ |  | 22.68 | mS/cm | 20 ± 0.2 (w) |  | F | 12:46 | [6 empty] | Battery % |  | 56.5
| pH 7 | 7 |  |  | 6.96 |  |  | 7 |  | ± 0.1 |  |  | 12:48 | ✔ | 6.33 |  | F | 10:31 |  | Set clock |  | Y
| pH 10 | 10 |  |  | 9.99 |  |  | 10 |  | ± 0.1 |  |  | 12:50 | ✔ | 5.7 |  | F | 10:32 |  | End time |  | 14:00
[5 empty] | pH 7 |  |  | 21.8 | mV | 0 ±50 mV |  | P |  |  | 59.1 | mV | F |  |  |  | Other/notes
[5 empty] | pH 10 |  |  | -149.5 | mV | -180 ±50 mV |  | P |  |  | 95.4 | mV | F |  | Turb probe was reading o.o.r. during call. Checked
[5 empty] | pH 4 |  |  |  | mV | +180 ±50 mV | [8 empty] | optical DO probe and port and discovered corrosion
[5 empty] | pH Slope |  |  | 171.3 |  | > 155 mV |  | P |  |  | -36.3 |  | F |  | Cleaned it up and probes all working correctly.
| pH 4 or 10 (ICV) | 10 |  |  |  | ✔ |  | 9.95 |  | ± 0.1 |  | P | 12:52
| BP at Depth Cal |  |  |  | 766.572 | mmHg | [8 empty] | 763.27 | mmHg
| Depth | 0.089 | Offset |  | 0.07 | m |  | 0.53681 | m | ±0.004 m |  | P | 12:53 |  | 0.065 | m | F | 10:25
| Turbidity | 0 | FNU |  | -0.8 |  |  | 0 | FNU | 0.00 ±0.3 |  |  | 11:44 | ✔ | -0.9 | FNU | F | 10:50 |  |  | General Maintenance
| Turbidity | 126 | FNU |  | 121.5 | FNU |  | 126 | FNU | 124 ±2.48 |  |  | 11:45 | ✔ | 137.6 | FNU | F | 10:54 | Lock grease and ram rod on deployment. Treated and cleaned DO probe port; had corrosion.
| Turbidity 124 (ICV) | 126 | FNU |  |  |  | ✔ | 125.8 | FNU | 124 ±2.48 |  | P | 11:46
| BP at ODO Cal |  |  |  | 766.572 | mmHg | [8 empty] | 765.81 | mmHg
| Optical DO % |  |  |  | 100.6 |  |  | 101 | % | ± 1% |  | P | 11:48 |  | 98.7 | % | P | 10:15
[6 empty] | DO gain |  | 0.97087 | mg/L (table) | 0.87-1.25 |  | P |  |  | 99 | %
| Optical DO mg/L (ICV) | 9.01 | mg/L (DL) |  | 20.75 | °C |  | 8.915 |  | ± 0.1 mg/L |  | P | 11:48 |  | 8.81 | mg/L (DL)
| Battery Voltage | [6 empty] | 11.1 | V | ~6V (EXO 2) |  | P |  |  | 20.06 | °C | P | 11:16
[6 empty] | Optional Parameters | [8 empty] | 10.3 | V | P
| Chlorophyll | 0 | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| Chlorophyll |  | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [5 empty] | °C
| Chlorophyll | 0 | µg/L |  |  | µg/L |  |  | µg/L | [6 empty] | µg/L |  |  |  |  | Post-Deployment Notes
| Chlorophyll |  | µg/L |  |  | µg/L |  |  | µg/L | [6 empty] | µg/L |  |  | Low pH readings; probe needs to be replaced. BP=763.27 ON 11/30/22; BP=765.81 ON 11/14/22
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [5 empty] | °C
| BGA | 0 | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| BGA |  | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [5 empty] | °C
Field Log
Fouling Presence | [9 empty] | Deployment |  |  | Retrieval |  |  |  | Field Notes:
A=algae, AM=amphipods, B=barnacles, BR=bryozoans, C=crabs, E=eggs, F=fish, FA=filamentous algae, G=gastropod, H=hydroids, I=isopod, MD=mud, S=sponges, SI=silt, SL=shell, SP=shrimp, T=tunicates, W=worm, O=other | [8 empty] | Date deploy/retrieval | 2022-10-24 |  |  | 2022-11-14 |  |  | Deployment:
[9 empty] | Time deploy/retrieval | 9:35 |  |  | 14:09 |  |  | Turbidity:
[9 empty] | First/last sample time | 9:30 |  |  | 14:00
| H= heavy, M= moderate, L= light | [7 empty] | Technician initials | JB |  |  | AL
Mount: None
[9 empty] | Depth (m) | 3.1 |  |  | 3.1 |  | m
Sonde Body: None | [8 empty] | Temp (°C) | 16.1 |  |  | 16.3 |  | °C
[9 empty] | DO % | 101.1 |  |  | 103.8 |  | %
Field guard: C-L, SP-M | [8 empty] | DO mg/L | 8.13 |  |  | 8.31 |  | mg/L
All Sesnors: | [8 empty] | SpCond (mS/cm) | 50.836 |  |  | 50.577 |  | mS/cm | Retrieval:
C/T: None | [8 empty] | Salinity (psu) | 33.4 |  |  | 33.21 |  | psu
pH:None | [8 empty] | pH | 8.15 |  |  | 8.11
DO:None | [11 empty] | Optional Samples
Turb:None | [8 empty] | pHmV |  | mV |  |  |  | mV
Wiper: None | [8 empty] | Conductivity |  | mS/cm |  |  |  | mS/cm
TAL-PE:
Other/notes:

## CPFS_111422
Aquatic Preserve Continuous Water Quality Monitoring Calibration Log
|  | File Name | CPFS_111422 | [5 empty] | .csv | [5 empty] | File Name | CPFS_111422.DAT |  |  |  |  | .bin
[7 empty] | Datasonde and Probe Identification Numbers | [12 empty] | Calibration Standards
|  | Sonde Code |  |  | Serial Number |  |  | Serial Number |  | Model Number | [5 empty] | Standard | Lot No. |  |  | Received | Opened | Exp. Date
| Datasonde | 145458 |  |  | 13C100686 |  | pH | 21A |  | 6151 |  |  |  |  | SpCond (mS/cm) | 20 | 4203H74 |  |  |  | 4/11/2022 | 3/2024
| Vented | N |  |  | Model |  | ODO | 14D100738 |  | 6150 |  |  |  |  | SpCond (mS/cm) | 50 | 2908A58 |  |  |  |  | 6/2021
| Nickname | AH18 |  |  | 6600 |  | Turb | 14C103217 |  | 6136 |  |  |  |  | pH | 4
| Preserve | CPAP |  |  |  |  | Cond | 16G100618 |  | 6560 |  |  |  |  | pH | 7 | 21M2S |  |  |  |  | 12/8/2023
| Station | CPFS |  |  |  |  | Chl | [7 empty] | pH | 10 | 22A1T |  |  |  |  | 1/3/2024
[7 empty] | Wiper | [7 empty] | Turb (FNU) | 124 | 22D22130237 |  |  |  |  | 4/2023
|  |  |  | Wiper Replaced | Y |  |  |  | Batteries Replaced | N | [5 empty] | Technician(s) | Jon Brucker
|  |  |  | Wiper Parked in Garage | Y |  |  |  | Date of Calibration | 2022-11-09 | [5 empty] | Date of Post-Calibration | 2022-12-08
| Pre-deployment Comments/ Failed-calibration values: | No DO probe installed. pH 10 o.o.r. - not able to calibrate to callibrate 10 (out of pH probes)
| Pre-Deployment (CAL & ICV) | [12 empty] | Post-Deployment (CCV) | [5 empty] | Programming
|  | Standards |  |  | Before Cal |  |  | Calibrated |  | Specs |  | P/ F | Time |  | Post value |  | P/ F | Time |  | Interval |  | 15
| Temp | 22.7 | °C (Therm) |  | 21.61 | °C (DL) |  |  |  | ±0.2(w) +0.1°C |  | F | 14:23 |  | 20.63 | °C (Therm) | F | 10:41 |  | Duration |  | 365 days
| SpCond in air |  |  |  | 0.535 |  |  |  |  | ≤ 0.002 mS/cm |  | F |  |  | 20.36 | °C (DL) |  |  |  | Free memory %
| SpCond | 50 | mS/cm |  | 48.91 | mS/cm |  | 50 | mS/cm | 50 ± 0.25 (n) |  |  | 14:27 | ✔ | 49.27 | mS/cm | P | 10:22 |  | Memory (days) |  | 147
[6 empty] | Cell Constant |  | 5.2118 |  | 0.419-0.519 (w) |  | P | [7 empty] | Start time |  | 13:14
| SpCond 20 (ICV) | 20 | mS/cm |  |  | ✔ |  | 19.21 | mS/cm | 20 ± 0.2 (w) |  | F | 14:29 | [6 empty] | Battery % |  | 58.1 (days)
| pH 7 | 7 |  |  | 6.93 |  |  | 7 |  | ± 0.1 |  |  | 14:30 | ✔ | 5.84 |  | P | 10:26 |  | Set clock |  | Y
| pH 10 | 10 | [5 empty] | 10 |  | ± 0.1 |  |  |  | ✔ | 7.85 |  | F | 10:29 |  | End time |  | 15:28
[5 empty] | pH 7 |  |  | 28.2 | mV | 0 ±50 mV |  |  |  |  | 85.1 | mV | F |  |  |  | Other/notes
[5 empty] | pH 10 |  |  |  | mV | -180 ±50 mV |  |  |  |  | -13.5 | mV | F |  | Turb probe was reading o.o.r. during call. Checked
[5 empty] | pH 4 |  |  |  | mV | +180 ±50 mV | [8 empty] | optical DO probe and port and discovered corrosion
[5 empty] | pH Slope |  |  |  |  | > 155 mV |  |  |  |  | 98.6 |  | F |  | Cleaned it up and probes all working correctly.
| pH 4 or 10 (ICV) | 10 |  |  |  | ✔ |  |  |  | ± 0.1 | [9 empty] | Battery: 58.1 days
| BP at Depth Cal |  |  |  | 766.04 | mmHg | [8 empty] | 755.25 | mmHg
| Depth | 0.082 | Offset |  | 0.045 | m |  | 0.082 | m | ±0.004 m |  | P | 14:46 |  | 0.137 | m | P | 10:22
| Turbidity | 0 | FNU |  | 0 |  |  | 0 | FNU | 0.00 ±0.3 |  |  | 14:41 | ✔ | 1.1 | FNU | F | 10:32 |  |  | General Maintenance
| Turbidity | 126 | FNU |  | 130.8 | FNU |  | 126 | FNU | 124 ±2.48 |  |  | 14:42 | ✔ | 129.3 | FNU | F | 10:35 | Lock grease and ram rod on deployment
| Turbidity 124 (ICV) | 126 | FNU |  |  |  | ✔ | 126.1 | FNU | 124 ±2.48 |  | P | 14:43
| BP at ODO Cal |  |  |  |  | mmHg | [8 empty] | 765.8 | mmHg
| Optical DO % | [7 empty] | % | ± 1% |  | P |  |  | 0 | % | P | 15:50
[6 empty] | DO gain |  |  | mg/L (table) | 0.87-1.25 |  | P |  |  | 0 | %
| Optical DO mg/L (ICV) |  | mg/L (DL) |  |  | °C |  |  |  | ± 0.1 mg/L |  | P |  |  | 0 | mg/L (DL)
| Battery Voltage | [6 empty] | 11.3 | V | ~6V (EXO 2) |  | P |  |  | 18.09 | °C | P | 15:51
[6 empty] | Optional Parameters | [8 empty] | 10.5 | V | P
| Chlorophyll | 0 | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| Chlorophyll |  | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [5 empty] | °C
| Chlorophyll | 0 | µg/L |  |  | µg/L |  |  | µg/L | [6 empty] | µg/L |  |  |  |  | Post-Deployment Notes
| Chlorophyll |  | µg/L |  |  | µg/L |  |  | µg/L | [6 empty] | µg/L
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [5 empty] | °C
| BGA | 0 | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| BGA |  | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [5 empty] | °C
Field Log
Fouling Presence | [9 empty] | Deployment |  |  | Retrieval |  |  |  | Field Notes:
A=algae, AM=amphipods, B=barnacles, BR=bryozoans, C=crabs, E=eggs, F=fish, FA=filamentous algae, G=gastropod, H=hydroids, I=isopod, MD=mud, S=sponges, SI=silt, SL=shell, SP=shrimp, T=tunicates, W=worm, O=other | [8 empty] | Date deploy/retrieval | 2022-11-14 |  |  | 2022-12-07 |  |  | Deployment:
[9 empty] | Time deploy/retrieval | 14:11 |  |  | 11:55 |  |  | Turbidity: 0.62
[9 empty] | First/last sample time | 14:15 |  |  | 11:45
| H= heavy, M= moderate, L= light | [7 empty] | Technician initials | AL |  |  | AL
Mount: None
[9 empty] | Depth (m) | 3.1 |  |  | 3.2 |  | m
Sonde Body: None | [8 empty] | Temp (°C) | 16.3 |  |  | 17.1 |  | °C
[9 empty] | DO % | 103.8 |  |  | 103.6 |  | %
Field guard: SP-H, C-M, G-L | [8 empty] | DO mg/L | 8.31 |  |  | 8.2 |  | mg/L
All Sesnors: | [8 empty] | SpCond (mS/cm) | 50577 |  |  | 49950 |  | mS/cm | Retrieval:
C/T: None | [8 empty] | Salinity (psu) | 33.21 |  |  | 32.77 |  | psu | Turbidity: 0.83
pH:None | [8 empty] | pH | 8.11 |  |  | 7.99
DO:None | [11 empty] | Optional Samples
Turb:None | [8 empty] | pHmV |  | mV |  |  |  | mV
Wiper: None | [8 empty] | Conductivity |  | mS/cm |  |  |  | mS/cm
TAL-PE:
Other/notes:

## CPFS_120722
Aquatic Preserve Continuous Water Quality Monitoring Calibration Log
|  | File Name | CPFS_120722 | [5 empty] | .csv | [5 empty] | File Name | CPFS_120722.DAT |  |  |  |  | .bin
[7 empty] | Datasonde and Probe Identification Numbers | [12 empty] | Calibration Standards
|  | Sonde Code |  |  | Serial Number |  |  | Serial Number |  | Model Number | [5 empty] | Standard | Lot No. |  |  | Received | Opened | Exp. Date
| Datasonde | 145458 |  |  | 13C100686 |  | pH | 21A |  | 6151 |  |  |  |  | SpCond (mS/cm) | 20 | 4203H74 |  |  |  | 4/11/2022 | 3/2024
| Vented | N |  |  | Model |  | ODO | 14D100738 |  | 6150 |  |  |  |  | SpCond (mS/cm) | 50 | 2908A58 |  |  |  |  | 6/2021
| Nickname | AH18 |  |  | 6600 |  | Turb | 14C103217 |  | 6136 |  |  |  |  | pH | 4
| Preserve | CPAP |  |  |  |  | Cond | 16G100618 |  | 6560 |  |  |  |  | pH | 7 | 21M2S |  |  |  |  | 12/8/2023
| Station | CPFS |  |  |  |  | Chl | [7 empty] | pH | 10 | 22A1T |  |  |  |  | 1/3/2024
[7 empty] | Wiper | [7 empty] | Turb (FNU) | 124 | 22D22130237 |  |  |  |  | 4/2023
|  |  |  | Wiper Replaced | Y |  |  |  | Batteries Replaced | Y | [5 empty] | Technician(s) | Jon Brucker
|  |  |  | Wiper Parked in Garage | Y |  |  |  | Date of Calibration | 2022-12-01 | [5 empty] | Date of Post-Calibration
| Pre-deployment Comments/ Failed-calibration values: | No DO probe installed. pH reading low; out of pH probes.
| Pre-Deployment (CAL & ICV) | [12 empty] | Post-Deployment (CCV) | [5 empty] | Programming
|  | Standards |  |  | Before Cal |  |  | Calibrated |  | Specs |  | P/ F | Time |  | Post value |  | P/ F | Time |  | Interval |  | 15
| Temp | 19.67 | °C (Therm) |  | 19.57 | °C (DL) |  |  |  | ±0.2(w) +0.1°C |  | P | 11:55 |  |  | °C (Therm) |  |  |  | Duration |  | 365 days
| SpCond in air |  |  |  | 0.076 |  |  |  |  | ≤ 0.002 mS/cm |  | F |  |  |  | °C (DL) |  |  |  | Free memory %
| SpCond | 50 | mS/cm |  | 49.86 | mS/cm |  | 50.1 | mS/cm | 50 ± 0.25 (n) |  |  | 11:08 | ✔ |  | mS/cm |  |  |  | Memory (days) |  | 106
[6 empty] | Cell Constant |  | 5.1448 |  | 0.419-0.519 (w) |  | F | [7 empty] | Start time
| SpCond 20 (ICV) | 20 | mS/cm |  |  | ✔ |  |  | mS/cm | 20 ± 0.2 (w) | [9 empty] | Battery % |  | 62.8 (days)
| pH 7 | 7 |  |  | 4.83 |  |  | 6.96 |  | ± 0.1 |  |  | 10:00 | ✔ | [5 empty] | Set clock |  | Y
| pH 10 | 10 |  |  | 8.14 |  |  | 10.18 |  | ± 0.1 |  |  | 10:05 | ✔ | [5 empty] | End time
[5 empty] | pH 7 |  |  | 149.1 | mV | 0 ±50 mV |  | F |  |  |  | mV |  |  |  |  | Other/notes
[5 empty] | pH 10 |  |  | 80.1 | mV | -180 ±50 mV |  | F |  |  |  | mV |  |  | Turb probe was reading o.o.r. during call. Checked
[5 empty] | pH 4 |  |  |  | mV | +180 ±50 mV | [8 empty] | optical DO probe and port and discovered corrosion
[5 empty] | pH Slope |  |  | 69 |  | > 155 mV |  | F | [6 empty] | Cleaned it up and probes all working correctly.
| pH 4 or 10 (ICV) | 10 |  |  |  | ✔ |  | 9.96 |  | ± 0.1 |  | P | 10:08 | [6 empty] | Battery: 62.8 days
| BP at Depth Cal |  |  |  | 771.14 | mmHg | [9 empty] | mmHg
| Depth |  | Offset |  | 0.135 | m |  | 0.15 | m | ±0.004 m |  | P | 11:30 |  |  | m
| Turbidity | 0 | FNU |  | -1.3 |  |  | 0 | FNU | 0.00 ±0.3 |  |  | 11:40 | ✔ |  | FNU |  |  |  |  | General Maintenance
| Turbidity | 126 | FNU |  | 126.9 | FNU |  | 126 | FNU | 124 ±2.48 |  |  | 11:44 | ✔ |  | FNU |  |  | Lock grease and ram rod on deployment
| Turbidity 124 (ICV) | 126 | FNU |  |  |  | ✔ | 125.4 | FNU | 124 ±2.48 |  | P | 11:46
| BP at ODO Cal |  |  |  |  | mmHg | [9 empty] | mmHg
| Optical DO % | [7 empty] | % | ± 1% |  | P |  |  |  | %
[6 empty] | DO gain |  |  | mg/L (table) | 0.87-1.25 |  | P |  |  |  | %
| Optical DO mg/L (ICV) |  | mg/L (DL) |  |  | °C |  |  |  | ± 0.1 mg/L |  | P |  |  |  | mg/L (DL)
| Battery Voltage | [6 empty] | 9.4 | V | ~6V (EXO 2) |  | P |  |  |  | °C
[6 empty] | Optional Parameters | [9 empty] | V
| Chlorophyll | 0 | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| Chlorophyll |  | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [5 empty] | °C
| Chlorophyll | 0 | µg/L |  |  | µg/L |  |  | µg/L | [6 empty] | µg/L |  |  |  |  | Post-Deployment Notes
| Chlorophyll |  | µg/L |  |  | µg/L |  |  | µg/L | [6 empty] | µg/L
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [5 empty] | °C
| BGA | 0 | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| BGA |  | RFU |  |  | RFU |  |  | RFU | [6 empty] | RFU
| Rhodamine WT Temp |  |  |  |  | °C |  |  | Linearity: | R2 ≥ 0.999 | [5 empty] | °C
Field Log
Fouling Presence | [9 empty] | Deployment |  |  | Retrieval |  |  |  | Field Notes:
A=algae, AM=amphipods, B=barnacles, BR=bryozoans, C=crabs, E=eggs, F=fish, FA=filamentous algae, G=gastropod, H=hydroids, I=isopod, MD=mud, S=sponges, SI=silt, SL=shell, SP=shrimp, T=tunicates, W=worm, O=other | [8 empty] | Date deploy/retrieval | 2022-12-07 | [5 empty] | Deployment:
[9 empty] | Time deploy/retrieval | 11:58 | [5 empty] | Turbidity: 0.83
[9 empty] | First/last sample time | 12:00
| H= heavy, M= moderate, L= light | [7 empty] | Technician initials | AL
Mount:
[9 empty] | Depth (m) | 3.2 |  |  |  |  | m
Sonde Body: | [8 empty] | Temp (°C) | 17.1 |  |  |  |  | °C
[9 empty] | DO % | 103.6 |  |  |  |  | %
Field guard: | [8 empty] | DO mg/L | 8.2 |  |  |  |  | mg/L
All Sesnors: | [8 empty] | SpCond (mS/cm) | 49.95 |  |  |  |  | mS/cm | Retrieval:
C/T: | [8 empty] | Salinity (psu) | 32.77 |  |  |  |  | psu | Turbidity:
pH: | [8 empty] | pH | 7.99
DO: | [11 empty] | Optional Samples
Turb: | [8 empty] | pHmV |  | mV |  |  |  | mV
Wiper: | [8 empty] | Conductivity |  | mS/cm |  |  |  | mS/cm
TAL-PE:
Other/notes: